## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 236
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.8485 | 0.059041 | 1.7304 | 1.9665 | 1.6713 | 2.0256 | 3.194% | 6.388% | 9.582% | 1.756 | 1.9409
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.7515 | 0.087372 | 1.5768 | 1.9263 | 1.4894 | 2.0136 | 4.988% | 9.977% | 14.96% | 1.664 | 1.8391
Cyanide Leach
Au, ppm | 1.7929 | 0.047766 | 1.6974 | 1.8884 | 1.6496 | 1.9362 | 2.664% | 5.328% | 7.993% | 1.7033 | 1.8825
PhotonAssay
Au, ppm | 1.8475 | 0.053436 | 1.7406 | 1.9544 | 1.6872 | 2.0078 | 2.892% | 5.785% | 8.677% | 1.7551 | 1.9399
4-Acid Digestion
Ag, ppm | 0.47814 | 0.057686 | 0.36277 | 0.59351 | 0.30508 | 0.6512 | 12.06% | 24.13% | 36.19% | 0.45423 | 0.50205
Al, wt.% | 6.5932 | 0.15166 | 6.2899 | 6.8965 | 6.1382 | 7.0481 | 2.3% | 4.6% | 6.901% | 6.2635 | 6.9228
As, ppm | 67.6896 | 3.9831 | 59.7234 | 75.6558 | 55.7404 | 79.6389 | 5.884% | 11.77% | 17.65% | 64.3051 | 71.0741
Ba, ppm | 245.7522 | 8.9752 | 227.8018 | 263.7027 | 218.8265 | 272.6779 | 3.652% | 7.304% | 10.96% | 233.4646 | 258.0398
Be, ppm | 0.45857 | 0.029291 | 0.39999 | 0.51715 | 0.37069 | 0.54644 | 6.388% | 12.78% | 19.16% | 0.43564 | 0.4815
Bi, ppm | 0.061389 | 0.0065661 | 0.048257 | 0.074521 | 0.04169 | 0.081087 | 10.7% | 21.39% | 32.09% | 0.058319 | 0.064458
Ca, wt.% | 5.9341 | 0.24588 | 5.4424 | 6.4259 | 5.1965 | 6.6717 | 4.143% | 8.287% | 12.43% | 5.6374 | 6.2308
Cd, ppm | 0.62376 | 0.044435 | 0.53489 | 0.71263 | 0.49046 | 0.75707 | 7.124% | 14.25% | 21.37% | 0.59257 | 0.65495
Ce, ppm | 13.9587 | 1.3597 | 11.2392 | 16.6782 | 9.8795 | 18.0379 | 9.741% | 19.48% | 29.22% | 13.2608 | 14.6566
Co, ppm | 41.4883 | 2.3618 | 36.7647 | 46.2119 | 34.4029 | 48.5738 | 5.693% | 11.39% | 17.08% | 39.4139 | 43.5627
Cr, ppm | 86.4016 | 4.255 | 77.8917 | 94.9115 | 73.6367 | 99.1664 | 4.925% | 9.849% | 14.77% | 82.0815 | 90.7217
Cs, ppm | 1.0081 | 0.072407 | 0.86325 | 1.1529 | 0.79084 | 1.2253 | 7.183% | 14.37% | 21.55% | 0.95766 | 1.0585
Cu, ppm | 169.6544 | 5.9576 | 157.7392 | 181.5696 | 151.7816 | 187.5272 | 3.512% | 7.023% | 10.53% | 161.1716 | 178.1371
Dy, ppm | 3.938 | 0.26366 | 3.4106 | 4.4653 | 3.147 | 4.7289 | 6.695% | 13.39% | 20.09% | 3.7411 | 4.1349
Er, ppm | 2.4286 | 0.15351 | 2.1216 | 2.7356 | 1.9681 | 2.8891 | 6.321% | 12.64% | 18.96% | 2.3072 | 2.55
Eu, ppm | 0.96342 | 0.029576 | 0.90427 | 1.0226 | 0.8747 | 1.0522 | 3.07% | 6.14% | 9.21% | 0.91525 | 1.0116
Fe, wt.% | 8.0044 | 0.2733 | 7.4578 | 8.551 | 7.1845 | 8.8243 | 3.414% | 6.829% | 10.24% | 7.6042 | 8.4047
Ga, ppm | 16.1386 | 1.6477 | 12.8432 | 19.434 | 11.1955 | 21.0817 | 10.21% | 20.42% | 30.63% | 15.3317 | 16.9455
Gd, ppm | 3.4747 | 0.23002 | 3.0146 | 3.9347 | 2.7846 | 4.1648 | 6.62% | 13.24% | 19.86% | 3.301 | 3.6484
Hf, ppm | 1.8009 | 0.081094 | 1.6387 | 1.9631 | 1.5577 | 2.0442 | 4.503% | 9.006% | 13.51% | 1.7109 | 1.891
Ho, ppm | 0.85917 | 0.032016 | 0.79514 | 0.9232 | 0.76312 | 0.95521 | 3.726% | 7.453% | 11.18% | 0.81621 | 0.90213
In, ppm | 0.079708 | 0.0081174 | 0.063473 | 0.095943 | 0.055356 | 0.10406 | 10.18% | 20.37% | 30.55% | 0.075723 | 0.083694
K, wt.% | 0.55896 | 0.019076 | 0.52081 | 0.59711 | 0.50173 | 0.61619 | 3.413% | 6.825% | 10.24% | 0.53101 | 0.58691
La, ppm | 5.9884 | 0.4624 | 5.0636 | 6.9132 | 4.6012 | 7.3756 | 7.722% | 15.44% | 23.16% | 5.689 | 6.2878
Li, ppm | 10.5106 | 0.61433 | 9.2819 | 11.7392 | 8.6676 | 12.3536 | 5.845% | 11.69% | 17.53% | 9.985 | 11.0361
Lu, ppm | 0.37163 | 0.028047 | 0.31554 | 0.42773 | 0.28749 | 0.45577 | 7.547% | 15.09% | 22.64% | 0.35305 | 0.39021
Mg, wt.% | 3.3955 | 0.10378 | 3.188 | 3.6031 | 3.0842 | 3.7069 | 3.056% | 6.113% | 9.169% | 3.2258 | 3.5653
Mn, wt.% | 0.13325 | 0.006762 | 0.11973 | 0.14678 | 0.11297 | 0.15354 | 5.075% | 10.15% | 15.22% | 0.12659 | 0.13992
Mo, ppm | 1.5601 | 0.19029 | 1.1795 | 1.9407 | 0.98925 | 2.131 | 12.2% | 24.39% | 36.59% | 1.4821 | 1.6381
Na, wt.% | 2.155 | 0.055982 | 2.043 | 2.2669 | 1.987 | 2.3229 | 2.598% | 5.196% | 7.794% | 2.0472 | 2.2627
Nb, ppm | 3.6281 | 0.2817 | 3.0647 | 4.1915 | 2.783 | 4.4732 | 7.764% | 15.53% | 23.29% | 3.4467 | 3.8095
Nd, ppm | 9.0119 | 0.462 | 8.0879 | 9.9359 | 7.6259 | 10.3979 | 5.126% | 10.25% | 15.38% | 8.5613 | 9.4625
Ni, ppm | 68.4509 | 4.2841 | 59.8828 | 77.019 | 55.5988 | 81.3031 | 6.259% | 12.52% | 18.78% | 65.0284 | 71.8735
P, wt.% | 0.045336 | 0.0025798 | 0.040176 | 0.050495 | 0.037597 | 0.053075 | 5.69% | 11.38% | 17.07% | 0.043069 | 0.047603
Pb, ppm | 30.4861 | 1.9905 | 26.5052 | 34.4671 | 24.5147 | 36.4575 | 6.529% | 13.06% | 19.59% | 28.9618 | 32.0104
Pr, ppm | 1.9404 | 0.10125 | 1.7379 | 2.1429 | 1.6367 | 2.2442 | 5.218% | 10.44% | 15.65% | 1.8434 | 2.0374
Rb, ppm | 14.0762 | 0.83178 | 12.4126 | 15.7397 | 11.5808 | 16.5715 | 5.909% | 11.82% | 17.73% | 13.3724 | 14.78
Re, ppm | 0.0026667 | 0.00086005 | 0.00094657 | 0.0043868 | 8.6515e-05 | 0.0052468 | 32.25% | 64.5% | 96.76% | 0.0025333 | 0.0028
S, wt.% | 0.44412 | 0.016634 | 0.41085 | 0.47739 | 0.39422 | 0.49402 | 3.745% | 7.491% | 11.24% | 0.42192 | 0.46633
Sb, ppm | 1.7628 | 0.14738 | 1.4681 | 2.0576 | 1.3207 | 2.205 | 8.36% | 16.72% | 25.08% | 1.6747 | 1.851
Sc, ppm | 38.5064 | 1.9102 | 34.686 | 42.3267 | 32.7758 | 44.2369 | 4.961% | 9.921% | 14.88% | 36.5811 | 40.4317
Se, ppm | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sm, ppm | 2.7077 | 0.17112 | 2.3654 | 3.0499 | 2.1943 | 3.221 | 6.32% | 12.64% | 18.96% | 2.5723 | 2.8431
Sn, ppm | 1.0822 | 0.13344 | 0.81534 | 1.3491 | 0.6819 | 1.4825 | 12.33% | 24.66% | 36.99% | 1.0281 | 1.1363
Sr, ppm | 95.1012 | 4.8512 | 85.3988 | 104.8037 | 80.5476 | 109.6549 | 5.101% | 10.2% | 15.3% | 90.3462 | 99.8563
Ta, ppm | 0.24791 | 0.028555 | 0.1908 | 0.30502 | 0.16224 | 0.33357 | 11.52% | 23.04% | 34.55% | 0.23551 | 0.2603
Tb, ppm | 0.61604 | 0.035636 | 0.54476 | 0.68731 | 0.50913 | 0.72294 | 5.785% | 11.57% | 17.35% | 0.58523 | 0.64684
Te, ppm | 0.11278 | 0.017528 | 0.077722 | 0.14783 | 0.060194 | 0.16536 | 15.54% | 31.08% | 46.63% | 0.10714 | 0.11842
Th, ppm | 1.1001 | 0.11373 | 0.8726 | 1.3275 | 0.75888 | 1.4412 | 10.34% | 20.68% | 31.02% | 1.0451 | 1.1551
Ti, wt.% | 0.63857 | 0.026485 | 0.5856 | 0.69154 | 0.55912 | 0.71803 | 4.148% | 8.295% | 12.44% | 0.60664 | 0.6705
Tl, ppm | 0.24219 | 0.017785 | 0.20662 | 0.27776 | 0.18884 | 0.29555 | 7.343% | 14.69% | 22.03% | 0.23008 | 0.2543
Tm, ppm | 0.35807 | 0.021597 | 0.31488 | 0.40127 | 0.29328 | 0.42287 | 6.032% | 12.06% | 18.09% | 0.34017 | 0.37598
U, ppm | 0.3461 | 0.043416 | 0.25926 | 0.43293 | 0.21585 | 0.47634 | 12.54% | 25.09% | 37.63% | 0.32879 | 0.3634
V, ppm | 279.1667 | 15.3471 | 248.4725 | 309.861 | 233.1254 | 325.2081 | 5.497% | 10.99% | 16.49% | 265.2084 | 293.1251
W, ppm | 30.5099 | 1.4347 | 27.6404 | 33.3794 | 26.2057 | 34.8141 | 4.703% | 9.405% | 14.11% | 28.9844 | 32.0354
Y, ppm | 21.654 | 0.84878 | 19.9564 | 23.3515 | 19.1076 | 24.2003 | 3.92% | 7.84% | 11.76% | 20.5713 | 22.7367
Yb, ppm | 2.3228 | 0.1518 | 2.0192 | 2.6264 | 1.8674 | 2.7782 | 6.535% | 13.07% | 19.61% | 2.2067 | 2.439
Zn, ppm | 143.6533 | 6.4896 | 130.6742 | 156.6324 | 124.1846 | 163.122 | 4.518% | 9.035% | 13.55% | 136.4706 | 150.836
Zr, ppm | 54.7689 | 3.9231 | 46.9228 | 62.615 | 42.9997 | 66.5381 | 7.163% | 14.33% | 21.49% | 52.0305 | 57.5074
Aqua Regia Digestion
Ag, ppm | 0.48763 | 0.02905 | 0.42953 | 0.54573 | 0.40048 | 0.57478 | 5.957% | 11.91% | 17.87% | 0.46325 | 0.51201
Al, wt.% | 3.155 | 0.16672 | 2.8216 | 3.4884 | 2.6549 | 3.6552 | 5.284% | 10.57% | 15.85% | 2.9973 | 3.3128
As, ppm | 66.6126 | 3.4001 | 59.8124 | 73.4128 | 56.4123 | 76.8129 | 5.104% | 10.21% | 15.31% | 63.282 | 69.9432
B, ppm | 90.2605 | 6.8946 | 76.4712 | 104.0498 | 69.5766 | 110.9444 | 7.639% | 15.28% | 22.92% | 85.7475 | 94.7735
Ba, ppm | 35.6851 | 2.9416 | 29.8019 | 41.5683 | 26.8602 | 44.5099 | 8.243% | 16.49% | 24.73% | 33.9008 | 37.4694
Be, ppm | 0.26335 | 0.020362 | 0.22262 | 0.30407 | 0.20226 | 0.32443 | 7.732% | 15.46% | 23.2% | 0.25018 | 0.27652
Bi, ppm | 0.061833 | 0.0058736 | 0.050086 | 0.07358 | 0.044213 | 0.079454 | 9.499% | 19% | 28.5% | 0.058742 | 0.064925
Ca, wt.% | 2.6373 | 0.20918 | 2.2189 | 3.0557 | 2.0098 | 3.2649 | 7.932% | 15.86% | 23.79% | 2.5054 | 2.7692
Cd, ppm | 0.61448 | 0.030922 | 0.55263 | 0.67632 | 0.52171 | 0.70724 | 5.032% | 10.06% | 15.1% | 0.58375 | 0.6452
Ce, ppm | 10.6291 | 0.39418 | 9.8408 | 11.4175 | 9.4466 | 11.8117 | 3.709% | 7.417% | 11.13% | 10.0977 | 11.1606
Co, ppm | 30.8067 | 2.2132 | 26.3802 | 35.2332 | 24.1669 | 37.4464 | 7.184% | 14.37% | 21.55% | 29.2664 | 32.347
Cr, ppm | 32.8153 | 1.8121 | 29.191 | 36.4396 | 27.3789 | 38.2517 | 5.522% | 11.04% | 16.57% | 31.1745 | 34.456
Cs, ppm | 0.72621 | 0.023339 | 0.67953 | 0.77289 | 0.65619 | 0.79623 | 3.214% | 6.428% | 9.642% | 0.6899 | 0.76252
Cu, ppm | 169.4907 | 6.2056 | 157.0796 | 181.9019 | 150.874 | 188.1075 | 3.661% | 7.323% | 10.98% | 161.0162 | 177.9653
Dy, ppm | 2.5829 | 0.20394 | 2.175 | 2.9908 | 1.9711 | 3.1947 | 7.896% | 15.79% | 23.69% | 2.4538 | 2.7121
Er, ppm | 1.5241 | 0.16365 | 1.1968 | 1.8514 | 1.0331 | 2.015 | 10.74% | 21.48% | 32.21% | 1.4479 | 1.6003
Eu, ppm | 0.57988 | 0.079993 | 0.41989 | 0.73986 | 0.3399 | 0.81986 | 13.79% | 27.59% | 41.38% | 0.55088 | 0.60887
Fe, wt.% | 5.951 | 0.21659 | 5.5178 | 6.3842 | 5.3012 | 6.6008 | 3.64% | 7.279% | 10.92% | 5.6535 | 6.2486
Ga, ppm | 11.6266 | 0.68704 | 10.2525 | 13.0007 | 9.5655 | 13.6877 | 5.909% | 11.82% | 17.73% | 11.0453 | 12.2079
Gd, ppm | 2.2411 | 0.14236 | 1.9564 | 2.5258 | 1.814 | 2.6682 | 6.352% | 12.7% | 19.06% | 2.129 | 2.3532
Ge, ppm | 0.14633 | 0.017905 | 0.11052 | 0.18214 | 0.09262 | 0.20005 | 12.24% | 24.47% | 36.71% | 0.13902 | 0.15365
Hf, ppm | 0.54117 | 0.045337 | 0.45049 | 0.63184 | 0.40516 | 0.67718 | 8.378% | 16.76% | 25.13% | 0.51411 | 0.56822
Hg, ppm | 0.046422 | 0.012361 | 0.0217 | 0.071144 | 0.0093395 | 0.083505 | 26.63% | 53.25% | 79.88% | 0.044101 | 0.048743
Ho, ppm | 0.52543 | 0.060868 | 0.4037 | 0.64717 | 0.34283 | 0.70804 | 11.58% | 23.17% | 34.75% | 0.49916 | 0.55171
In, ppm | 0.038569 | 0.003268 | 0.032033 | 0.045105 | 0.028766 | 0.048373 | 8.473% | 16.95% | 25.42% | 0.036641 | 0.040498
K, wt.% | 0.13424 | 0.0084762 | 0.11728 | 0.15119 | 0.10881 | 0.15966 | 6.314% | 12.63% | 18.94% | 0.12752 | 0.14095
La, ppm | 4.7234 | 0.2622 | 4.199 | 5.2478 | 3.9368 | 5.51 | 5.551% | 11.1% | 16.65% | 4.4872 | 4.9596
Li, ppm | 9.2802 | 0.43989 | 8.4004 | 10.16 | 7.9605 | 10.5999 | 4.74% | 9.48% | 14.22% | 8.8162 | 9.7442
Lu, ppm | 0.18221 | 0.016938 | 0.14833 | 0.21609 | 0.1314 | 0.23302 | 9.296% | 18.59% | 27.89% | 0.1731 | 0.19132
Mg, wt.% | 1.7239 | 0.048805 | 1.6263 | 1.8215 | 1.5774 | 1.8703 | 2.831% | 5.662% | 8.493% | 1.6377 | 1.8101
Mn, wt.% | 0.071399 | 0.00401 | 0.063379 | 0.079419 | 0.059369 | 0.083429 | 5.616% | 11.23% | 16.85% | 0.067829 | 0.074969
Mo, ppm | 1.4887 | 0.071404 | 1.3459 | 1.6315 | 1.2745 | 1.7029 | 4.796% | 9.593% | 14.39% | 1.4143 | 1.5632
Na, wt.% | 0.26515 | 0.016892 | 0.23136 | 0.29893 | 0.21447 | 0.31583 | 6.371% | 12.74% | 19.11% | 0.25189 | 0.27841
Nb, ppm | 0.15752 | 0.046074 | 0.06537 | 0.24967 | 0.019295 | 0.29574 | 29.25% | 58.5% | 87.75% | 0.14964 | 0.16539
Nd, ppm | 6.8859 | 0.32649 | 6.2329 | 7.5388 | 5.9064 | 7.8653 | 4.741% | 9.483% | 14.22% | 6.5416 | 7.2302
Ni, ppm | 48.8399 | 3.2853 | 42.2693 | 55.4105 | 38.984 | 58.6958 | 6.727% | 13.45% | 20.18% | 46.3979 | 51.2819
P, wt.% | 0.04488 | 0.0014648 | 0.04195 | 0.047809 | 0.040485 | 0.049274 | 3.264% | 6.528% | 9.791% | 0.042636 | 0.047124
Pb, ppm | 30.4578 | 2.062 | 26.3338 | 34.5818 | 24.2719 | 36.6438 | 6.77% | 13.54% | 20.31% | 28.9349 | 31.9807
Pr, ppm | 1.4299 | 0.12984 | 1.1702 | 1.6896 | 1.0404 | 1.8194 | 9.08% | 18.16% | 27.24% | 1.3584 | 1.5014
Rb, ppm | 5.8727 | 0.42772 | 5.0173 | 6.7282 | 4.5896 | 7.1559 | 7.283% | 14.57% | 21.85% | 5.5791 | 6.1664
Re, ppm | 0.0021333 | 0.00039408 | 0.0013452 | 0.0029215 | 0.0009511 | 0.0033156 | 18.47% | 36.94% | 55.42% | 0.0020267 | 0.00224
S, wt.% | 0.44514 | 0.025234 | 0.39467 | 0.49561 | 0.36944 | 0.52085 | 5.669% | 11.34% | 17.01% | 0.42289 | 0.4674
Sb, ppm | 0.89803 | 0.31107 | 0.27589 | 1.5202 | 0 | 1.8312 | 34.64% | 69.28% | 103.9% | 0.85312 | 0.94293
Sc, ppm | 6.472 | 0.66303 | 5.146 | 7.7981 | 4.4829 | 8.4611 | 10.24% | 20.49% | 30.73% | 6.1484 | 6.7956
Se, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sm, ppm | 1.8878 | 0.18395 | 1.5199 | 2.2557 | 1.3359 | 2.4396 | 9.744% | 19.49% | 29.23% | 1.7934 | 1.9822
Sn, ppm | 0.68423 | 0.049347 | 0.58554 | 0.78292 | 0.53619 | 0.83227 | 7.212% | 14.42% | 21.64% | 0.65002 | 0.71844
Sr, ppm | 29.7497 | 2.6716 | 24.4066 | 35.0929 | 21.735 | 37.7644 | 8.98% | 17.96% | 26.94% | 28.2622 | 31.2372
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.38707 | 0.02084 | 0.34539 | 0.42875 | 0.32455 | 0.44959 | 5.384% | 10.77% | 16.15% | 0.36772 | 0.40642
Te, ppm | 0.10545 | 0.013332 | 0.07879 | 0.13212 | 0.065457 | 0.14545 | 12.64% | 25.29% | 37.93% | 0.10018 | 0.11073
Th, ppm | 0.87875 | 0.057138 | 0.76448 | 0.99303 | 0.70734 | 1.0502 | 6.502% | 13% | 19.51% | 0.83482 | 0.92269
Ti, wt.% | 0.38698 | 0.060976 | 0.26502 | 0.50893 | 0.20405 | 0.5699 | 15.76% | 31.51% | 47.27% | 0.36763 | 0.40632
Tl, ppm | 0.12677 | 0.0091055 | 0.10856 | 0.14498 | 0.099454 | 0.15409 | 7.183% | 14.37% | 21.55% | 0.12043 | 0.13311
Tm, ppm | 0.19545 | 0.021004 | 0.15344 | 0.23745 | 0.13243 | 0.25846 | 10.75% | 21.49% | 32.24% | 0.18567 | 0.20522
U, ppm | 0.24262 | 0.014149 | 0.21432 | 0.27092 | 0.20017 | 0.28507 | 5.832% | 11.66% | 17.5% | 0.23049 | 0.25475
V, ppm | 158.2314 | 10.497 | 137.2373 | 179.2254 | 126.7403 | 189.7224 | 6.634% | 13.27% | 19.9% | 150.3198 | 166.1429
W, ppm | 22.6049 | 2.9498 | 16.7053 | 28.5045 | 13.7555 | 31.4543 | 13.05% | 26.1% | 39.15% | 21.4746 | 23.7351
Y, ppm | 12.9081 | 0.84393 | 11.2203 | 14.596 | 10.3763 | 15.4399 | 6.538% | 13.08% | 19.61% | 12.2627 | 13.5535
Yb, ppm | 1.2904 | 0.13127 | 1.0278 | 1.5529 | 0.89657 | 1.6842 | 10.17% | 20.35% | 30.52% | 1.2259 | 1.3549
Zn, ppm | 132.8469 | 7.3422 | 118.1625 | 147.5312 | 110.8204 | 154.8734 | 5.527% | 11.05% | 16.58% | 126.2045 | 139.4892
Zr, ppm | 17.8391 | 2.0927 | 13.6537 | 22.0245 | 11.561 | 24.1172 | 11.73% | 23.46% | 35.19% | 16.9471 | 18.731
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 236
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.8485 | 1.8256 | 1.8714 | 1.8392 | 1.8577
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.7515 | 1.7118 | 1.7913 | 1.7419 | 1.7611
Cyanide Leach
Au, Gold (ppm) | 1.7929 | 1.7703 | 1.8155 | 1.7894 | 1.7964
PhotonAssay
Au, Gold (ppm) | 1.8475 | 1.8289 | 1.866 | 1.8448 | 1.8502
4-Acid Digestion
Ag, Silver (ppm) | 0.47814 | 0.4326 | 0.52368 | 0.45207 | 0.50421
Al, Aluminium (wt.%) | 6.5932 | 6.4499 | 6.7364 | 6.4938 | 6.6925
As, Arsenic (ppm) | 67.6896 | 64.8166 | 70.5627 | 65.9474 | 69.4318
Ba, Barium (ppm) | 245.7522 | 238.8185 | 252.6859 | 241.58 | 249.9244
Be, Beryllium (ppm) | 0.45857 | 0.42785 | 0.48929 | 0.41675 | 0.50039
Bi, Bismuth (ppm) | 0.061389 | 0.05299 | 0.069788 | IND | IND
Ca, Calcium (wt.%) | 5.9341 | 5.7827 | 6.0855 | 5.8406 | 6.0276
Cd, Cadmium (ppm) | 0.62376 | 0.57955 | 0.66797 | 0.5938 | 0.65372
Ce, Cerium (ppm) | 13.9587 | 13.0941 | 14.8233 | 13.4873 | 14.4301
Co, Cobalt (ppm) | 41.4883 | 40.1524 | 42.8242 | 40.5613 | 42.4153
Cr, Chromium (ppm) | 86.4016 | 82.7263 | 90.0769 | 83.6179 | 89.1853
Cs, Caesium (ppm) | 1.0081 | 0.93022 | 1.0859 | 0.96312 | 1.053
Cu, Copper (ppm) | 169.6544 | 165.0799 | 174.2288 | 166.1231 | 173.1857
Dy, Dysprosium (ppm) | 3.938 | 3.6995 | 4.1764 | 3.7817 | 4.0942
Er, Erbium (ppm) | 2.4286 | 2.2868 | 2.5704 | 2.3518 | 2.5054
Eu, Europium (ppm) | 0.96342 | 0.91349 | 1.0134 | 0.92676 | 1.0001
Fe, Iron (wt.%) | 8.0044 | 7.7917 | 8.2172 | 7.8738 | 8.1351
Ga, Gallium (ppm) | 16.1386 | 15.2113 | 17.066 | 15.7349 | 16.5423
Gd, Gadolinium (ppm) | 3.4747 | 3.2319 | 3.7175 | 3.3337 | 3.6157
Hf, Hafnium (ppm) | 1.8009 | 1.6826 | 1.9193 | 1.7032 | 1.8987
Ho, Holmium (ppm) | 0.85917 | 0.80736 | 0.91097 | 0.83423 | 0.8841
In, Indium (ppm) | 0.079708 | 0.07148 | 0.087937 | 0.0734 | 0.086016
K, Potassium (wt.%) | 0.55896 | 0.54026 | 0.57767 | 0.55006 | 0.56786
La, Lanthanum (ppm) | 5.9884 | 5.6718 | 6.3049 | 5.8139 | 6.1628
Li, Lithium (ppm) | 10.5106 | 10.0801 | 10.9411 | 10.1821 | 10.839
Lu, Lutetium (ppm) | 0.37163 | 0.33924 | 0.40403 | 0.35125 | 0.39202
Mg, Magnesium (wt.%) | 3.3955 | 3.3151 | 3.476 | 3.3491 | 3.442
Mn, Manganese (wt.%) | 0.13325 | 0.12883 | 0.13768 | 0.13106 | 0.13545
Mo, Molybdenum (ppm) | 1.5601 | 1.4267 | 1.6936 | 1.4842 | 1.6361
Na, Sodium (wt.%) | 2.155 | 2.0939 | 2.216 | 2.1157 | 2.1942
Nb, Niobium (ppm) | 3.6281 | 3.3968 | 3.8594 | 3.4829 | 3.7733
Nd, Neodymium (ppm) | 9.0119 | 8.5551 | 9.4688 | 8.8196 | 9.2043
Ni, Nickel (ppm) | 68.4509 | 65.8669 | 71.0349 | 66.8284 | 70.0735
P, Phosphorus (wt.%) | 0.045336 | 0.043806 | 0.046865 | 0.044275 | 0.046397
Pb, Lead (ppm) | 30.4861 | 29.1292 | 31.843 | 29.5789 | 31.3934
Pr, Praseodymium (ppm) | 1.9404 | 1.7947 | 2.0861 | 1.8806 | 2.0002
Rb, Rubidium (ppm) | 14.0762 | 13.432 | 14.7203 | 13.7408 | 14.4116
Re, Rhenium (ppm) | 0.0026667 | 0.0011433 | 0.00419 | IND | IND
S, Sulphur (wt.%) | 0.44412 | 0.43037 | 0.45787 | 0.43392 | 0.45433
Sb, Antimony (ppm) | 1.7628 | 1.6477 | 1.878 | 1.6708 | 1.8549
Sc, Scandium (ppm) | 38.5064 | 37.1998 | 39.8129 | 37.3266 | 39.6861
Se, Selenium (ppm) | < 5 | IND | IND | IND | IND
Sm, Samarium (ppm) | 2.7077 | 2.4744 | 2.941 | 2.6178 | 2.7975
Sn, Tin (ppm) | 1.0822 | 0.94543 | 1.219 | IND | IND
Sr, Strontium (ppm) | 95.1012 | 92.1398 | 98.0627 | 93.3934 | 96.8091
Ta, Tantalum (ppm) | 0.24791 | 0.22285 | 0.27297 | 0.22885 | 0.26697
Tb, Terbium (ppm) | 0.61604 | 0.574 | 0.65807 | 0.58994 | 0.64213
Te, Tellurium (ppm) | 0.11278 | 0.086992 | 0.13856 | IND | IND
Th, Thorium (ppm) | 1.1001 | 1.0239 | 1.1762 | 1.0525 | 1.1476
Ti, Titanium (wt.%) | 0.63857 | 0.61944 | 0.65771 | 0.62269 | 0.65445
Tl, Thallium (ppm) | 0.24219 | 0.22687 | 0.25752 | 0.22957 | 0.25482
Tm, Thulium (ppm) | 0.35807 | 0.33042 | 0.38573 | 0.34398 | 0.37217
U, Uranium (ppm) | 0.3461 | 0.3084 | 0.38379 | 0.32177 | 0.37042
V, Vanadium (ppm) | 279.1667 | 268.7538 | 289.5797 | 270.8319 | 287.5016
W, Tungsten (ppm) | 30.5099 | 29.3577 | 31.662 | 29.4654 | 31.5543
Y, Yttrium (ppm) | 21.654 | 20.9559 | 22.352 | 21.2648 | 22.0431
Yb, Ytterbium (ppm) | 2.3228 | 2.1785 | 2.4671 | 2.2757 | 2.37
Zn, Zinc (ppm) | 143.6533 | 139.4721 | 147.8345 | 140.9873 | 146.3193
Zr, Zirconium (ppm) | 54.7689 | 52.1709 | 57.3669 | 52.9224 | 56.6154
Aqua Regia Digestion
Ag, Silver (ppm) | 0.48763 | 0.46021 | 0.51505 | 0.46749 | 0.50778
Al, Aluminium (wt.%) | 3.155 | 3.0411 | 3.2689 | 3.0916 | 3.2184
As, Arsenic (ppm) | 66.6126 | 64.3099 | 68.9153 | 65.1942 | 68.0311
B, Boron (ppm) | 90.2605 | 83.6966 | 96.8244 | 88.2216 | 92.2993
Ba, Barium (ppm) | 35.6851 | 33.8248 | 37.5454 | 34.672 | 36.6982
Be, Beryllium (ppm) | 0.26335 | 0.23855 | 0.28815 | 0.23285 | 0.29384
Bi, Bismuth (ppm) | 0.061833 | 0.054902 | 0.068765 | IND | IND
Ca, Calcium (wt.%) | 2.6373 | 2.5197 | 2.7549 | 2.5723 | 2.7023
Cd, Cadmium (ppm) | 0.61448 | 0.57842 | 0.65054 | 0.57834 | 0.65061
Ce, Cerium (ppm) | 10.6291 | 10.2538 | 11.0045 | 10.3617 | 10.8966
Co, Cobalt (ppm) | 30.8067 | 29.4977 | 32.1157 | 29.9414 | 31.672
Cr, Chromium (ppm) | 32.8153 | 31.5121 | 34.1184 | 31.9937 | 33.6369
Cs, Caesium (ppm) | 0.72621 | 0.69891 | 0.7535 | 0.70866 | 0.74375
Cu, Copper (ppm) | 169.4907 | 165.0542 | 173.9272 | 165.9272 | 173.0543
Dy, Dysprosium (ppm) | 2.5829 | 2.304 | 2.8618 | 2.4471 | 2.7187
Er, Erbium (ppm) | 1.5241 | 1.2889 | 1.7592 | 1.3092 | 1.7389
Eu, Europium (ppm) | 0.57988 | 0.46662 | 0.69313 | 0.53886 | 0.62089
Fe, Iron (wt.%) | 5.951 | 5.7993 | 6.1027 | 5.8606 | 6.0414
Ga, Gallium (ppm) | 11.6266 | 11.1472 | 12.106 | 11.2329 | 12.0203
Gd, Gadolinium (ppm) | 2.2411 | 2.0171 | 2.4651 | 2.124 | 2.3582
Ge, Germanium (ppm) | 0.14633 | 0.12265 | 0.17002 | IND | IND
Hf, Hafnium (ppm) | 0.54117 | 0.49805 | 0.58428 | 0.51763 | 0.5647
Hg, Mercury (ppm) | 0.046422 | 0.03416 | 0.058685 | 0.026825 | 0.06602
Ho, Holmium (ppm) | 0.52543 | 0.43752 | 0.61335 | IND | IND
In, Indium (ppm) | 0.038569 | 0.032274 | 0.044865 | 0.034082 | 0.043057
K, Potassium (wt.%) | 0.13424 | 0.12794 | 0.14053 | 0.12983 | 0.13865
La, Lanthanum (ppm) | 4.7234 | 4.5316 | 4.9152 | 4.5945 | 4.8523
Li, Lithium (ppm) | 9.2802 | 8.9258 | 9.6346 | 9.0281 | 9.5324
Lu, Lutetium (ppm) | 0.18221 | 0.16347 | 0.20095 | IND | IND
Mg, Magnesium (wt.%) | 1.7239 | 1.6866 | 1.7611 | 1.6975 | 1.7502
Mn, Manganese (wt.%) | 0.071399 | 0.068935 | 0.073864 | 0.070375 | 0.072424
Mo, Molybdenum (ppm) | 1.4887 | 1.4167 | 1.5607 | 1.4336 | 1.5439
Na, Sodium (wt.%) | 0.26515 | 0.25513 | 0.27517 | 0.25704 | 0.27326
Nb, Niobium (ppm) | 0.15752 | 0.11844 | 0.19659 | IND | IND
Nd, Neodymium (ppm) | 6.8859 | 6.4011 | 7.3706 | 6.6766 | 7.0952
Ni, Nickel (ppm) | 48.8399 | 46.8957 | 50.7841 | 47.8102 | 49.8696
P, Phosphorus (wt.%) | 0.04488 | 0.043487 | 0.046273 | 0.044083 | 0.045677
Pb, Lead (ppm) | 30.4578 | 29.2094 | 31.7062 | 29.2908 | 31.6248
Pr, Praseodymium (ppm) | 1.4299 | 1.2449 | 1.6149 | 1.3234 | 1.5364
Rb, Rubidium (ppm) | 5.8727 | 5.5545 | 6.191 | 5.6547 | 6.0908
Re, Rhenium (ppm) | 0.0021333 | 0.0012724 | 0.0029943 | IND | IND
S, Sulphur (wt.%) | 0.44514 | 0.42867 | 0.46162 | 0.43442 | 0.45587
Sb, Antimony (ppm) | 0.89803 | 0.73544 | 1.0606 | 0.82029 | 0.97577
Sc, Scandium (ppm) | 6.472 | 6.0282 | 6.9159 | 6.2155 | 6.7285
Se, Selenium (ppm) | < 1 | IND | IND | IND | IND
Sm, Samarium (ppm) | 1.8878 | 1.6513 | 2.1242 | 1.765 | 2.0106
Sn, Tin (ppm) | 0.68423 | 0.61953 | 0.74893 | IND | IND
Sr, Strontium (ppm) | 29.7497 | 28.2457 | 31.2537 | 28.9271 | 30.5724
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.38707 | 0.36211 | 0.41203 | 0.36673 | 0.40741
Te, Tellurium (ppm) | 0.10545 | 0.088907 | 0.122 | IND | IND
Th, Thorium (ppm) | 0.87875 | 0.81147 | 0.94603 | 0.83328 | 0.92423
Ti, Titanium (wt.%) | 0.38698 | 0.3557 | 0.41825 | 0.37499 | 0.39896
Tl, Thallium (ppm) | 0.12677 | 0.11823 | 0.13531 | IND | IND
Tm, Thulium (ppm) | 0.19545 | 0.16649 | 0.2244 | IND | IND
U, Uranium (ppm) | 0.24262 | 0.23033 | 0.2549 | 0.23078 | 0.25446
V, Vanadium (ppm) | 158.2314 | 151.5344 | 164.9283 | 154.1627 | 162.3
W, Tungsten (ppm) | 22.6049 | 20.573 | 24.6368 | 21.8352 | 23.3746
Y, Yttrium (ppm) | 12.9081 | 12.366 | 13.4502 | 12.5188 | 13.2975
Yb, Ytterbium (ppm) | 1.2904 | 1.132 | 1.4488 | 1.1709 | 1.4099
Zn, Zinc (ppm) | 132.8469 | 128.6489 | 137.0449 | 129.99 | 135.7037
Zr, Zirconium (ppm) | 17.8391 | 16.4569 | 19.2213 | 17.0355 | 18.6426
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 236
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.12259 | Hg | ppm | < 2
Aqua Regia Digestion
Pd | ppb | < 10 | Pt | ppb | 9.8
Borate Fusion XRF
Al2O3 | wt.% | 12.665 | Fe2O3 | wt.% | 11.81 | S | wt.% | 0.4425
As | ppm | 70 | K2O | wt.% | 0.6685 | SiO2 | wt.% | 52.785
BaO | ppm | 285 | MgO | wt.% | 5.85 | Sn | ppm | 7.5
CaO | wt.% | 8.555 | MnO | wt.% | 0.1765 | Sr | ppm | 126.8392
Cl | ppm | 2705 | Na2O | wt.% | 2.895 | TiO2 | wt.% | 1.102
Co | ppm | 50 | Ni | ppm | 75 | V2O5 | ppm | 570
Cr2O3 | ppm | 145 | P2O5 | wt.% | 0.1095 | Zn | ppm | 155
Cu | ppm | 185 | Pb | ppm | 35 | Zr | ppm | 77.7338
Thermogravimetry
LOI1000 | wt.% | 3.27
Infrared Combustion
C | wt.% | 0.195 | S | wt.% | 0.38
Laser Ablation ICP-MS
Ag | ppm | 0.5 | Hf | ppm | 2.105 | Sm | ppm | 2.785
As | ppm | 66.5 | Ho | ppm | 0.92 | Sn | ppm | 1.1
Ba | ppm | 258 | In | ppm | 0.075 | Sr | ppm | 95.2
Be | ppm | 0.35 | La | ppm | 6.12 | Ta | ppm | 0.27
Bi | ppm | 0.06 | Lu | ppm | 0.37 | Tb | ppm | 0.63
Cd | ppm | 0.7 | Mn | wt.% | 0.14 | Te | ppm | < 0.2
Ce | ppm | 13.65 | Mo | ppm | 1.6 | Th | ppm | 1.12
Co | ppm | 44.4 | Nb | ppm | 3.795 | Ti | wt.% | 0.659
Cr | ppm | 102.5 | Nd | ppm | 9.535 | Tl | ppm | < 0.2
Cs | ppm | 1 | Ni | ppm | 75 | Tm | ppm | 0.375
Cu | ppm | 172 | Pb | ppm | 33 | U | ppm | 0.375
Dy | ppm | 4.075 | Pr | ppm | 2.07 | V | ppm | 299.5
Er | ppm | 2.595 | Rb | ppm | 14.1 | W | ppm | 32.75
Eu | ppm | 0.99 | Re | ppm | < 0.01 | Y | ppm | 22.95
Ga | ppm | 15.8 | Sb | ppm | 1.85 | Yb | ppm | 2.53
Gd | ppm | 3.52 | Sc | ppm | 39.5 | Zn | ppm | 152.5
Ge | ppm | 1.425 | Se | ppm | < 5 | Zr | ppm | 72
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 236
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 236
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Ostrea Mineral Laboratories, Inc., Laguna, Philippines
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 236 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 13 | 13 | 207034 | 0.086901 | 1.8534 | 1.9011
2 | 1 | ANSLu | 1 | 15 | 15 | 207035 | 0.087444 | 1.9042 | 1.9038
2 | 1 | ANSLu | 1 | 16 | 16 | 207036 | 0.085415 | 1.9431 | 1.9059
2 | 1 | ANSLu | 1 | 3 | 3 | 207037 | 0.083299 | 1.9515 | 1.9063
2 | 1 | ANSLu | 1 | 9 | 9 | 207038 | 0.08511 | 1.8219 | 1.8994
2 | 1 | ANSLu | 1 | 14 | 14 | 207039 | 0.083589 | 1.7816 | 1.8972
2 | 1 | ANSLu | 1 | 4 | 4 | 207040 | 0.088186 | 1.9368 | 1.9056
2 | 1 | ANSLu | 1 | 17 | 17 | 207041 | 0.087484 | 1.8594 | 1.9014
2 | 1 | ANSLu | 1 | 11 | 11 | 207042 | 0.087389 | 1.9984 | 1.9088
2 | 1 | ANSLu | 1 | 1 | 1 | 207043 | 0.083758 | 1.9529 | 1.9064
2 | 1 | ANSLu | 1 | 6 | 6 | 207044 | 0.087695 | 1.9052 | 1.9039
2 | 1 | ANSLu | 1 | 2 | 2 | 207045 | 0.082173 | 1.93 | 1.9052
2 | 1 | ANSLu | 1 | 12 | 12 | 207046 | 0.087432 | 1.8936 | 1.9032
2 | 1 | ANSLu | 1 | 8 | 8 | 207047 | 0.084447 | 1.9875 | 1.9083
2 | 1 | ANSLu | 1 | 18 | 18 | 207048 | 0.086499 | 1.9064 | 1.9039
2 | 1 | ANSLu | 1 | 7 | 7 | 207049 | 0.086286 | 1.8844 | 1.9027
2 | 1 | ANSLu | 1 | 19 | 19 | 207050 | 0.087078 | 1.9687 | 1.9073
2 | 1 | ANSLu | 1 | 10 | 10 | 207051 | 0.085123 | 1.8578 | 1.9013
2 | 1 | ANSLu | 1 | 5 | 5 | 207052 | 0.084554 | 1.8518 | 1.901
2 | 1 | ANSLu | 1 | 20 | 20 | 207053 | 0.087196 | 1.8872 | 1.9029
Mean | [6 empty] | 0.085853 | 1.9038 | 1.9038
Median | [7 empty] | 1.9047 | 1.9038
SD | [7 empty] | 0.05615 | 0.0030038
RSD | [7 empty] | 2.949% | 0.1578%
RSD Upscale Factor | [7 empty] | 0.053495
Parameters | :
Nominal Mass Value |  | 30 | gm
Note: Upscaled Metrics are also available in this CSG for analytes with INAA results for Au

## Fire Assay
| Analytical results for Au in OREAS 236 (Certified Value 1.85 ppm) | [28 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | 2
|  |  | 0.085g | 30g | 40g | 40g | 30g | 30g | 30g | 30g | 25g | 50g | 25g | 30g | 25g | 25g | 30g | 50g | 30g | 50g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 1.9529 | 1.85 | 1.9 | 1.92 | 1.771 | 1.85 | 1.8 | 1.845 | 1.928 | 1.813 | 1.88 | 1.87 | 1.802 | 1.84 | 1.76 | 1.834 | 1.75 | 1.79 | 1.92 | 1.89 | 1.718 | 1.89 | 1.866 | 1.93 | 1.855 | 1.888 | 1.785 | 1
| 1 | 2 | 1.93 | 1.885 | 1.88 | 1.91 | 1.756 | 1.84 | 1.8 | 1.786 | 1.922 | 1.7906 | 1.92 | 1.87 | 1.709 | 1.855 | 1.77 | 1.818 | 1.77 | 1.8 | 1.945 | 1.88 | 1.798 | 1.86 | 1.897 | 1.93 | 1.845 | 1.804 | 1.845 | #N/A
| 1 | 3 | 1.9515 | 1.87 | 1.95 | 1.91 | 1.82 | 1.87 | 1.81 | 1.775 | 1.955 | 1.81 | 1.91 | 1.89 | 1.776 | 1.824 | 1.75 | 1.851 | 1.81 | 1.8 | 1.89 | 1.89 | 1.792 | 1.9 | 1.866 | 1.97 | 1.86 | 1.847 | 1.8 | 16
| 1 | 4 | 1.9368 | 1.87 | 1.96 | 1.9 | 1.801 | 1.83 | 1.8 | 1.762 | 1.952 | 1.85 | 1.93 | 1.88 | 1.754 | 1.773 | 1.8 | 1.873 | 1.78 | 1.8 | 1.935 | 1.83 | 1.757 | 1.87 | 1.865 | 1.97 | 1.885 | 1.88 | 1.845 | 1.8485
| 1 | 5 | 1.8518 | 1.845 | 1.96 | 1.93 | 1.798 | 1.9 | 1.82 | 1.823 | 1.949 | 1.85 | 1.88 | 1.89 | 1.763 | 1.871 | 1.74 | 1.852 | 1.83 | 1.76 | 1.945 | 1.87 | 1.767 | 1.81 | 1.885 | 1.87 | 1.86 | 1.806 | 1.795 | 7
| 1 | 6 | 1.9052 | 1.88 | 1.87 | 1.93 | 1.806 | 1.86 | 1.77 | 1.778 | 1.943 | 1.87 | 1.94 | 1.88 | 1.752 | 1.864 | 1.77 | 1.833 | 1.86 | 1.64 | 1.93 | 1.88 | 1.79 | 1.83 | 1.852 | 1.89 | 1.87 | 1.873 | 1.8
|  | 7 | 1.8844
|  | 8 | 1.9875
|  | 9 | 1.8219
|  | 10 | 1.8578
|  | 11 | 1.9984
|  | 12 | 1.8936
|  | 13 | 1.8534
|  | 14 | 1.7816
|  | 15 | 1.9042
|  | 16 | 1.9431
|  | 17 | 1.8594
|  | 18 | 1.9064
|  | 19 | 1.9687
|  | 20 | 1.8872
| Mean |  | 1.9038 | 1.8667 | 1.92 | 1.9167 | 1.792 | 1.8583 | 1.8 | 1.7948 | 1.9415 | 1.8306 | 1.91 | 1.88 | 1.7593 | 1.8378 | 1.765 | 1.8435 | 1.8 | 1.765 | 1.9275 | 1.8733 | 1.7703 | 1.86 | 1.8718 | 1.9267 | 1.8625 | 1.8497 | 1.8117
| Median |  | 1.9047 | 1.87 | 1.925 | 1.915 | 1.7995 | 1.855 | 1.8 | 1.782 | 1.946 | 1.8315 | 1.915 | 1.88 | 1.7585 | 1.8475 | 1.765 | 1.8425 | 1.795 | 1.795 | 1.9325 | 1.88 | 1.7785 | 1.865 | 1.866 | 1.93 | 1.86 | 1.86 | 1.8
| Std Dev. |  | 0.05615 | 0.016021 | 0.041473 | 0.012111 | 0.023808 | 0.024833 | 0.016733 | 0.032071 | 0.013517 | 0.030471 | 0.025298 | 0.0089443 | 0.030748 | 0.035997 | 0.020736 | 0.019233 | 0.040988 | 0.063166 | 0.020676 | 0.022509 | 0.030164 | 0.034641 | 0.016216 | 0.040825 | 0.013693 | 0.037238 | 0.026394
| Rel.Std.Dev. |  | 2.949% | 0.8583% | 2.16% | 0.6319% | 1.329% | 1.336% | 0.9296% | 1.787% | 0.6962% | 1.665% | 1.325% | 0.4758% | 1.748% | 1.959% | 1.175% | 1.043% | 2.277% | 3.579% | 1.073% | 1.202% | 1.704% | 1.862% | 0.8663% | 2.119% | 0.7352% | 2.013% | 1.457%
| PDM3 |  | 2.993% | 0.9854% | 3.871% | 3.69% | -3.054% | 0.5346% | -2.621% | -2.901% | 5.034% | -0.9662% | 3.33% | 1.707% | -4.821% | -0.5745% | -4.515% | -0.2679% | -2.621% | -4.515% | 4.276% | 1.346% | -4.226% | 0.6247% | 1.265% | 4.231% | 0.76% | 0.0657% | -1.99%
| Z-Score (Absolute) |  | NA | 0.16 | 0.82 | 0.78 | 0.77 | 0.05 | 0.67 | 0.74 | 1.09 | 0.29 | 0.7 | 0.32 | 1.18 | 0.2 | 1.11 | 0.13 | 0.67 | 1.11 | 0.92 | 0.24 | 1.04 | 0.07 | 0.22 | 0.91 | 0.11 | 0.05 | 0.53
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 236 (Certified Value 1.75 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | ppm
|  |  | INAA | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*OES | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*MS | 2
|  |  | 0.085g | 25g | 40g | 30g | 25g | 10g | 25g | 50g | 25g | 50g | 25g | 30g | 30g | 25g | 50g | 25g | 30g | 15g | 50g | 15g | 25g | 20g | 25g | 3
| 1 | 1 | 1.9529 | 1.69 | 1.8 | 1.639 | 1.41 | 1.78 | 1.71 | 1.71 | 1.803 | 1.714 | 1.762 | 1.69 | 1.869 | > 0.5 | 1.887 | 1.73 | 1.6 | 1.9371 | 1.82 | 1.798 | 1.85 | 1.7835 | 1.65 | 1
| 1 | 2 | 1.93 | 1.67 | 1.83 | 1.636 | 1.47 | 1.64 | 1.77 | 1.71 | 1.775 | 1.724 | 1.709 | 1.81 | 1.8795 | > 0.5 | 1.856 | 1.74 | 1.6 | 1.9904 | 1.82 | 1.808 | 1.72 | 1.7623 | 1.66 | #N/A
| 1 | 3 | 1.9515 | 1.68 | 1.87 | 1.633 | 1.42 | 1.72 | 1.75 | 1.6738 | 1.806 | 1.706 | 1.696 | 1.8 | 1.904 | > 0.5 | 1.872 | 1.75 | 1.6 | 2.0102 | 1.86 | 1.828 | 1.84 | 1.7717 | 1.59 | 16
| 1 | 4 | 1.9368 | 1.7 | 1.87 | 1.614 | 1.48 | 1.74 | 1.7 | 1.8 | 1.828 | 1.723 | 1.801 | 1.75 | 1.9215 | > 0.5 | 1.872 | 1.77 | 1.6 | 2.0227 | 1.82 | 1.868 | 1.9 | 1.694 | 1.59 | 1.7515
| 1 | 5 | 1.8518 | 1.66 | 1.82 | 1.619 | 1.41 | 1.77 | 1.73 | 1.71 | 1.793 | 1.729 | 1.775 | 1.77 | 1.913 | > 0.5 | 1.904 | 1.75 | 1.6 | 1.9474 | 1.78 | 1.828 | 1.85 | 1.7242 | 1.65 | 9
| 1 | 6 | 1.9052 | 1.69 | 1.8 | 1.628 | 1.4 | 1.6 | 1.75 | 1.83 | 1.773 | 1.731 | 1.696 | 1.74 | 1.8375 | > 0.5 | 1.949 | 1.76 | 1.6 | 2.0064 | 1.75 | 1.798 | 1.67 | 1.7494 | 1.63
|  | 7 | 1.8844
|  | 8 | 1.9875
|  | 9 | 1.8219
|  | 10 | 1.8578
|  | 11 | 1.9984
|  | 12 | 1.8936
|  | 13 | 1.8534
|  | 14 | 1.7816
|  | 15 | 1.9042
|  | 16 | 1.9431
|  | 17 | 1.8594
|  | 18 | 1.9064
|  | 19 | 1.9687
|  | 20 | 1.8872
| Mean |  | 1.9038 | 1.6817 | 1.8317 | 1.6282 | 1.4317 | 1.7083 | 1.735 | 1.739 | 1.7963 | 1.7212 | 1.7398 | 1.76 | 1.8874 |  | 1.89 | 1.75 | 1.6 | 1.9857 | 1.8083 | 1.8213 | 1.805 | 1.7475 | 1.6283
| Median |  | 1.9047 | 1.685 | 1.825 | 1.6305 | 1.415 | 1.73 | 1.74 | 1.71 | 1.798 | 1.7235 | 1.7355 | 1.76 | 1.8918 |  | 1.8795 | 1.75 | 1.6 | 1.9984 | 1.82 | 1.818 | 1.845 | 1.7559 | 1.64
| Std Dev. |  | 0.05615 | 0.01472 | 0.031885 | 0.0098675 | 0.034303 | 0.072778 | 0.026646 | 0.061279 | 0.020743 | 0.0094956 | 0.045305 | 0.043818 | 0.031574 |  | 0.033136 | 0.014142 | 2.4324e-16 | 0.035352 | 0.038166 | 0.026583 | 0.089163 | 0.033195 | 0.031252
| Rel.Std.Dev. |  | 2.949% | 0.8753% | 1.741% | 0.606% | 2.396% | 4.26% | 1.536% | 3.524% | 1.155% | 0.5517% | 2.604% | 2.49% | 1.673% |  | 1.753% | 0.8081% | 1.52e-14% | 1.78% | 2.111% | 1.46% | 4.94% | 1.9% | 1.919%
| PDM3 |  | 8.693% | -3.989% | 4.575% | -7.043% | -18.26% | -2.466% | -0.9437% | -0.7172% | 2.558% | -1.734% | -0.6678% | 0.4836% | 7.758% |  | 7.906% | -0.08734% | -8.651% | 13.37% | 3.243% | 3.985% | 3.053% | -0.2291% | -7.034%
| Z-Score (Absolute) |  | NA | 0.73 | 0.93 | 1.32 | 3.5 | 0.43 | 0.14 | 0.09 | 0.54 | 0.29 | 0.09 | 0.14 | 1.55 | NA | 1.58 | 0.03 | 1.64 | 2.64 | 0.67 | 0.82 | 0.64 | 0 | 1.32
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 236 (Certified Value 1.79 ppm) | [20 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.11 | 1.14 | 1.15 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 30g | 50g | 50g | 200g | 60g | 200g | 50g | 20g | 50g | 200g | 30g | 05g | 30g | 200g | 200g | 05g | 200g | 30g | 3
| 1 | 1 | 1.9529 | 1.82 | 1.95 | 1.84 | 1.4867 | 1.76 | 1.81 | 1.7925 | 1.79 | 1.81 | 1.782 | 1.76 | 1.8 | 1.94 | 1.8892 | 1.8 | 1.7 | 1.76 | 1.71 | 1
| 1 | 2 | 1.93 | 1.86 | 1.99 | 1.86 | 1.39 | 1.78 | 1.84 | 1.84 | 1.76 | 1.79 | 1.769 | 1.72 | 1.8 | 1.8 | 1.8788 | 1.78 | 1.8 | 1.77 | 1.77 | #N/A
| 1 | 3 | 1.9515 | 1.83 | 2.08 | 1.84 | 1.56 | 1.74 | 1.84 | 1.8125 | 1.73 | 1.77 | 1.83 | 1.73 | 1.8 | 1.76 | 1.8658 | 1.83 | 1.8 | 1.77 | 1.77 | 16
| 1 | 4 | 1.9368 | 1.84 | 2.06 | 1.84 | 1.5133 | 1.74 | 1.83 | 1.795 | 1.72 | 1.85 | 1.817 | 1.7 | 1.8 | 1.76 | 1.8723 | 1.79 | 1.8 | 1.77 | 1.7 | 1.7929
| 1 | 5 | 1.8518 | 1.83 | 2.06 | 1.86 | 1.47 | 1.74 | 1.84 | 1.8675 | 1.78 | 1.76 | 1.821 | 1.72 | 1.8 | 1.77 | 1.8654 | 1.76 | 1.8 | 1.76 | 1.77 | 11
| 1 | 6 | 1.9052 | 1.89 | 1.9 | 1.85 | 1.4333 | 1.75 | 1.86 | 1.77 | 1.72 | 1.77 | 1.861 | 1.73 | 1.8 | 1.78 | 1.8552 | 1.75 | 1.7 | 1.77 | 1.76
|  | 7 | 1.8844
|  | 8 | 1.9875
|  | 9 | 1.8219
|  | 10 | 1.8578
|  | 11 | 1.9984
|  | 12 | 1.8936
|  | 13 | 1.8534
|  | 14 | 1.7816
|  | 15 | 1.9042
|  | 16 | 1.9431
|  | 17 | 1.8594
|  | 18 | 1.9064
|  | 19 | 1.9687
|  | 20 | 1.8872
| Mean |  | 1.9038 | 1.845 | 2.0067 | 1.8483 | 1.4756 | 1.7517 | 1.8367 | 1.8129 | 1.75 | 1.7917 | 1.8133 | 1.7267 | 1.8 | 1.8017 | 1.8711 | 1.785 | 1.7667 | 1.7667 | 1.7467
| Median |  | 1.9047 | 1.835 | 2.025 | 1.845 | 1.4783 | 1.745 | 1.84 | 1.8037 | 1.745 | 1.78 | 1.819 | 1.725 | 1.8 | 1.775 | 1.8691 | 1.785 | 1.8 | 1.77 | 1.765
| Std Dev. |  | 0.05615 | 0.025884 | 0.072019 | 0.0098319 | 0.059691 | 0.016021 | 0.01633 | 0.035476 | 0.030984 | 0.033714 | 0.033375 | 0.019664 | 0 | 0.069402 | 0.011862 | 0.02881 | 0.05164 | 0.005164 | 0.03266
| Rel.Std.Dev. |  | 2.949% | 1.403% | 3.589% | 0.5319% | 4.045% | 0.9146% | 0.8891% | 1.957% | 1.771% | 1.882% | 1.841% | 1.139% | 0% | 3.852% | 0.6339% | 1.614% | 2.923% | 0.2923% | 1.87%
| PDM3 |  | 6.185% | 2.906% | 11.92% | 3.092% | -17.7% | -2.3% | 2.441% | 1.116% | -2.393% | -0.06875% | 1.14% | -3.694% | 0.396% | 0.489% | 4.364% | -0.4406% | -1.463% | -1.463% | -2.579%
| Z-Score (Absolute) |  | NA | 0.78 | 3.35 | 0.83 | 5.08 | 0.7 | 0.65 | 0.27 | 0.73 | 0.07 | 0.28 | 1.1 | 0.07 | 0.09 | 1.19 | 0.17 | 0.46 | 0.46 | 0.78
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 236 (Certified Value 1.85 ppm) | [9 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 1.9052 | 1.8462 | 1.8789 | 1.818 | 1.8161 | 1.8313 | 2.0012 | 1.888 | 1
| 1 | 2 | 1.9431 | 1.83 | 1.85 | 1.8642 | 1.821 | 1.7481 | 1.9636 | 1.8296 | #N/A
| 1 | 3 | 1.93 | 1.7931 | 1.8326 | 1.9582 | 1.763 | 1.8679 | 1.8768 | 1.7955 | 16
| 1 | 4 | 1.8518 | 1.9541 | 1.8379 | 1.93 | 1.7347 | 1.8453 | 1.8769 | 1.9169 | 1.8475
| 1 | 5 | 1.9515 | 1.8482 | 1.7893 | 1.843 | 1.816 | 1.7888 | 1.8132 | 1.9115 | 13
| 1 | 6 | 1.8534 | 2.0805 | 1.789 | 1.9007 | 1.7951 | 1.8297 | 1.824 | 1.9248
| 2 | 7 | 1.8936 | 1.8745 | 1.8747 | 1.8448 | 1.9637 | 1.8923 | 1.8416 | 1.8292
| 2 | 8 | 1.9064 | 1.8517 | 1.8532 | 1.8401 | 1.8638 | 1.8297 | 1.8481 | 1.8342
| 2 | 9 | 1.9984 | 1.8956 | 1.8155 | 1.8305 | 1.8039 | 1.8456 | 1.8606 | 1.8408
| 2 | 10 | 1.9529 | 1.828 | 1.9251 | 1.8357 | 1.8442 | 1.8536 | 1.8759 | 1.9469
| 2 | 11 | 1.8844 | 1.8256 | 1.8224 | 1.8746 | 1.7938 | 1.8641 | 1.8668 | 1.8405
| 2 | 12 | 1.7816 | 1.8499 | 1.8415 | 1.8797 | 1.8131 | 1.8126 |  | 1.8711
| 3 | 13 | 1.8594 | 1.87 | 1.9157 | 1.8866
| 3 | 14 | 1.8872 | 1.87 | 1.8552 | 1.8302
| 3 | 15 | 1.8219 | 1.9 | 1.9399 | 1.8647
| 3 | 16 | 1.9368 | 1.85 | 1.8267 | 1.8175
| 3 | 17 | 1.8578 | 1.88 | 1.8088 | 1.9523
| 3 | 18 | 1.9687 | 1.78 | 1.8275 | 1.8878
| 4 | 19 | 1.9875 | 1.82 | 1.92 | 1.8574
| 4 | 20 | 1.9042 | 1.76 | 1.92 | 1.8232
| 4 | 21 |  | 1.79 | 1.84 | 1.788
| 4 | 22 |  | 1.79 | 1.93 | 1.7025
| 4 | 23 |  | 1.84 | 1.83 | 1.7398
| 4 | 24 |  | 1.85 | 1.68 | 1.7852
| 5 | 25 |  |  | 1.92 | 1.8454
| 5 | 26 |  |  | 1.83 | 1.8264
| 5 | 27 |  |  | 1.92 | 1.8446
| 5 | 28 |  |  | 1.87 | 1.8173
| 5 | 29 |  |  | 1.8 | 1.768
| 5 | 30 |  |  | 1.76 | 1.8089
| 6 | 31 |  |  | 1.84 | 1.9212
| 6 | 32 |  |  | 1.78 | 1.8637
| 6 | 33 |  |  | 1.79 | 1.8579
| 6 | 34 |  |  | 1.92 | 1.8329
| 6 | 35 |  |  | 1.78 | 1.7743
| 6 | 36 |  |  | 1.9 | 1.8001
| Mean |  | 1.9038 | 1.8532 | 1.8476 | 1.8421 | 1.819 | 1.8341 | 1.8772 | 1.8691
| Median |  | 1.9047 | 1.849 | 1.84 | 1.8415 | 1.8146 | 1.8383 | 1.8668 | 1.8559
| Std Dev. |  | 0.05615 | 0.064788 | 0.058467 | 0.054365 | 0.056649 | 0.038171 | 0.056802 | 0.047666
| Rel.Std.Dev. |  | 2.949% | 3.496% | 3.164% | 2.951% | 3.114% | 2.081% | 3.026% | 2.55%
| PDM3 |  | 3.047% | 0.3105% | 0.006857% | -0.2918% | -1.54% | -0.7263% | 1.606% | 1.169%
| Z-Score (Absolute) |  | NA | 0.28 | 0 | 0.27 | 1.42 | 0.67 | 1.47 | 1.07
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 236 (Certified Value 0.478 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.5 | 0.5 | 0.8385 | 0.38 | 0.5 | 0.49 | 0.66 | 0.6 | < 5 | 0.5 | 0.45 |  | 0.5 | < 0.1 | 0.46 | 0.4 | 0.44 | 1 | 0.39 | 0.53 | 0.47 | 0.47 |  | 1
| 1 | 2 | 0.5 | 0.5 | 0.8415 | 0.37 | 0.5 | 0.5 | 0.59 | 0.7 | < 5 | 0.5 | 0.46 |  | 0.51 | < 0.1 | 0.47 | 0.4 | 0.433 | 0.9 | 0.38 | 0.49 | 0.45 | 0.53 |  | 18
| 1 | 3 | 0.49 | 0.5 | 0.8265 | 0.42 | 0.5 | 0.51 | 0.69 | 0.6 | < 5 | 0.5 | 0.45 |  | 0.51 | < 0.1 | 0.45 | 0.4 | 0.469 | 0.8 | 0.36 | 0.53 | 0.47 | 0.54 |  | 16
| 1 | 4 | 0.51 | 0.5 | 0.836 | 0.41 | 0.5 | 0.49 | 0.57 | 0.6 | < 5 | 0.5 | 0.48 |  | 0.52 | < 0.1 | 0.48 | 0.4 | 0.44 | 0.9 | 0.38 | 0.52 | 0.48 | 0.5 |  | 0.47814
| 1 | 5 | 0.48 | 0.5 | 0.8415 | 0.4 | 0.5 | 0.51 | 0.58 | 0.7 | < 5 | 0.5 | 0.47 |  | 0.52 | < 0.1 | 0.46 | 0.4 | 0.393 | 1 | 0.38 | 0.5 | 0.48 | 0.53 |  | 15
| 1 | 6 | 0.51 | 0.5 | 0.849 | 0.4 | 0.5 | 0.52 | 0.55 | 0.6 | < 5 | 0.5 | 0.54 |  | 0.53 | < 0.1 | 0.48 | 0.4 | 0.459 | 0.9 | 0.35 | 0.5 | 0.48 | 0.49
| Mean |  | 0.49833 | 0.5 | 0.83883 | 0.39667 | 0.5 | 0.50333 | 0.60667 | 0.63333 |  | 0.5 | 0.475 |  | 0.515 |  | 0.46667 | 0.4 | 0.439 | 0.91667 | 0.37333 | 0.51167 | 0.47167 | 0.51
| Median |  | 0.5 | 0.5 | 0.84 | 0.4 | 0.5 | 0.505 | 0.585 | 0.6 |  | 0.5 | 0.465 |  | 0.515 |  | 0.465 | 0.4 | 0.44 | 0.9 | 0.38 | 0.51 | 0.475 | 0.515
| Std Dev. |  | 0.01169 | 0 | 0.0074543 | 0.018619 | 0 | 0.012111 | 0.055377 | 0.05164 |  | 0 | 0.033912 |  | 0.010488 |  | 0.012111 | 6.0809e-17 | 0.026283 | 0.075277 | 0.015055 | 0.017224 | 0.01169 | 0.027568
| Rel.Std.Dev. |  | 2.346% | 0% | 0.8887% | 4.694% | 0% | 2.406% | 9.128% | 8.154% |  | 0% | 7.139% |  | 2.037% |  | 2.595% | 1.52e-14% | 5.987% | 8.212% | 4.033% | 3.366% | 2.479% | 5.406%
| PDM3 |  | 4.224% | 4.572% | 75.44% | -17.04% | 4.572% | 5.269% | 26.88% | 32.46% |  | 4.572% | -0.6565% |  | 7.709% |  | -2.399% | -16.34% | -8.186% | 91.72% | -21.92% | 7.012% | -1.354% | 6.664%
| Z-Score (Absolute) |  | 0.04 | NA | 4.75 | 1.46 | NA | 0.04 | 1.49 | NA | 28.08 | NA | 0.36 | NA | 0.2 | 6.33 | 0.48 | NA | 0.87 | 5.84 | 1.79 | 0.15 | 0.41 | 0.13
Results from laboratories 2, 7, 11, 13 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 236 (Certified Value 6.59 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.83 | 6.58 | 5.99 | 6.42 | 7.66 | 6.57 | 6.3947 | 6.43 | 6.713 | 7.155 | 6.4955 | 6.49 | 6.6177 | 6.63 | 6.44 | 6.62 | 6.52 | 6.5 | 6.81 | 6.4 | 6.83 | 6.31 |  | 1
| 1 | 2 | 6.53 | 6.65 | 6.065 | 6.55 | 7.24 | 6.61 | 6.3935 | 6.45 | 6.716 | 7.142 | 6.6183 | 6.38 | 6.6223 | 6.72 | 6.54 | 6.58 | 6.74 | 6.59 | 6.77 | 6.5 | 6.95 | 6.5 |  | #N/A
| 1 | 3 | 6.36 | 6.83 | 6.095 | 6.36 | 7.54 | 6.78 | 6.4269 | 6.52 | 6.748 | 7.1523 | 6.5888 | 6.51 | 6.6427 | 6.69 | 6.29 | 6.52 | 6.94 | 6.45 | 6.81 | 6.58 | 6.73 | 6.66 |  | 16
| 1 | 4 | 6.42 | 6.79 | 6.0225 | 6.51 | 7.41 | 6.7 | 6.3461 | 6.55 | 6.764 | 7.1403 | 6.6433 | 6.69 | 6.638 | 6.62 | 6.44 | 6.54 | 6.58 | 6.44 | 6.81 | 6.84 | 6.89 | 6.44 |  | 6.5932
| 1 | 5 | 6.36 | 6.7 | 6.05 | 6.66 | 7.33 | 6.68 | 6.3681 | 6.68 | 6.706 | 7.1554 | 6.5112 | 6.61 | 6.6874 | 6.59 | 6.45 | 6.59 | 5.87 | 6.55 | 6.8 | 6.6 | 6.93 | 6.49 |  | 16
| 1 | 6 | 6.4 | 6.74 | 6.11 | 6.49 | 7.44 | 6.66 | 6.4701 | 6.62 | 6.748 | 7.1411 | 6.5022 | 6.85 | 6.5429 | 6.61 | 6.63 | 6.54 | 6.44 | 6.43 | 6.79 | 6.51 | 6.91 | 6.45
| Mean |  | 6.4833 | 6.715 | 6.0554 | 6.4983 | 7.4367 | 6.6667 | 6.3999 | 6.5417 | 6.7325 | 7.1477 | 6.5599 | 6.5883 | 6.6252 | 6.6433 | 6.465 | 6.565 | 6.515 | 6.4933 | 6.7983 | 6.5717 | 6.8733 | 6.475
| Median |  | 6.41 | 6.72 | 6.0575 | 6.5 | 7.425 | 6.67 | 6.3941 | 6.535 | 6.732 | 7.1471 | 6.55 | 6.56 | 6.6301 | 6.625 | 6.445 | 6.56 | 6.55 | 6.475 | 6.805 | 6.545 | 6.9 | 6.47
| Std Dev. |  | 0.18096 | 0.091815 | 0.044789 | 0.1042 | 0.14922 | 0.073121 | 0.043888 | 0.096626 | 0.02378 | 0.0072895 | 0.064884 | 0.16642 | 0.047295 | 0.050465 | 0.11397 | 0.037815 | 0.36253 | 0.06532 | 0.016021 | 0.14919 | 0.081404 | 0.11327
| Rel.Std.Dev. |  | 2.791% | 1.367% | 0.7396% | 1.603% | 2.007% | 1.097% | 0.6858% | 1.477% | 0.3532% | 0.102% | 0.9891% | 2.526% | 0.7139% | 0.7596% | 1.763% | 0.576% | 5.565% | 1.006% | 0.2357% | 2.27% | 1.184% | 1.749%
| PDM3 |  | -1.666% | 1.848% | -8.156% | -1.439% | 12.79% | 1.115% | -2.932% | -0.7813% | 2.113% | 8.41% | -0.505% | -0.07354% | 0.4852% | 0.7607% | -1.944% | -0.4274% | -1.186% | -1.514% | 3.112% | -0.3263% | 4.249% | -1.792%
| Z-Score (Absolute) |  | 0.64 | 1.11 | 3.88 | 0.53 | 6.57 | 0.74 | 1.27 | 0.2 | 1.24 | 4.38 | 0.06 | 0.15 | 0.43 | 0.57 | 0.78 | 0.03 | 0.4 | 0.57 | 1.74 | 0.03 | 2.31 | 0.71
| Analytical results for As in OREAS 236 (Certified Value 68 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 73 | 73 | 60.606 | 63 | 70.1 | 70.1 | 66.9 | 74.2 | 72 | 52.274 | 62 | 43.5 | 66 | 64.8 | 68 | 68.6 | 64 | 64 | 70 | 62.2 | 66.5 |  |  | 1
| 1 | 2 | 71.2 | 73 | 60.788 | 64 | 70.8 | 70.5 | 67.1 | 74.1 | 73 | 51.393 | 64 | 44 | 68 | 66.2 | 69 | 65.6 | 65 | 66 | 68 | 63.1 | 70.8 |  |  | 6
| 1 | 3 | 68.6 | 72 | 60.186 | 62 | 71.5 | 72.5 | 65.6 | 75 | 73 | 52.114 | 65 | 49.7 | 65 | 65 | 69 | 71.3 | 60 | 69 | 71.4 | 60.9 | 71.4 |  |  | 16
| 1 | 4 | 68.9 | 72 | 60.144 | 64 | 69.4 | 72.1 | 66.9 | 73.9 | 73 | 51.978 | 65 | 51.3 | 67 | 65.8 | 70 | 65.3 | 64 | 63 | 69.5 | 63.5 | 70.7 |  |  | 67.6896
| 1 | 5 | 69.2 | 70 | 59.9072 | 62 | 69.4 | 71.6 | 67.7 | 73 | 73 | 52.063 | 64 | 54.5 | 68 | 65.8 | 69 | 60.9 | 64 | 68 | 71.7 | 65.8 | 70.6 |  |  | 17
| 1 | 6 | 70.6 | 68 | 59.465 | 63 | 71.3 | 72.7 | 67 | 75.1 | 72 | 51.634 | 63 | 61.9 | 66 | 66.9 | 70 | 68 | 64 | 66 | 68.5 | 61.2 | 70.1
| Mean |  | 70.25 | 71.3333 | 60.1827 | 63 | 70.4167 | 71.5833 | 66.8667 | 74.2167 | 72.6667 | 51.9093 | 63.8333 | 50.8167 | 66.6667 | 65.75 | 69.1667 | 66.6167 | 63.5 | 66 | 69.85 | 62.7833 | 70.0167
| Median |  | 69.9 | 72 | 60.165 | 63 | 70.45 | 71.85 | 66.95 | 74.15 | 73 | 52.0205 | 64 | 50.5 | 66.5 | 65.8 | 69 | 66.8 | 64 | 66 | 69.75 | 62.65 | 70.65
| Std Dev. |  | 1.6873 | 1.9664 | 0.47709 | 0.89443 | 0.92394 | 1.0704 | 0.68896 | 0.77309 | 0.5164 | 0.33034 | 1.169 | 6.8977 | 1.2111 | 0.77395 | 0.75277 | 3.5561 | 1.7607 | 2.2804 | 1.4977 | 1.7949 | 1.7725
| Rel.Std.Dev. |  | 2.402% | 2.757% | 0.7927% | 1.42% | 1.312% | 1.495% | 1.03% | 1.042% | 0.7106% | 0.6364% | 1.831% | 13.57% | 1.817% | 1.177% | 1.088% | 5.338% | 2.773% | 3.455% | 2.144% | 2.859% | 2.532%
| PDM3 |  | 3.783% | 5.383% | -11.09% | -6.928% | 4.029% | 5.752% | -1.216% | 9.643% | 7.353% | -23.31% | -5.697% | -24.93% | -1.511% | -2.865% | 2.182% | -1.585% | -6.189% | -2.496% | 3.192% | -7.248% | 3.438%
| Z-Score (Absolute) |  | 0.67 | 0.88 | 1.22 | 0.69 | 0.71 | 0.93 | 0.04 | 1.42 | 1.13 | 2.78 | 0.53 | 2.98 | 0 | 0.17 | 0.47 | 0.01 | 0.6 | 0.13 | 0.6 | 0.73 | 0.63
| Analytical results for Ba in OREAS 236 (Certified Value 246 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 260 | 247 | 297 | 230 | 254.6 | 255 | 258 | 223.645 | 238 | 257 | 248.283 | 229 | 234.7 | 240 | 231 | 243 | 239 | 254 | 250 | 246 | 240 |  |  | 1
| 1 | 2 | 260 | 251 | 294 | 231 | 252.1 | 251 | 258 | 221.789 | 239 | 254 | 248.67 | 241 | 237.9 | 250 | 235 | 242 | 248 | 251 | 250 | 250 | 250 |  |  | 20
| 1 | 3 | 250 | 249 | 304 | 238 | 249.2 | 254 | 257 | 223.645 | 244 | 253 | 249.948 | 233 | 231.8 | 240 | 229 | 249 | 243 | 252 | 260 | 245 | 260 |  |  | 16
| 1 | 4 | 250 | 250 | 303 | 236 | 249.5 | 250 | 254 | 222.978 | 254 | 258 | 249.75 | 233 | 230.6 | 240 | 234 | 229 | 238 | 252 | 260 | 248 | 250 |  |  | 245.7522
| 1 | 5 | 250 | 248 | 303 | 231 | 257.3 | 245 | 256 | 223.546 | 239 | 253 | 254.964 | 231 | 230.7 | 240 | 233 | 212 | 242 | 273 | 260 | 249 | 250 |  |  | 18
| 1 | 6 | 250 | 248 | 299 | 232 | 254.7 | 254 | 254 | 222.857 | 240 | 278 | 245.736 | 232 | 230.5 | 250 | 233 | 240 | 247 | 254 | 250 | 249 | 250
| Mean |  | 253.3333 | 248.8333 | 300 | 233 | 252.9 | 251.5 | 256.1667 | 223.0767 | 242.3333 | 258.8333 | 249.5585 | 233.1667 | 232.7 | 243.3333 | 232.5 | 235.8333 | 242.8333 | 256 | 255 | 247.8333 | 250
| Median |  | 250 | 248.5 | 301 | 231.5 | 253.35 | 252.5 | 256.5 | 223.262 | 239.5 | 255.5 | 249.21 | 232.5 | 231.25 | 240 | 233 | 241 | 242.5 | 253 | 255 | 248.5 | 250
| Std Dev. |  | 5.164 | 1.472 | 4 | 3.2249 | 3.2056 | 3.7283 | 1.8348 | 0.71866 | 6.0882 | 9.6212 | 3.0475 | 4.1191 | 3.0033 | 5.164 | 2.1679 | 13.3778 | 4.0702 | 8.4143 | 5.4772 | 1.9408 | 6.3246
| Rel.Std.Dev. |  | 2.038% | 0.5915% | 1.333% | 1.384% | 1.268% | 1.482% | 0.7163% | 0.3222% | 2.512% | 3.717% | 1.221% | 1.767% | 1.291% | 2.122% | 0.9325% | 5.673% | 1.676% | 3.287% | 2.148% | 0.7831% | 2.53%
| PDM3 |  | 3.085% | 1.254% | 22.07% | -5.189% | 2.909% | 2.339% | 4.238% | -9.227% | -1.391% | 5.323% | 1.549% | -5.121% | -5.311% | -0.9843% | -5.393% | -4.036% | -1.188% | 4.17% | 3.763% | 0.8468% | 1.728%
| Z-Score (Absolute) |  | 0.47 | 0 | 5.31 | 1.64 | 0.42 | 0.28 | 0.76 | 2.67 | 0.67 | 1.04 | 0.08 | 1.63 | 1.67 | 0.57 | 1.69 | 1.35 | 0.62 | 0.74 | 0.64 | 0.1 | 0.12
| Analytical results for Be in OREAS 236 (Certified Value 0.46 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.49 | 0.5 | < 5 | 0.5 | 0.2 | < 1 | 0.45 | 0.33 | < 0.5 | < 5 | < 0.5 | 0.4 | 0.47333 | 0.4 | 0.42 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.52 | 0.49 | 0.44 |  | 1
| 1 | 2 | 0.44 | 0.5 | < 5 | 0.5 | 0.2 | < 1 | 0.48 | 0.33 | < 0.5 | < 5 | < 0.5 | 0.4 | 0.50526 | 0.5 | 0.42 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.46 | 0.48 | 0.45 |  | 21
| 1 | 3 | 0.47 | 0.4 | < 5 | 0.6 | 0.2 | < 1 | 0.45 | 0.54 | < 0.5 | < 5 | < 0.5 | 0.4 | 0.4798 | 0.4 | 0.4 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.43 | 0.48 | 0.48 |  | 16
| 1 | 4 | 0.48 | 0.5 | < 5 | 0.6 | 0.2 | 1 | 0.42 | 0.47 | < 0.5 | < 5 | < 0.5 | 0.4 | 0.50921 | 0.5 | 0.44 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.44 | 0.48 | 0.46 |  | 0.45857
| 1 | 5 | 0.45 | 0.5 | < 5 | 0.5 | 0.2 | < 1 | 0.44 | 0.37 | < 0.5 | < 5 | < 0.5 | 0.4 | 0.48393 | 0.4 | 0.45 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.42 | 0.52 | 0.45 |  | 19
| 1 | 6 | 0.43 | 0.5 | < 5 | 0.5 | 0.2 | < 1 | 0.46 | 0.46 | < 0.5 | < 5 | < 0.5 | 0.5 | 0.48118 | 0.5 | 0.44 | < 0.5 | < 1 | < 0.5 | 0.47 | 0.45 | 0.47 | 0.42
| Mean |  | 0.46 | 0.48333 |  | 0.53333 | 0.2 | 1 | 0.45 | 0.41667 |  |  |  | 0.41667 | 0.48878 | 0.45 | 0.42833 |  |  |  | 0.47 | 0.45333 | 0.48667 | 0.45
| Median |  | 0.46 | 0.5 |  | 0.5 | 0.2 | 1 | 0.45 | 0.415 |  |  |  | 0.4 | 0.48256 | 0.45 | 0.43 |  |  |  | 0.47 | 0.445 | 0.48 | 0.45
| Std Dev. |  | 0.023664 | 0.040825 |  | 0.05164 | 3.0405e-17 |  | 0.02 | 0.086178 |  |  |  | 0.040825 | 0.014762 | 0.054772 | 0.018348 |  |  |  | 6.0809e-17 | 0.03559 | 0.017512 | 0.02
| Rel.Std.Dev. |  | 5.144% | 8.447% |  | 9.682% | 1.52e-14% |  | 4.444% | 20.68% |  |  |  | 9.798% | 3.02% | 12.17% | 4.284% |  |  |  | 1.294e-14% | 7.851% | 3.598% | 4.444%
| PDM3 |  | 0.3121% | 5.4% |  | 16.3% | -56.39% | 118.1% | -1.869% | -9.138% |  |  |  | -9.138% | 6.589% | -1.869% | -6.593% |  |  |  | 2.493% | -1.142% | 6.127% | -1.869%
| Z-Score (Absolute) |  | 0.12 | NA | 37.64 | NA | NA | 2.39 | 0.06 | 0.67 | 3.74 | 37.64 | 3.74 | NA | 0.65 | NA | 0.46 | 3.74 | 0.86 | 3.74 | 0.31 | 0 | 0.61 | 0.06
Results from laboratories 2, 4, 5, 14 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 236 (Certified Value 0.061 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.07 | < 0.1 | < 5 | 0.07 | 0.21 | < 0.1 | 0.06 | 0.08 | 0.06 | < 5 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 14 | 0.04 | 0.06 | 0.07 | 0.06 |  |  | 1
| 1 | 2 | 0.07 | < 0.1 | < 5 | 0.09 | 0.21 | < 0.1 | 0.06 | 0.07 | 0.07 | < 5 | 0.06 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 15 | 0.04 | 0.06 | 0.06 | 0.06 |  |  | 22
| 1 | 3 | 0.07 | < 0.1 | < 5 | 0.07 | 0.21 | < 0.1 | 0.06 | 0.09 | 0.07 | < 5 | 0.06 | 0.05 | 0.06 | 0.05 | 0.07 | 0.06 | 16 | 0.03 | 0.07 | 0.07 | 0.06 |  |  | 16
| 1 | 4 | 0.07 | < 0.1 | < 5 | 0.09 | 0.2 | < 0.1 | 0.07 | 0.09 | 0.07 | < 5 | 0.07 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06 | 18 | 0.03 | 0.06 | 0.07 | 0.06 |  |  | 0.061389
| 1 | 5 | 0.06 | < 0.1 | < 5 | 0.08 | 0.19 | < 0.1 | 0.06 | 0.08 | 0.06 | < 5 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 14 | 0.04 | 0.06 | 0.07 | 0.07 |  |  | 20
| 1 | 6 | 0.07 | < 0.1 | < 5 | 0.09 | 0.2 | < 0.1 | 0.06 | 0.09 | 0.07 | < 5 | 0.06 | 0.07 | 0.06 | 0.05 | 0.07 | 0.06 | 21 | 0.03 | 0.06 | 0.07 | 0.06
| Mean |  | 0.068333 |  |  | 0.081667 | 0.20333 |  | 0.061667 | 0.083333 | 0.066667 |  | 0.058333 | 0.06 | 0.058333 | 0.05 | 0.063333 | 0.058333 | 16.3333 | 0.035 | 0.061667 | 0.068333 | 0.061667
| Median |  | 0.07 |  |  | 0.085 | 0.205 |  | 0.06 | 0.085 | 0.07 |  | 0.06 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 15.5 | 0.035 | 0.06 | 0.07 | 0.06
| Std Dev. |  | 0.0040825 |  |  | 0.0098319 | 0.008165 |  | 0.0040825 | 0.008165 | 0.005164 |  | 0.0075277 | 0.0063246 | 0.0040825 | 7.6012e-18 | 0.005164 | 0.0040825 | 2.7325 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 5.974% |  |  | 12.04% | 4.016% |  | 6.62% | 9.798% | 7.746% |  | 12.9% | 10.54% | 6.999% | 1.52e-14% | 8.154% | 6.999% | 16.73% | 15.65% | 6.62% | 5.974% | 6.62%
| PDM3 |  | 11.31% |  |  | 33.03% | 231.2% |  | 0.4525% | 35.75% | 8.597% |  | -4.977% | -2.262% | -4.977% | -18.55% | 3.167% | -4.977% | 26506% | -42.99% | 0.4525% | 11.31% | 0.4525%
| Z-Score (Absolute) |  | 0.67 | 1.18 | 246.63 | 2.02 | 14.33 | 1.18 | 0 | 2.19 | 0.51 | 246.63 | 0.34 | 0.17 | 0.34 | 1.18 | 0.17 | 0.34 | NA | 2.7 | 0 | 0.67 | 0
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 236 (Certified Value 5.93 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.11 | 5.9 | 5.25 | 5.55 | 6.35 | 6.28 | 5.8366 | 5.74 | 5.985 | 6.3698 | 5.7809 | 5.75 | 5.8323 | 5.59 | 5.96 | 6.48 | 5.74 | 5.45 | 6.04 | 5.81 | 5.97 | 5.79 |  | 1
| 1 | 2 | 5.81 | 5.95 | 5.25 | 5.73 | 6.14 | 6.33 | 5.9055 | 5.82 | 6.005 | 6.266 | 5.9027 | 5.67 | 5.8707 | 5.66 | 5.99 | 6.34 | 5.74 | 5.54 | 5.97 | 5.88 | 6.13 | 6.02 |  | #N/A
| 1 | 3 | 5.7 | 6.04 | 5.28 | 5.59 | 6.32 | 6.32 | 5.9075 | 5.89 | 6.079 | 6.3498 | 5.8816 | 5.66 | 5.9043 | 5.66 | 5.83 | 6.4 | 5.78 | 5.42 | 6.03 | 5.96 | 5.95 | 6.09 |  | 16
| 1 | 4 | 5.76 | 5.98 | 5.275 | 5.7 | 6.2 | 6.36 | 5.8084 | 5.92 | 6.056 | 6.266 | 5.9076 | 5.82 | 5.9432 | 5.65 | 6 | 6.4 | 5.65 | 5.39 | 6.01 | 6.15 | 6.07 | 5.95 |  | 5.9341
| 1 | 5 | 5.67 | 5.91 | 5.29 | 5.72 | 6.17 | 6.34 | 5.997 | 6.01 | 6.098 | 6.3798 | 5.823 | 5.74 | 5.8529 | 5.55 | 5.98 | 6.42 | 5.16 | 5.53 | 6 | 5.97 | 6 | 5.98 |  | 21
| 1 | 6 | 5.78 | 5.98 | 5.29 | 5.64 | 6.23 | 6.24 | 5.9185 | 5.99 | 5.992 | 6.276 | 5.8066 | 6.56 | 5.8153 | 5.66 | 6.11 | 6.4 | 5.64 | 5.44 | 6.04 | 5.85 | 6.03 | 5.91
| Mean |  | 5.805 | 5.96 | 5.2725 | 5.655 | 6.235 | 6.3117 | 5.8956 | 5.895 | 6.0358 | 6.3179 | 5.8504 | 5.8667 | 5.8698 | 5.6283 | 5.9783 | 6.4067 | 5.6183 | 5.4617 | 6.015 | 5.9367 | 6.025 | 5.9567
| Median |  | 5.77 | 5.965 | 5.2775 | 5.67 | 6.215 | 6.325 | 5.9065 | 5.905 | 6.0305 | 6.3129 | 5.8523 | 5.745 | 5.8618 | 5.655 | 5.985 | 6.4 | 5.695 | 5.445 | 6.02 | 5.92 | 6.015 | 5.965
| Std Dev. |  | 0.15808 | 0.051769 | 0.018371 | 0.073959 | 0.083606 | 0.044008 | 0.066529 | 0.10252 | 0.048147 | 0.054186 | 0.053815 | 0.34465 | 0.047421 | 0.047081 | 0.089759 | 0.045019 | 0.23121 | 0.06047 | 0.027386 | 0.1216 | 0.066858 | 0.10231
| Rel.Std.Dev. |  | 2.723% | 0.8686% | 0.3484% | 1.308% | 1.341% | 0.6972% | 1.128% | 1.739% | 0.7977% | 0.8577% | 0.9198% | 5.875% | 0.8079% | 0.8365% | 1.501% | 0.7027% | 4.115% | 1.107% | 0.4553% | 2.048% | 1.11% | 1.718%
| PDM3 |  | -2.176% | 0.4361% | -11.15% | -4.704% | 5.07% | 6.362% | -0.6494% | -0.6592% | 1.714% | 6.468% | -1.411% | -1.137% | -1.084% | -5.153% | 0.7451% | 7.963% | -5.322% | -7.962% | 1.363% | 0.04292% | 1.531% | 0.3799%
| Z-Score (Absolute) |  | 0.68 | 0.27 | 3.95 | 1.6 | 1.95 | 2.42 | 0.13 | 0.13 | 0.73 | 2.46 | 0.4 | 0.3 | 0.28 | 1.76 | 0.38 | 3.01 | 1.83 | 2.79 | 0.61 | 0.13 | 0.67 | 0.25
| Analytical results for Cd in OREAS 236 (Certified Value 0.62 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.69 | 0.5 | 0.84646 | 0.61 | 0.63 | 0.7 | 0.68 | 0.7 | 0.72 | < 5 | 0.58 | 0.6 | 0.68 | 0.9 | 0.57 | 0.58 | 0.6 | < 1 | < 1 | 0.65 | 0.6 | 0.57 |  | 1
| 1 | 2 | 0.65 | 0.5 | 0.85109 | 0.62 | 0.64 | 0.7 | 0.7 | 0.58 | 0.67 | < 5 | 0.59 | 0.6 | 0.72 | 0.9 | 0.57 | 0.57 | 0.58 | < 1 | < 1 | 0.59 | 0.57 | 0.63 |  | 23
| 1 | 3 | 0.64 | 0.5 | 0.84967 | 0.62 | 0.61 | 0.6 | 0.72 | 0.72 | 0.66 | < 5 | 0.58 | 0.5 | 0.67 | 0.9 | 0.55 | 0.58 | 0.61 | < 1 | < 1 | 0.65 | 0.61 | 0.61 |  | 16
| 1 | 4 | 0.64 | 0.5 | 0.84456 | 0.61 | 0.62 | 0.7 | 0.66 | 0.78 | 0.72 | < 5 | 0.62 | 0.6 | 0.69 | 0.9 | 0.59 | 0.6 | 0.58 | < 1 | < 1 | 0.61 | 0.6 | 0.61 |  | 0.62376
| 1 | 5 | 0.63 | 0.5 | 0.85187 | 0.61 | 0.6 | 0.8 | 0.68 | 0.55 | 0.66 | < 5 | 0.56 | 0.6 | 0.7 | 0.9 | 0.55 | 0.6 | 0.52 | < 1 | < 1 | 0.6 | 0.62 | 0.62 |  | 22
| 1 | 6 | 0.66 | 0.5 | 0.83833 | 0.6 | 0.63 | 0.8 | 0.61 | 0.63 | 0.66 | < 5 | 0.59 | 0.8 | 0.65 | 0.9 | 0.58 | 0.59 | 0.6 | < 1 | < 1 | 0.62 | 0.59 | 0.61
| Mean |  | 0.65167 | 0.5 | 0.847 | 0.61167 | 0.62167 | 0.71667 | 0.675 | 0.66 | 0.68167 |  | 0.58667 | 0.61667 | 0.685 | 0.9 | 0.56833 | 0.58667 | 0.58167 |  |  | 0.62 | 0.59833 | 0.60833
| Median |  | 0.645 | 0.5 | 0.84806 | 0.61 | 0.625 | 0.7 | 0.68 | 0.665 | 0.665 |  | 0.585 | 0.6 | 0.685 | 0.9 | 0.57 | 0.585 | 0.59 |  |  | 0.615 | 0.6 | 0.61
| Std Dev. |  | 0.02137 | 0 | 0.0050785 | 0.0075277 | 0.01472 | 0.075277 | 0.037815 | 0.088318 | 0.029944 |  | 0.019664 | 0.098319 | 0.02429 | 0 | 0.016021 | 0.012111 | 0.032506 |  |  | 0.025298 | 0.017224 | 0.020412
| Rel.Std.Dev. |  | 3.279% | 0% | 0.5996% | 1.231% | 2.368% | 10.5% | 5.602% | 13.38% | 4.393% |  | 3.352% | 15.94% | 3.546% | 0% | 2.819% | 2.064% | 5.588% |  |  | 4.08% | 2.879% | 3.355%
| PDM3 |  | 4.474% | -19.84% | 35.79% | -1.939% | -0.3359% | 14.89% | 8.214% | 5.81% | 9.283% |  | -5.947% | -1.137% | 9.818% | 44.29% | -8.886% | -5.947% | -6.749% |  |  | -0.6031% | -4.077% | -2.473%
| Z-Score (Absolute) |  | 0.6 | NA | 3.9 | 0.07 | 0.1 | NA | 1 | 0.74 | 1.11 | 31.76 | 0.49 | NA | 1.17 | NA | 0.8 | 0.49 | 0.58 | 1.95 | 1.95 | 0.07 | 0.3 | 0.13
Results from laboratories 2, 7, 14 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 236 (Certified Value 14 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 14.9 | 14.7 | 14.13 | 17 | 14 | 13.68 | 13.92 | 7.589 | 13.7 | 11.9802 | 11.63 | 11.5 | 13.55 | 13 | 17 | 14.94 | 14.25 | 14.58 | 13.05 |  |  |  |  | 1
| 1 | 2 | 14.7 | 15 | 14.81 | 17.4 | 14 | 14.1 | 13.56 | 7.565 | 13.6 | 11.779 | 12.05 | 11.4 | 14 | 12.88 | 18 | 15.24 | 14.55 | 14.54 | 13.55 |  |  |  |  | 24
| 1 | 3 | 14.45 | 15 | 14.61 | 17.3 | 14 | 13.86 | 13.8 | 7.51 | 13.6 | 11.8309 | 12.01 | 12.4 | 13.15 | 13.07 | 16 | 15.15 | 15.2 | 14.29 | 14.25 |  |  |  |  | 16
| 1 | 4 | 14.8 | 14.8 | 14.91 | 16.9 | 14 | 13.93 | 13.51 | 7.61 | 13.6 | 11.9054 | 12.06 | 11.8 | 14.05 | 12.25 | 17 | 15.06 | 14.75 | 14.78 | 13.45 |  |  |  |  | 13.9587
| 1 | 5 | 14.1 | 14.9 | 14.22 | 16.8 | 15 | 14.14 | 13.53 | 7.55 | 13.8 | 11.1385 | 12.22 | 12.9 | 14 | 11.29 | 17 | 16.49 | 15 | 15.4 | 13.35 |  |  |  |  | 23
| 1 | 6 | 14.8 | 15.2 | 15.06 | 16.5 | 14 | 14.12 | 13.55 | 7.42 | 14 | 12.0197 | 11.97 | 13.6 | 14.05 | 12.65 | 16 | 15.02 | 14.6 | 14.54 | 13.4
| Mean |  | 14.625 | 14.9333 | 14.6233 | 16.9833 | 14.1667 | 13.9717 | 13.645 | 7.5407 | 13.7167 | 11.7756 | 11.99 | 12.2667 | 13.8 | 12.5233 | 16.8333 | 15.3167 | 14.725 | 14.6883 | 13.5083
| Median |  | 14.75 | 14.95 | 14.71 | 16.95 | 14 | 14.015 | 13.555 | 7.5575 | 13.65 | 11.8681 | 12.03 | 12.1 | 14 | 12.765 | 17 | 15.105 | 14.675 | 14.56 | 13.425
| Std Dev. |  | 0.29958 | 0.17512 | 0.37787 | 0.33116 | 0.40825 | 0.18225 | 0.17167 | 0.068298 | 0.16021 | 0.32473 | 0.19586 | 0.86641 | 0.37148 | 0.67307 | 0.75277 | 0.58415 | 0.33875 | 0.38191 | 0.40052
| Rel.Std.Dev. |  | 2.048% | 1.173% | 2.584% | 1.95% | 2.882% | 1.304% | 1.258% | 0.9057% | 1.168% | 2.758% | 1.634% | 7.063% | 2.692% | 5.375% | 4.472% | 3.814% | 2.3% | 2.6% | 2.965%
| PDM3 |  | 4.773% | 6.982% | 4.762% | 21.67% | 1.49% | 0.093% | -2.247% | -45.98% | -1.734% | -15.64% | -14.1% | -12.12% | -1.137% | -10.28% | 20.59% | 9.729% | 5.49% | 5.227% | -3.226%
| Z-Score (Absolute) |  | 0.6 | 0.83 | 0.6 | 2.32 | NA | 0.13 | 0.11 | 4.56 | 0.06 | 1.48 | 1.32 | 1.12 | 0 | 0.93 | NA | 1.11 | 0.67 | 0.65 | 0.21
Results from laboratories 7 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 236 (Certified Value 41.5 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 42.8 | 45 | 31.644 | 39.4 | 40.9 | 43.8 | 40.2 | 44.4 | 44.7 | 37.967 | 40.8 | 40.8 | 43.0463 | 44.9 | 40.7 | 38.7 | 38 | 42.4 | 39 | 40 | 39.9 | 43.9 | 41.2 | 1
| 1 | 2 | 40.8 | 45 | 31.8435 | 40.5 | 39.3 | 44.2 | 40.7 | 44.5 | 45.2 | 36.857 | 40.9 | 39.8 | 42.1399 | 45 | 42.2 | 38.8 | 37 | 42.5 | 40 | 40 | 40.3 | 43.4 | 42.3 | 10
| 1 | 3 | 39.4 | 45 | 31.9568 | 41.7 | 41.4 | 44.1 | 40.1 | 44.4 | 45 | 36.98 | 41.9 | 40.4 | 42.6372 | 45.1 | 39.8 | 38.2 | 37 | 43.6 | 38 | 41 | 40.7 | 42.8 | 43.8 | 16
| 1 | 4 | 41.1 | 45 | 32.1063 | 42.3 | 40.4 | 44.3 | 40.6 | 44.2 | 45.1 | 37.25 | 41.8 | 40.7 | 41.9795 | 45.9 | 39.5 | 39.2 | 37 | 41.6 | 39 | 41 | 41.2 | 43.7 | 41.8 | 41.4883
| 1 | 5 | 39.7 | 45 | 32.0872 | 42 | 39 | 44.9 | 41.9 | 43.6 | 45 | 37.33 | 41.2 | 40.8 | 41.4559 | 42.8 | 39.1 | 39 | 39 | 37.9 | 40 | 40 | 41 | 43.5 | 43.1 | 24
| 1 | 6 | 40.8 | 45 | 32.2073 | 40.7 | 42 | 45.2 | 41.4 | 45.1 | 44.7 | 37.52 | 41.4 | 43.7 | 42.2739 | 44.8 | 39.2 | 39.7 | 38 | 41.5 | 39 | 41 | 39.6 | 43.1 | 40.9
| Mean |  | 40.7667 | 45 | 31.9742 | 41.1 | 40.5 | 44.4167 | 40.8167 | 44.3667 | 44.95 | 37.3173 | 41.3333 | 41.0333 | 42.2555 | 44.75 | 40.0833 | 38.9333 | 37.6667 | 41.5833 | 39.1667 | 40.5 | 40.45 | 43.4 | 42.1833
| Median |  | 40.8 | 45 | 32.022 | 41.2 | 40.65 | 44.25 | 40.65 | 44.4 | 45 | 37.29 | 41.3 | 40.75 | 42.2069 | 44.95 | 39.65 | 38.9 | 37.5 | 42 | 39 | 40.5 | 40.5 | 43.45 | 42.05
| Std Dev. |  | 1.2044 | 0 | 0.2054 | 1.0973 | 1.1764 | 0.52694 | 0.70261 | 0.48442 | 0.20736 | 0.39838 | 0.45461 | 1.3604 | 0.54776 | 1.033 | 1.1856 | 0.50465 | 0.8165 | 1.957 | 0.75277 | 0.54772 | 0.62849 | 0.4 | 1.1161
| Rel.Std.Dev. |  | 2.954% | 0% | 0.6424% | 2.67% | 2.905% | 1.186% | 1.721% | 1.092% | 0.4613% | 1.068% | 1.1% | 3.315% | 1.296% | 2.308% | 2.958% | 1.296% | 2.168% | 4.706% | 1.922% | 1.352% | 1.554% | 0.9217% | 2.646%
| PDM3 |  | -1.739% | 8.464% | -22.93% | -0.9359% | -2.382% | 7.058% | -1.619% | 6.938% | 8.344% | -10.05% | -0.3735% | -1.097% | 1.849% | 7.862% | -3.386% | -6.158% | -9.211% | 0.229% | -5.596% | -2.382% | -2.503% | 4.608% | 1.675%
| Z-Score (Absolute) |  | 0.15 | 2.19 | 5 | 0.04 | 0.29 | 1.87 | 0.12 | 1.84 | 2.16 | 2.05 | 0.17 | 0 | 0.67 | 2.05 | 0.52 | 1.16 | 1.86 | 0.3 | 1.03 | 0.29 | 0.32 | 1.31 | 0.63
| Analytical results for Cr in OREAS 236 (Certified Value 86 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 84 | 80 | 81.263 | 86 | 105 | 103 | 90 | 100 | 86.496 | 75 | 86 | 81 | 89 | 89 | 88 | 95 | 91 | 86 | 86 | 85 | 81 |  |  | 1
| 1 | 2 | 81 | 80 | 83.648 | 89 | 109 | 103 | 88 | 102 | 86.199 | 76 | 83 | 82 | 88 | 84 | 85 | 94 | 91 | 87 | 85 | 90 | 85 |  |  | 26
| 1 | 3 | 81 | 80 | 82 | 87 | 109 | 101 | 85 | 103 | 86.146 | 77 | 92 | 88 | 94 | 77 | 85 | 96 | 90 | 86 | 88 | 90 | 86 |  |  | 16
| 1 | 4 | 83 | 80 | 82.2 | 88 | 103 | 101 | 91 | 100 | 86.246 | 78 | 87 | 87 | 88 | 88 | 85 | 94 | 93 | 88 | 90 | 86 | 84 |  |  | 86.4016
| 1 | 5 | 82 | 80 | 82.677 | 89 | 105 | 104 | 92 | 102 | 86.355 | 79 | 117 | 82 | 89 | 90 | 84 | 86 | 92 | 88 | 88 | 84 | 82 |  |  | 25
| 1 | 6 | 85 | 80 | 83.4915 | 87 | 105 | 103 | 93 | 103 | 86.241 | 76 | 109 | 87 | 84 | 87 | 86 | 93 | 92 | 87 | 87 | 83 | 84
| Mean |  | 82.6667 | 80 | 82.5466 | 87.6667 | 106 | 102.5 | 89.8333 | 101.6667 | 86.2805 | 76.8333 | 95.6667 | 84.5 | 88.6667 | 85.8333 | 85.5 | 93 | 91.5 | 87 | 87.3333 | 86.3333 | 83.6667
| Median |  | 82.5 | 80 | 82.4385 | 87.5 | 105 | 103 | 90.5 | 102 | 86.2435 | 76.5 | 89.5 | 84.5 | 88.5 | 87.5 | 85 | 94 | 91.5 | 87 | 87.5 | 85.5 | 84
| Std Dev. |  | 1.633 | 0 | 0.91536 | 1.2111 | 2.4495 | 1.2247 | 2.9269 | 1.3663 | 0.12607 | 1.472 | 13.9666 | 3.1464 | 3.2042 | 4.7924 | 1.3784 | 3.5777 | 1.0488 | 0.89443 | 1.7512 | 3.0111 | 1.8619
| Rel.Std.Dev. |  | 1.975% | 0% | 1.109% | 1.381% | 2.311% | 1.195% | 3.258% | 1.344% | 0.1461% | 1.916% | 14.6% | 3.724% | 3.614% | 5.583% | 1.612% | 3.847% | 1.146% | 1.028% | 2.005% | 3.488% | 2.225%
| PDM3 |  | -4.323% | -7.409% | -4.462% | 1.464% | 22.68% | 18.63% | 3.972% | 17.67% | -0.1402% | -11.07% | 10.72% | -2.201% | 2.622% | -0.6577% | -1.043% | 7.637% | 5.901% | 0.6926% | 1.078% | -0.079% | -3.165%
| Z-Score (Absolute) |  | 0.98 | NA | 1.01 | 0.11 | 4.12 | 3.35 | 0.58 | 3.17 | 0.19 | 2.26 | 1.86 | 0.58 | 0.33 | 0.29 | 0.36 | 1.28 | 0.95 | 0.04 | 0.04 | 0.18 | 0.77
Results from laboratory 2 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 236 (Certified Value 1.01 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.04 | 1 | 1.11 | 1.08 | 0.7 | 0.96 | 1.01 | 1 | 0.9 | 0.9 | 1.099 | 1 | 1.1 | 0.88 | 0.9 | 0.9 | 0.98 | 0.99 | 0.98 | 0.93 |  |  |  | 1
| 1 | 2 | 1 | 1 | 1.12 | 1.08 | 0.7 | 0.98 | 0.95 | 1 | 1 | 0.91 | 1.0467 | 1 | 1.1 | 0.89 | 0.8 | 0.9 | 1.01 | 0.97 | 0.98 | 0.97 |  |  |  | 27
| 1 | 3 | 1 | 1 | 1.15 | 1.11 | 0.8 | 0.97 | 1.18 | 1.1 | 1 | 0.94 | 1.0881 | 1 | 1.1 | 0.91 | 0.8 | 0.9 | 1.01 | 1.02 | 0.97 | 0.99 |  |  |  | 16
| 1 | 4 | 1 | 0.9 | 1.19 | 1.07 | 0.8 | 0.99 | 1.03 | 1.1 | 1 | 0.9 | 1.11 | 1 | 1.1 | 0.89 | 0.8 | 0.9 | 1 | 0.99 | 1 | 0.96 |  |  |  | 1.0081
| 1 | 5 | 1 | 1 | 1.02 | 1.05 | 1 | 1.03 | 1.08 | 1 | 0.9 | 0.98 | 1.0404 | 1 | 1.1 | 0.88 | 0.8 | 0.8 | 1.07 | 1.02 | 1.04 | 0.96 |  |  |  | 26
| 1 | 6 | 1.01 | 1 | 1.15 | 1.13 | 0.8 | 0.99 | 1 | 1 | 0.9 | 1.05 | 1.0764 | 1 | 1.1 | 0.9 | 0.9 | 0.9 | 1 | 0.96 | 0.97 | 0.97
| Mean |  | 1.0083 | 0.98333 | 1.1233 | 1.0867 | 0.8 | 0.98667 | 1.0417 | 1.0333 | 0.95 | 0.94667 | 1.0768 | 1 | 1.1 | 0.89167 | 0.83333 | 0.88333 | 1.0117 | 0.99167 | 0.99 | 0.96333
| Median |  | 1 | 1 | 1.135 | 1.08 | 0.8 | 0.985 | 1.02 | 1 | 0.95 | 0.925 | 1.0822 | 1 | 1.1 | 0.89 | 0.8 | 0.9 | 1.005 | 0.99 | 0.98 | 0.965
| Std Dev. |  | 0.016021 | 0.040825 | 0.05785 | 0.028752 | 0.10954 | 0.024221 | 0.079854 | 0.05164 | 0.054772 | 0.059217 | 0.028133 | 0 | 2.4324e-16 | 0.01169 | 0.05164 | 0.040825 | 0.030605 | 0.024833 | 0.026833 | 0.019664
| Rel.Std.Dev. |  | 1.589% | 4.152% | 5.15% | 2.646% | 13.69% | 2.455% | 7.666% | 4.997% | 5.766% | 6.255% | 2.613% | 0% | 2.211e-14% | 1.311% | 6.197% | 4.622% | 3.025% | 2.504% | 2.71% | 2.041%
| PDM3 |  | 0.0268% | -2.453% | 11.43% | 7.797% | -20.64% | -2.123% | 3.333% | 2.507% | -5.76% | -6.091% | 6.815% | -0.7999% | 9.12% | -11.55% | -17.33% | -12.37% | 0.3575% | -1.627% | -1.792% | -4.437%
| Z-Score (Absolute) |  | 0.14 | NA | 2.12 | 1.49 | NA | 0.23 | 0.72 | NA | NA | 0.92 | 1.32 | NA | NA | 1.87 | NA | NA | 0.2 | 0.14 | 0.17 | 0.63
Results from laboratories 2, 7, 11, 13, 17, 19 and 20 were removed due to their 0.1 ppm reading resolution. Results from laboratory 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 236 (Certified Value 170 ppm) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 176.5 | 166 | 164.024 | 167.4 | 187 | 163 | 166.5 | 174.4 | 166 | 166 | 165 | 167.8448 | 179 | 164 | 166.5 | 161 | 167.7 | 165 | 174 | 178.5 | 176.7 | 167.5 |  | 1
| 1 | 2 | 170 | 170 | 163.31 | 166 | 181 | 161 | 170.7 | 171.8 | 166 | 172 | 160 | 171.8192 | 185 | 169.5 | 167.5 | 161 | 166.4 | 167 | 175 | 175.5 | 179.9 | 175 |  | 12
| 1 | 3 | 167 | 174 | 161.98 | 170.1 | 184 | 172 | 167.9 | 173.9 | 167 | 171 | 161 | 173.0888 | 184 | 159.8 | 167 | 164 | 167.9 | 163 | 172 | 180 | 175.6 | 180 |  | 16
| 1 | 4 | 168 | 176 | 162.428 | 171.9 | 183 | 176 | 165.1 | 172.5 | 168 | 171 | 163 | 172.0584 | 180 | 160.2 | 168.5 | 163 | 162.4 | 162 | 169 | 182 | 177.7 | 175 |  | 169.6544
| 1 | 5 | 164 | 168 | 161.462 | 174.2 | 181 | 175 | 174.1 | 167.2 | 168 | 166 | 166 | 173.124 | 179 | 161.6 | 167.5 | 162 | 147.6 | 165 | 170 | 179.5 | 178.2 | 176.5 |  | 27
| 1 | 6 | 169 | 170 | 160.748 | 175.6 | 186 | 168 | 171.2 | 173.4 | 167 | 169 | 171 | 170.062 | 181 | 160 | 172.5 | 163 | 163 | 167 | 174 | 174.5 | 178.4 | 173.5
| Mean |  | 169.0833 | 170.6667 | 162.3253 | 170.8667 | 183.6667 | 169.1667 | 169.25 | 172.2 | 167 | 169.1667 | 164.3333 | 171.3329 | 181.3333 | 162.5167 | 168.25 | 162.3333 | 162.5 | 164.8333 | 172.3333 | 178.3333 | 177.75 | 174.5833
| Median |  | 168.5 | 170 | 162.204 | 171 | 183.5 | 170 | 169.3 | 172.95 | 167 | 170 | 164 | 171.9388 | 180.5 | 160.9 | 167.5 | 162.5 | 164.7 | 165 | 173 | 179 | 177.95 | 175
| Std Dev. |  | 4.1763 | 3.7238 | 1.2015 | 3.7649 | 2.5033 | 6.2423 | 3.3477 | 2.6237 | 0.89443 | 2.6394 | 3.9833 | 2.0414 | 2.582 | 3.7653 | 2.1852 | 1.2111 | 7.6639 | 2.0412 | 2.4221 | 2.8402 | 1.4816 | 4.116
| Rel.Std.Dev. |  | 2.47% | 2.182% | 0.7402% | 2.203% | 1.363% | 3.69% | 1.978% | 1.524% | 0.5356% | 1.56% | 2.424% | 1.191% | 1.424% | 2.317% | 1.299% | 0.746% | 4.716% | 1.238% | 1.405% | 1.593% | 0.8335% | 2.358%
| PDM3 |  | -0.3366% | 0.5967% | -4.32% | 0.7146% | 8.259% | -0.2875% | -0.2383% | 1.5% | -1.565% | -0.2875% | -3.136% | 0.9894% | 6.884% | -4.207% | -0.8278% | -4.315% | -4.217% | -2.842% | 1.579% | 5.116% | 4.772% | 2.905%
| Z-Score (Absolute) |  | 0.02 | 0.26 | 1.24 | 0.3 | 2.6 | 0.01 | 0.01 | 0.54 | 0.4 | 0.01 | 0.88 | 0.38 | 2.18 | 1.2 | 0.17 | 1.24 | 1.21 | 0.79 | 0.56 | 1.64 | 1.54 | 0.97
| Analytical results for Dy in OREAS 236 (Certified Value 3.94 ppm) | [24 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 4.2 | 3.96 | 3.98 | 4.25 | 3.9 | 3.6756 | 3.86 | 3.6 | 3.7 | 4.12 | [13 empty] | 1
| 1 | 2 | 4.1 | 4.11 | 3.95 | 4.3 | 3.8 | 3.5866 | 3.89 | 3.5 | 3.7 | 4.21 | [13 empty] | 29
| 1 | 3 | 4 | 4.22 | 3.95 | 4.22 | 3.8 | 3.5888 | 3.92 | 3.8 | 3.7 | 4.2 | [13 empty] | 16
| 1 | 4 | 4.05 | 4.11 | 4.02 | 4.17 | 3.9 | 3.5689 | 3.9 | 3.8 | 3.4 | 4.2 | [13 empty] | 3.938
| 1 | 5 | 4.05 | 3.96 | 4 | 4.08 | 4 | 3.2847 | 3.81 | 4.3 | 3.2 | 4.13 | [13 empty] | 28
| 1 | 6 | 4.15 | 4.18 | 4.05 | 4.29 | 4 | 3.7831 | 3.88 | 4.5 | 3.6 | 4.12
| Mean |  | 4.0917 | 4.09 | 3.9917 | 4.2183 | 3.9 | 3.5813 | 3.8767 | 3.9167 | 3.55 | 4.1633
| Median |  | 4.075 | 4.11 | 3.99 | 4.235 | 3.9 | 3.5877 | 3.885 | 3.8 | 3.65 | 4.165
| Std Dev. |  | 0.073598 | 0.10918 | 0.039707 | 0.082805 | 0.089443 | 0.16604 | 0.038297 | 0.39707 | 0.20736 | 0.044121
| Rel.Std.Dev. |  | 1.799% | 2.669% | 0.9948% | 1.963% | 2.293% | 4.636% | 0.9879% | 10.14% | 5.841% | 1.06%
| PDM3 |  | 3.903% | 3.861% | 1.364% | 7.12% | -0.964% | -9.057% | -1.557% | -0.5408% | -9.852% | 5.723%
| Z-Score (Absolute) |  | 0.68 | 0.67 | 0.19 | 1.3 | 0.27 | 1.84 | 0.38 | 0.19 | 1.99 | 1.03
| Analytical results for Er in OREAS 236 (Certified Value 2.43 ppm) | [24 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 2.45 | 2.37 | 2.45 | 2.63 | 2.5 | 2.2511 | 2.42 | 1.7 | 2.2 | 2.59 | [13 empty] | 1
| 1 | 2 | 2.45 | 2.49 | 2.5 | 2.67 | 2.4 | 2.2132 | 2.38 | 1.6 | 2.2 | 2.61 | [13 empty] | 30
| 1 | 3 | 2.45 | 2.34 | 2.49 | 2.7 | 2.3 | 2.2279 | 2.46 | 1.7 | 2.2 | 2.58 | [13 empty] | 16
| 1 | 4 | 2.55 | 2.46 | 2.49 | 2.65 | 2.3 | 2.2593 | 2.44 | 1.7 | 2.1 | 2.59 | [13 empty] | 2.4286
| 1 | 5 | 2.55 | 2.42 | 2.48 | 2.6 | 2.4 | 2.1791 | 2.47 | 1.6 | 1.9 | 2.62 | [13 empty] | 29
| 1 | 6 | 2.5 | 2.59 | 2.53 | 2.62 | 2.3 | 2.2741 | 2.41 | 1.7 | 2.1 | 2.58
| Mean |  | 2.4917 | 2.445 | 2.49 | 2.645 | 2.3667 | 2.2341 | 2.43 | 1.6667 | 2.1167 | 2.595
| Median |  | 2.475 | 2.44 | 2.49 | 2.64 | 2.35 | 2.2395 | 2.43 | 1.7 | 2.15 | 2.59
| Std Dev. |  | 0.04916 | 0.090056 | 0.026077 | 0.036194 | 0.08165 | 0.034706 | 0.033466 | 0.05164 | 0.1169 | 0.016432
| Rel.Std.Dev. |  | 1.973% | 3.683% | 1.047% | 1.368% | 3.45% | 1.553% | 1.377% | 3.098% | 5.523% | 0.6332%
| PDM3 |  | 2.597% | 0.675% | 2.528% | 8.91% | -2.55% | -8.008% | 0.05737% | -31.37% | -12.84% | 6.851%
| Z-Score (Absolute) |  | 0.32 | 0.04 | 0.31 | 1.23 | 0.42 | 1.2 | 0.04 | 4.55 | 1.89 | 0.93
| Analytical results for Eu in OREAS 236 (Certified Value 0.96 ppm) | [24 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.95 | 0.98 | 0.96 | 1.03 | 0.94 | 0.95705 | 0.93 | 0.5 | 0.9 | 0.96 | [13 empty] | 1
| 1 | 2 | 0.9 | 1.01 | 0.96 | 1.08 | 0.89 | 0.93072 | 0.93 | 0.5 | 0.9 | 0.98 | [13 empty] | 31
| 1 | 3 | 0.95 | 1 | 0.99 | 1.02 | 0.95 | 0.94432 | 0.94 | 0.5 | 0.9 | 0.98 | [13 empty] | 16
| 1 | 4 | 0.95 | 1.01 | 0.97 | 0.96 | 0.96 | 0.97867 | 0.96 | 0.5 | 0.8 | 0.97 | [13 empty] | 0.96342
| 1 | 5 | 0.95 | 0.99 | 0.97 | 0.94 | 0.94 | 0.95141 | 0.93 | 0.5 | 0.8 | 0.98 | [13 empty] | 30
| 1 | 6 | 0.95 | 1.03 | 0.98 | 0.97 | 0.98 | 0.8982 | 0.95 | 0.5 | 0.8 | 0.96
| Mean |  | 0.94167 | 1.0033 | 0.97167 | 1 | 0.94333 | 0.9434 | 0.94 | 0.5 | 0.85 | 0.97167
| Median |  | 0.95 | 1.005 | 0.97 | 0.995 | 0.945 | 0.94787 | 0.935 | 0.5 | 0.85 | 0.975
| Std Dev. |  | 0.020412 | 0.017512 | 0.01169 | 0.052536 | 0.030111 | 0.027201 | 0.012649 | 0 | 0.054772 | 0.0098319
| Rel.Std.Dev. |  | 2.168% | 1.745% | 1.203% | 5.254% | 3.192% | 2.883% | 1.346% | 0% | 6.444% | 1.012%
| PDM3 |  | -2.258% | 4.142% | 0.8555% | 3.796% | -2.085% | -2.079% | -2.431% | -48.1% | -11.77% | 0.8555%
| Z-Score (Absolute) |  | 0.71 | 2.05 | 0.63 | 1.91 | 0.64 | 0.63 | 0.79 | NA | NA | 0.63
Results from laboratories 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 236 (Certified Value 8 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 8.19 | 8.2 | 7.2065 | 7.29 | 9.05 | 7.95 | 7.96 | 8.07 | 8.42 | 9.3467 | 7.79 | 8.08 | 8.0013 | 7.75 | 7.88 | 8.11 | 7.72 | 7.81 | 8.24 | 7.74 | 8.39 | 7.77 |  | 1
| 1 | 2 | 7.87 | 8.31 | 7.187 | 7.66 | 8.63 | 7.97 | 7.96 | 8.06 | 8.43 | 9.1798 | 7.92 | 7.88 | 8.0259 | 7.94 | 8.03 | 8.11 | 7.74 | 7.59 | 8.36 | 7.85 | 8.53 | 8.03 |  | #N/A
| 1 | 3 | 7.67 | 8.3 | 7.2075 | 7.51 | 9.01 | 7.98 | 7.93 | 8.15 | 8.49 | 9.2468 | 7.86 | 7.96 | 8.0738 | 7.85 | 7.83 | 8.08 | 7.91 | 7.49 | 8.37 | 8.05 | 8.31 | 8.21 |  | 16
| 1 | 4 | 7.78 | 8.24 | 7.1345 | 7.46 | 8.83 | 7.99 | 7.87 | 8.2 | 8.49 | 9.2798 | 7.93 | 8.12 | 8.0943 | 7.77 | 7.88 | 8.08 | 7.54 | 7.77 | 8.36 | 8.26 | 8.5 | 8.01 |  | 8.0044
| 1 | 5 | 7.7 | 8.19 | 7.1175 | 7.59 | 8.71 | 7.98 | 8.08 | 8.37 | 8.55 | 9.3467 | 8.08 | 8.13 | 7.9408 | 7.66 | 7.91 | 8.02 | 6.85 | 7.96 | 8.31 | 8.01 | 8.47 | 8 |  | 31
| 1 | 6 | 7.78 | 8.38 | 7.1885 | 7.58 | 8.89 | 8.04 | 7.99 | 8.28 | 8.53 | 9.1798 | 7.81 | 8.41 | 7.9789 | 7.77 | 8.1 | 7.99 | 7.54 | 7.49 | 8.3 | 7.9 | 8.48 | 7.95
| Mean |  | 7.8317 | 8.27 | 7.1736 | 7.515 | 8.8533 | 7.985 | 7.965 | 8.1883 | 8.485 | 9.2633 | 7.8983 | 8.0967 | 8.0192 | 7.79 | 7.9383 | 8.065 | 7.55 | 7.685 | 8.3233 | 7.9683 | 8.4467 | 7.995
| Median |  | 7.78 | 8.27 | 7.1878 | 7.545 | 8.86 | 7.98 | 7.96 | 8.175 | 8.49 | 9.2633 | 7.89 | 8.1 | 8.0136 | 7.77 | 7.895 | 8.08 | 7.63 | 7.68 | 8.335 | 7.955 | 8.475 | 8.005
| Std Dev. |  | 0.18904 | 0.073212 | 0.038234 | 0.13004 | 0.16464 | 0.030166 | 0.069498 | 0.12123 | 0.052058 | 0.075379 | 0.10534 | 0.18206 | 0.057907 | 0.095289 | 0.10381 | 0.049295 | 0.37008 | 0.1916 | 0.050067 | 0.18126 | 0.08165 | 0.14167
| Rel.Std.Dev. |  | 2.414% | 0.8853% | 0.533% | 1.73% | 1.86% | 0.3778% | 0.8725% | 1.481% | 0.6135% | 0.8137% | 1.334% | 2.249% | 0.7221% | 1.223% | 1.308% | 0.6112% | 4.902% | 2.493% | 0.6015% | 2.275% | 0.9666% | 1.772%
| PDM3 |  | -2.158% | 3.318% | -10.38% | -6.114% | 10.61% | -0.2427% | -0.4926% | 2.298% | 6.004% | 15.73% | -1.325% | 1.152% | 0.1841% | -2.679% | -0.8257% | 0.7567% | -5.677% | -3.991% | 3.984% | -0.451% | 5.525% | -0.1178%
| Z-Score (Absolute) |  | 0.54 | 0.95 | 2.76 | 1.61 | 2.92 | 0.02 | 0.08 | 0.67 | 1.68 | 4.31 | 0.31 | 0.36 | 0.1 | 0.68 | 0.17 | 0.25 | 1.49 | 1.03 | 1.13 | 0.07 | 1.55 | 0.02
| Analytical results for Ga in OREAS 236 (Certified Value 16.1 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 16.2 | 16.2 | 18.6197 | 14.7 | 19.5 | 15 | 14.67 | 16.35 | 15.8 | 15.3 | 14.7 | 20.4597 | 17.3 | 13.95 | 13.8 | 14.51 | 17 | 18.99 | 15.75 | 16.92 | 15.45 |  |  | 1
| 1 | 2 | 15.55 | 16.2 | 18.902 | 14.8 | 20 | 15 | 14.57 | 16.56 | 16 | 15.5 | 14.5 | 20.1985 | 16.9 | 14.05 | 14.2 | 14.57 | 18 | 19.26 | 15.35 | 16.94 | 16.05 |  |  | 32
| 1 | 3 | 15.05 | 16.2 | 18.914 | 15.3 | 19.7 | 15 | 14.85 | 16.74 | 16 | 15.7 | 14.6 | 20.3377 | 16.9 | 13.4 | 14.1 | 14.82 | 15 | 19.26 | 16.6 | 16.64 | 16.55 |  |  | 16
| 1 | 4 | 15.45 | 16 | 18.634 | 15.6 | 19.4 | 15 | 14.69 | 17.16 | 15.7 | 16 | 14.8 | 20.0294 | 17.4 | 14.15 | 14.3 | 14.19 | 19 | 19.08 | 15.75 | 17.26 | 15.9 |  |  | 16.1386
| 1 | 5 | 15.3 | 16.2 | 18.366 | 15.4 | 19.6 | 15 | 14.53 | 15.42 | 15.7 | 15.5 | 15 | 21.6924 | 16.4 | 14 | 14.7 | 13.01 | 19 | 19.31 | 16.1 | 17.83 | 15.9 |  |  | 32
| 1 | 6 | 15.5 | 16.4 | 18.954 | 15.6 | 19.7 | 15 | 14.98 | 16.57 | 15.7 | 15.4 | 15 | 20.7082 | 17.3 | 14.65 | 14.5 | 14.12 | 19 | 19.1 | 15.6 | 16.85 | 15.6
| Mean |  | 15.5083 | 16.2 | 18.7316 | 15.2333 | 19.65 | 15 | 14.715 | 16.4667 | 15.8167 | 15.5667 | 14.7667 | 20.571 | 17.0333 | 14.0333 | 14.2667 | 14.2033 | 17.8333 | 19.1667 | 15.8583 | 17.0733 | 15.9083
| Median |  | 15.475 | 16.2 | 18.768 | 15.35 | 19.65 | 15 | 14.68 | 16.565 | 15.75 | 15.5 | 14.75 | 20.3987 | 17.1 | 14.025 | 14.25 | 14.35 | 18.5 | 19.18 | 15.75 | 16.93 | 15.9
| Std Dev. |  | 0.38395 | 0.12649 | 0.23129 | 0.39328 | 0.20736 | 0 | 0.17108 | 0.58033 | 0.1472 | 0.25033 | 0.20656 | 0.59586 | 0.37771 | 0.40083 | 0.31411 | 0.63874 | 1.6021 | 0.12738 | 0.43751 | 0.42103 | 0.38395
| Rel.Std.Dev. |  | 2.476% | 0.7808% | 1.235% | 2.582% | 1.055% | 0% | 1.163% | 3.524% | 0.9306% | 1.608% | 1.399% | 2.897% | 2.217% | 2.856% | 2.202% | 4.497% | 8.984% | 0.6646% | 2.759% | 2.466% | 2.414%
| PDM3 |  | -3.905% | 0.3804% | 16.07% | -5.609% | 21.76% | -7.055% | -8.821% | 2.033% | -1.995% | -3.544% | -8.501% | 27.46% | 5.544% | -13.04% | -11.6% | -11.99% | 10.5% | 18.76% | -1.737% | 5.792% | -1.427%
| Z-Score (Absolute) |  | 0.21 | 0.2 | 1.69 | 0.37 | 2.24 | NA | 0.67 | 0.36 | 0.02 | 0.17 | 0.64 | 2.78 | 0.69 | 1.08 | 0.94 | 0.98 | NA | 1.95 | 0 | 0.72 | 0.03
Results from laboratories 7 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 236 (Certified Value 3.47 ppm) | [24 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 3.6 | 3.62 | 3.46 | 3.77 | 3.2 | 3.2742 | 3.4 | 2.2 | 3.3 | 3.68 | [13 empty] | 1
| 1 | 2 | 3.6 | 3.78 | 3.55 | 3.64 | 3.3 | 3.241 | 3.3 | 2.4 | 3.3 | 3.74 | [13 empty] | 33
| 1 | 3 | 3.6 | 3.83 | 3.48 | 3.88 | 3.1 | 3.3539 | 3.4 | 2.5 | 3.2 | 3.67 | [13 empty] | 16
| 1 | 4 | 3.6 | 3.63 | 3.51 | 3.53 | 3.2 | 3.2972 | 3.5 | 2.4 | 3.1 | 3.66 | [13 empty] | 3.4747
| 1 | 5 | 3.6 | 3.7 | 3.51 | 3.67 | 3.1 | 3.2233 | 3.4 | 2.3 | 2.8 | 3.7 | [13 empty] | 33
| 1 | 6 | 3.6 | 3.79 | 3.44 | 3.79 | 3.5 | 3.3738 | 3.4 | 2.5 | 3.1 | 3.64
| Mean |  | 3.6 | 3.725 | 3.4917 | 3.7133 | 3.2333 | 3.2939 | 3.4 | 2.3833 | 3.1333 | 3.6817
| Median |  | 3.6 | 3.74 | 3.495 | 3.72 | 3.2 | 3.2857 | 3.4 | 2.4 | 3.15 | 3.675
| Std Dev. |  | 0 | 0.088261 | 0.039707 | 0.12469 | 0.15055 | 0.060285 | 0.063246 | 0.1169 | 0.18619 | 0.034881
| Rel.Std.Dev. |  | 0% | 2.369% | 1.137% | 3.358% | 4.656% | 1.83% | 1.86% | 4.905% | 5.942% | 0.9474%
| PDM3 |  | 3.606% | 7.204% | 0.4884% | 6.868% | -6.946% | -5.203% | -2.15% | -31.41% | -9.824% | 5.957%
| Z-Score (Absolute) |  | 0.46 | 0.84 | 0.14 | 0.8 | 0.64 | 0.46 | 0.14 | 3.2 | 0.94 | 0.71
| Analytical results for Ge in OREAS 236 (Indicative Value 0.12 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 0.06 | 0.4 | 0.6 | 3.4 | 1.5 | < 0.05 | 2.4 | 0.1 | 0.2 | 0.06 | 1.4 | < 10 | 4.52 | 0.11 | 0.06 | 0.1 | [7 empty] | 1
| 1 | 2 | 0.07 | 0.4 | 0.6 | 3.3 | 1.6 | 0.08 | 2 | 0.1 | 0.1 | 0.06 | 1.5 | < 10 | 4.64 | 0.08 | 0.06 | 0.08 | [7 empty] | 1
| 1 | 3 | 0.06 | 0.3 | 0.6 | 3.4 | 1.4 | < 0.05 | 2.4 | 0.2 | 0.1 | 0.12 | 1.4 | < 10 | 4.66 | 0.07 | 0.06 | 0.08 | [7 empty] | 16
| 1 | 4 | 0.06 | 0.4 | 0.6 | 3.4 | 1.5 | < 0.05 | 2 | 0.2 | < 0.1 | 0.06 | 1.4 | < 10 | 4.6 | 0.1 | 0.07 | 0.08 | [7 empty] | 0.12259
| 1 | 5 | 0.06 | 0.5 | 0.5 | 3.4 | 1.6 | < 0.05 | 2 | 0.3 | 0.2 | 0.06 | 1.5 | < 10 | 5.03 | 0.07 | 0.06 | 0.06 | [7 empty] | 7
| 1 | 6 | 0.06 | 0.3 | 0.6 | 3.4 | 1.5 | < 0.05 | 2.2 | 0.1 | 0.6 | 0.06 | 1.4 | < 10 | 4.6 | 0.09 | 0.05 | 0.07
| Mean |  | 0.061667 | 0.38333 | 0.58333 | 3.3833 | 1.5167 | 0.08 | 2.1667 | 0.16667 | 0.24 | 0.07 | 1.4333 |  | 4.675 | 0.086667 | 0.06 | 0.078333
| Median |  | 0.06 | 0.4 | 0.6 | 3.4 | 1.5 | 0.08 | 2.1 | 0.15 | 0.2 | 0.06 | 1.4 |  | 4.62 | 0.085 | 0.06 | 0.08
| Std Dev. |  | 0.0040825 | 0.075277 | 0.040825 | 0.040825 | 0.075277 |  | 0.19664 | 0.08165 | 0.20736 | 0.024495 | 0.05164 |  | 0.18042 | 0.01633 | 0.0063246 | 0.013292
| Rel.Std.Dev. |  | 6.62% | 19.64% | 6.999% | 1.207% | 4.963% |  | 9.076% | 48.99% | 86.4% | 34.99% | 3.603% |  | 3.859% | 18.84% | 10.54% | 16.97%
| PDM3 |  | -49.7% | 212.7% | 375.8% | 2660% | 1137% | -34.74% | 1667% | 35.95% | 95.77% | -42.9% | 1069% |  | 3713% | -29.31% | -51.06% | -36.1%
| Z-Score (Absolute) |  | 0.63 | 0.24 | 0.78 | 8.37 | 3.31 | 0.71 | 5.07 | 0.35 | 0.24 | 0.61 | 3.08 | 12.75 | 11.87 | 0.57 | 0.64 | 0.59
| Analytical results for Hf in OREAS 236 (Certified Value 1.8 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.9 | 2 | 2.78 | 1.72 | 2.4 | 1.7 | 1.85 | 1.8 | 2 | 1 | 1.776 | 1.8 | 1.7 | 1.6 | 1.54 | 2.1 | 1.9 | 1.8 | 1.8 |  |  |  |  | 1
| 1 | 2 | 1.8 | 1.8 | 2.6 | 1.83 | 2.5 | 1.72 | 1.87 | 1.8 | 1.8 | 1 | 1.7877 | 1.76 | 1.6 | 1.6 | 1.56 | 2.1 | 1.9 | 1.7 | 1.8 |  |  |  |  | 18
| 1 | 3 | 1.7 | 1.8 | 2.67 | 1.84 | 2.5 | 1.78 | 1.89 | 1.8 | 1.8 | 1.5 | 1.8053 | 1.79 | 1.8 | 1.6 | 1.56 | 2.1 | 1.9 | 1.8 | 1.8 |  |  |  |  | 16
| 1 | 4 | 1.8 | 1.8 | 2.56 | 1.77 | 2.4 | 1.7 | 1.81 | 1.7 | 1.9 | 1.5 | 1.8515 | 1.82 | 1.7 | 1.6 | 1.46 | 2.1 | 1.8 | 1.8 | 1.8 |  |  |  |  | 1.8009
| 1 | 5 | 1.8 | 1.8 | 2.86 | 1.78 | 2.2 | 1.72 | 1.84 | 1.8 | 2 | 1.7 | 1.7995 | 1.86 | 1.6 | 1.5 | 1.36 | 2.2 | 1.8 | 1.8 | 1.8 |  |  |  |  | 34
| 1 | 6 | 1.7 | 2 | 2.38 | 1.83 | 2.4 | 1.8 | 1.92 | 1.7 | 1.8 | 1.8 | 1.8229 | 1.83 | 1.6 | 1.6 | 1.48 | 2.1 | 1.8 | 1.7 | 1.9
| Mean |  | 1.7833 | 1.8667 | 2.6417 | 1.795 | 2.4 | 1.7367 | 1.8633 | 1.7667 | 1.8833 | 1.4167 | 1.8072 | 1.81 | 1.6667 | 1.5833 | 1.4933 | 2.1167 | 1.85 | 1.7667 | 1.8167
| Median |  | 1.8 | 1.8 | 2.635 | 1.805 | 2.4 | 1.72 | 1.86 | 1.8 | 1.85 | 1.5 | 1.8024 | 1.81 | 1.65 | 1.6 | 1.51 | 2.1 | 1.85 | 1.8 | 1.8
| Std Dev. |  | 0.075277 | 0.10328 | 0.16999 | 0.046797 | 0.10954 | 0.04274 | 0.038816 | 0.05164 | 0.098319 | 0.34303 | 0.026942 | 0.034641 | 0.08165 | 0.040825 | 0.077632 | 0.040825 | 0.054772 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 4.221% | 5.533% | 6.435% | 2.607% | 4.564% | 2.461% | 2.083% | 2.923% | 5.22% | 24.21% | 1.491% | 1.914% | 4.899% | 2.578% | 5.199% | 1.929% | 2.961% | 2.923% | 2.247%
| PDM3 |  | -0.9774% | 3.65% | 46.68% | -0.3296% | 33.26% | -3.569% | 3.465% | -1.903% | 4.575% | -21.34% | 0.346% | 0.5033% | -7.456% | -12.08% | -17.08% | 17.53% | 2.724% | -1.903% | 0.8735%
| Z-Score (Absolute) |  | 0.27 | 0.67 | 9.46 | 0.14 | 6.72 | 0.8 | 0.64 | 0.46 | 0.86 | 4.43 | 0 | 0.03 | 1.59 | 2.54 | 3.56 | 3.51 | 0.49 | 0.46 | 0.11
| Analytical results for Hg in OREAS 236 (Indicative Value < 2 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | Lab | [22 empty] | 1
| No | No | 1.24 | [22 empty] | ppm
|  |  | 4A*OES/MS | [22 empty] | 2
2
| 1 | 1 | < 2 | [22 empty] | 1
| 1 | 2 | < 2 | [22 empty] | 1
| 1 | 3 | < 2 | [22 empty] | 16
| 1 | 4 | < 2 | [22 empty] | < 2
| 1 | 5 | < 2 | [22 empty] | 7
| 1 | 6 | < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 236 (Certified Value 0.86 ppm) | [24 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.84 | 0.84 | 0.85 | 0.92 | 0.8 | 0.78176 | 0.83 | 0.6 | 0.7 | 0.9 | [13 empty] | 1
| 1 | 2 | 0.84 | 0.87 | 0.84 | 0.87 | 0.8 | 0.77043 | 0.81 | 0.5 | 0.8 | 0.9 | [13 empty] | 19
| 1 | 3 | 0.84 | 0.87 | 0.85 | 0.91 | 0.8 | 0.75028 | 0.84 | 0.6 | 0.8 | 0.91 | [13 empty] | 16
| 1 | 4 | 0.82 | 0.85 | 0.83 | 0.89 | 0.9 | 0.75181 | 0.85 | 0.6 | 0.7 | 0.91 | [13 empty] | 0.85917
| 1 | 5 | 0.82 | 0.83 | 0.86 | 0.85 | 0.9 | 0.69492 | 0.81 | 0.6 | 0.7 | 0.91 | [13 empty] | 35
| 1 | 6 | 0.84 | 0.87 | 0.86 | 0.88 | 0.9 | 0.78353 | 0.82 | 0.6 | 0.7 | 0.9
| Mean |  | 0.83333 | 0.855 | 0.84833 | 0.88667 | 0.85 | 0.75546 | 0.82667 | 0.58333 | 0.73333 | 0.905
| Median |  | 0.84 | 0.86 | 0.85 | 0.885 | 0.85 | 0.76112 | 0.825 | 0.6 | 0.7 | 0.905
| Std Dev. |  | 0.010328 | 0.017607 | 0.01169 | 0.02582 | 0.054772 | 0.03289 | 0.01633 | 0.040825 | 0.05164 | 0.0054772
| Rel.Std.Dev. |  | 1.239% | 2.059% | 1.378% | 2.912% | 6.444% | 4.354% | 1.975% | 6.999% | 7.042% | 0.6052%
| PDM3 |  | -3.007% | -0.485% | -1.261% | 3.201% | -1.067% | -12.07% | -3.783% | -32.1% | -14.65% | 5.335%
| Z-Score (Absolute) |  | 0.47 | 0.21 | 0 | 1.19 | NA | 2.89 | 0.67 | NA | NA | 1.76
Results from laboratories 14, 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 236 (Certified Value 0.08 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.087 | 0.1 | < 2 | 0.08 | 0.08 | 0.09 | 0.08 | 0.094 | 0.09 | 0.08 | < 0.1 | 0.09 | 0.068 | 0.07 | 0.08 | 11 | 0.071 | 0.08 | 0.081 | 0.075 |  |  |  | 1
| 1 | 2 | 0.081 | 0.08 | < 2 | 0.08 | 0.08 | 0.08 | 0.07 | 0.088 | 0.07 | 0.08 | < 0.1 | 0.08 | 0.07 | 0.07 | 0.08 | 17 | 0.071 | 0.079 | 0.076 | 0.076 |  |  |  | 20
| 1 | 3 | 0.079 | 0.1 | < 2 | 0.08 | 0.08 | 0.09 | 0.08 | 0.131 | 0.09 | 0.08 | < 0.1 | 0.09 | 0.073 | 0.07 | 0.08 | 16 | 0.072 | 0.087 | 0.078 | 0.082 |  |  |  | 16
| 1 | 4 | 0.084 | 0.1 | < 2 | 0.08 | 0.08 | 0.1 | 0.08 | 0.099 | 0.08 | 0.07 | < 0.1 | 0.09 | 0.071 | 0.07 | 0.07 | 16 | 0.072 | 0.082 | 0.083 | 0.077 |  |  |  | 0.079708
| 1 | 5 | 0.08 | 0.1 | < 2 | 0.08 | 0.08 | 0.09 | 0.07 | 0.098 | 0.09 | 0.07 | < 0.1 | 0.09 | 0.073 | 0.07 | 0.07 | 17 | 0.071 | 0.077 | 0.085 | 0.081 |  |  |  | 36
| 1 | 6 | 0.081 | 0.1 | < 2 | 0.08 | 0.08 | 0.08 | 0.08 | 0.116 | 0.07 | 0.07 | < 0.1 | 0.09 | 0.073 | 0.07 | 0.07 | 18 | 0.072 | 0.076 | 0.08 | 0.078
| Mean |  | 0.082 | 0.096667 |  | 0.08 | 0.08 | 0.088333 | 0.076667 | 0.10433 | 0.081667 | 0.075 |  | 0.088333 | 0.071333 | 0.07 | 0.075 | 15.8333 | 0.0715 | 0.080167 | 0.0805 | 0.078167
| Median |  | 0.081 | 0.1 |  | 0.08 | 0.08 | 0.09 | 0.08 | 0.0985 | 0.085 | 0.075 |  | 0.09 | 0.072 | 0.07 | 0.075 | 16.5 | 0.0715 | 0.0795 | 0.0805 | 0.0775
| Std Dev. |  | 0.0029665 | 0.008165 |  | 0 | 0 | 0.0075277 | 0.005164 | 0.016058 | 0.0098319 | 0.0054772 |  | 0.0040825 | 0.0020656 | 0 | 0.0054772 | 2.4833 | 0.00054772 | 0.0039707 | 0.0032711 | 0.0027869
| Rel.Std.Dev. |  | 3.618% | 8.447% |  | 0% | 0% | 8.522% | 6.736% | 15.39% | 12.04% | 7.303% |  | 4.622% | 2.896% | 0% | 7.303% | 15.68% | 0.766% | 4.953% | 4.063% | 3.565%
| PDM3 |  | 2.875% | 21.28% |  | 0.3659% | 0.3659% | 10.82% | -3.816% | 30.89% | 2.457% | -5.907% |  | 10.82% | -10.51% | -12.18% | -5.907% | 19764% | -10.3% | 0.575% | 0.9932% | -1.934%
| Z-Score (Absolute) |  | 0.27 | 2.25 | 124.07 | 0 | 0 | 1.12 | 0.45 | 3.28 | 0.22 | 0.67 | 4.05 | 1.12 | 1.17 | 1.35 | 0.67 | NA | 1.15 | 0.02 | 0.07 | 0.25
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 236 (Certified Value 0.559 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.57 | 0.55 | 0.816 | 0.57 | 0.69 | 0.6 | 0.5566 | 0.54 | 0.565 | 0.68884 | 0.5496 | 0.56 | 0.57285 | 0.58 | 0.54 | 0.555 | 0.54 | 0.56 | 0.54 | 0.53 | 0.6 | 0.53 |  | 1
| 1 | 2 | 0.55 | 0.57 | 0.836 | 0.57 | 0.65 | 0.6 | 0.5557 | 0.55 | 0.559 | 0.68594 | 0.5609 | 0.57 | 0.57567 | 0.59 | 0.55 | 0.542 | 0.54 | 0.56 | 0.55 | 0.54 | 0.62 | 0.55 |  | #N/A
| 1 | 3 | 0.53 | 0.56 | 0.824 | 0.55 | 0.67 | 0.61 | 0.5591 | 0.55 | 0.559 | 0.68887 | 0.5588 | 0.55 | 0.58033 | 0.59 | 0.53 | 0.552 | 0.56 | 0.55 | 0.55 | 0.55 | 0.6 | 0.56 |  | 16
| 1 | 4 | 0.54 | 0.55 | 0.816 | 0.57 | 0.66 | 0.61 | 0.5551 | 0.55 | 0.562 | 0.68694 | 0.5648 | 0.56 | 0.5793 | 0.58 | 0.53 | 0.551 | 0.53 | 0.56 | 0.55 | 0.56 | 0.61 | 0.54 |  | 0.55896
| 1 | 5 | 0.53 | 0.56 | 0.82133 | 0.56 | 0.65 | 0.61 | 0.5551 | 0.56 | 0.565 | 0.68847 | 0.5521 | 0.57 | 0.57292 | 0.61 | 0.53 | 0.548 | 0.49 | 0.56 | 0.57 | 0.55 | 0.61 | 0.54 |  | 37
| 1 | 6 | 0.53 | 0.56 | 0.832 | 0.55 | 0.67 | 0.6 | 0.5557 | 0.57 | 0.567 | 0.68638 | 0.5512 | 0.6 | 0.58707 | 0.59 | 0.55 | 0.548 | 0.54 | 0.56 | 0.55 | 0.54 | 0.61 | 0.55
| Mean |  | 0.54167 | 0.55833 | 0.82422 | 0.56167 | 0.665 | 0.605 | 0.55622 | 0.55333 | 0.56283 | 0.68757 | 0.55623 | 0.56833 | 0.57802 | 0.59 | 0.53833 | 0.54933 | 0.53333 | 0.55833 | 0.55167 | 0.545 | 0.60833 | 0.545
| Median |  | 0.535 | 0.56 | 0.82267 | 0.565 | 0.665 | 0.605 | 0.5557 | 0.55 | 0.5635 | 0.6877 | 0.55545 | 0.565 | 0.57749 | 0.59 | 0.535 | 0.5495 | 0.54 | 0.56 | 0.55 | 0.545 | 0.61 | 0.545
| Std Dev. |  | 0.016021 | 0.0075277 | 0.0082803 | 0.0098319 | 0.015166 | 0.0054772 | 0.0015158 | 0.010328 | 0.0033714 | 0.0013098 | 0.0061347 | 0.017224 | 0.0054247 | 0.010954 | 0.0098319 | 0.0044572 | 0.023381 | 0.0040825 | 0.0098319 | 0.010488 | 0.0075277 | 0.010488
| Rel.Std.Dev. |  | 2.958% | 1.348% | 1.005% | 1.75% | 2.281% | 0.9053% | 0.2725% | 1.866% | 0.599% | 0.1905% | 1.103% | 3.031% | 0.9385% | 1.857% | 1.826% | 0.8114% | 4.384% | 0.7312% | 1.782% | 1.924% | 1.237% | 1.924%
| PDM3 |  | -3.094% | -0.1124% | 47.46% | 0.484% | 18.97% | 8.236% | -0.4911% | -1.007% | 0.6927% | 23.01% | -0.4881% | 1.677% | 3.41% | 5.553% | -3.69% | -1.723% | -4.585% | -0.1124% | -1.305% | -2.498% | 8.833% | -2.498%
| Z-Score (Absolute) |  | 0.84 | 0 | 13.45 | 0.17 | 5.39 | 2.36 | 0.11 | 0.25 | 0.23 | 6.54 | 0.11 | 0.51 | 1 | 1.6 | 1.01 | 0.46 | 1.26 | 0 | 0.34 | 0.67 | 2.53 | 0.67
| Analytical results for La in OREAS 236 (Certified Value 5.99 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 6.4 | 6.1 | 7.665 | 6.2 | 6.2 | 6 | 5.89 | 6.2 | < 10 | 5 | 6.8541 | 5.2 | 4.5 | 5.8 | 5.6 | 7 | 6.3 | 6.2 | 6 | 5.5 |  |  |  | 1
| 1 | 2 | 6.2 | 6.3 | 7.889 | 6.5 | 6.3 | 6 | 6.03 | 6 | < 10 | 4.9 | 6.7064 | 5.3 | 4.8 | 5.9 | 5.6 | 8 | 6.4 | 6.2 | 5.9 | 5.7 |  |  |  | 21
| 1 | 3 | 6 | 6.2 | 7.7477 | 6.4 | 6.5 | 6 | 5.96 | 6 | < 10 | 5 | 6.856 | 5.4 | 5 | 5.1 | 5.7 | 7 | 6.4 | 6.5 | 5.9 | 6 |  |  |  | 16
| 1 | 4 | 6.1 | 6.2 | 7.7023 | 6.6 | 6.2 | 6 | 5.92 | 5.9 | < 10 | 5 | 6.7318 | 5.4 | 4.7 | 6 | 5.2 | 7 | 6.3 | 6.4 | 6.1 | 5.8 |  |  |  | 5.9884
| 1 | 5 | 6 | 6.2 | 7.6517 | 6.2 | 6.1 | 6 | 5.98 | 5.9 | < 10 | 4.9 | 6.6328 | 5.5 | 4.7 | 6 | 4.9 | 7 | 6.4 | 6.4 | 6.3 | 5.7 |  |  |  | 38
| 1 | 6 | 6.1 | 6.3 | 7.621 | 6.5 | 6.4 | 6 | 5.99 | 6 | < 10 | 6.3 | 6.7231 | 5.3 | 5.1 | 5.9 | 5.4 | 7 | 6.4 | 6.2 | 6 | 5.7
| Mean |  | 6.1333 | 6.2167 | 7.7128 | 6.4 | 6.2833 | 6 | 5.9617 | 6 |  | 5.1833 | 6.7507 | 5.35 | 4.8 | 5.7833 | 5.4 | 7.1667 | 6.3667 | 6.3167 | 6.0333 | 5.7333
| Median |  | 6.1 | 6.2 | 7.6836 | 6.45 | 6.25 | 6 | 5.97 | 6 |  | 5 | 6.7275 | 5.35 | 4.75 | 5.9 | 5.5 | 7 | 6.4 | 6.3 | 6 | 5.7
| Std Dev. |  | 0.15055 | 0.075277 | 0.096769 | 0.16733 | 0.1472 | 0 | 0.050365 | 0.10954 |  | 0.54924 | 0.088043 | 0.10488 | 0.21909 | 0.34303 | 0.30332 | 0.40825 | 0.05164 | 0.13292 | 0.15055 | 0.1633
| Rel.Std.Dev. |  | 2.455% | 1.211% | 1.255% | 2.615% | 2.343% | 0% | 0.8448% | 1.826% |  | 10.6% | 1.304% | 1.96% | 4.564% | 5.931% | 5.617% | 5.696% | 0.8111% | 2.104% | 2.495% | 2.848%
| PDM3 |  | 2.421% | 3.812% | 28.8% | 6.874% | 4.925% | 0.194% | -0.4461% | 0.194% |  | -13.44% | 12.73% | -10.66% | -19.84% | -3.424% | -9.825% | 19.68% | 6.317% | 5.482% | 0.7507% | -4.259%
| Z-Score (Absolute) |  | 0.24 | 0.41 | 3.52 | 0.8 | 0.55 | NA | 0.11 | 0.03 | 2.11 | 1.73 | 1.52 | 1.38 | 2.52 | 0.48 | 1.28 | NA | 0.73 | 0.62 | 0.03 | 0.59
Results from laboratories 7 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 236 (Certified Value 10.5 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 10.7 | 11.5 | 13.342 | 12 | 14.7 | 11 | 11.2 | 10.5 | 10.35 | 10.4 | 10.6 | 10.7109 | 11 | 10.3 | 9.2 | 9.9 | 9.8 | 11 | 10.9 | 10.1 | 11.3 | 10.5 |  | 1
| 1 | 2 | 10.3 | 11.5 | 13.067 | 12 | 13.2 | 11.1 | 11.3 | 10.4 | 11.78 | 10.6 | 10.3 | 10.4795 | 12 | 10.2 | 9.5 | 9.9 | 9.9 | 10 | 10.8 | 9.9 | 11.2 | 11 |  | #N/A
| 1 | 3 | 10.2 | 11.5 | 13.3048 | 12 | 14.3 | 10.7 | 11.8 | 10.3 | 10.44 | 10.4 | 10.4 | 10.5786 | 12 | 9.7 | 9.5 | 9.9 | 10.2 | 10 | 10.8 | 10.2 | 11 | 11.2 |  | 16
| 1 | 4 | 10.4 | 11.5 | 13.11 | 12 | 13 | 10.7 | 11 | 10.3 | 10.53 | 10.5 | 10.5 | 10.349 | 12 | 9.8 | 9.7 | 9.8 | 9.6 | 10 | 10.6 | 10.3 | 11.3 | 10.8 |  | 10.5106
| 1 | 5 | 10.3 | 11.5 | 13.01 | 12 | 13.5 | 10.9 | 11.7 | 10.5 | 10.67 | 10.4 | 10.7 | 9.5093 | 12 | 9.6 | 9.6 | 9.8 | 8.9 | 10 | 10.9 | 10.2 | 11.6 | 10 |  | 39
| 1 | 6 | 10.4 | 11.5 | 13.3085 | 11 | 12.7 | 11.2 | 11.4 | 10.3 | 11.25 | 10.2 | 10.9 | 10.4634 | 12 | 9.6 | 9.7 | 10.3 | 9.7 | 11 | 10.6 | 10.1 | 11 | 10.5
| Mean |  | 10.3833 | 11.5 | 13.1904 | 11.8333 | 13.5667 | 10.9333 | 11.4 | 10.3833 | 10.8367 | 10.4167 | 10.5667 | 10.3485 | 11.8333 | 9.8667 | 9.5333 | 9.9333 | 9.6833 | 10.3333 | 10.7667 | 10.1333 | 11.2333 | 10.6667
| Median |  | 10.35 | 11.5 | 13.2074 | 12 | 13.35 | 10.95 | 11.35 | 10.35 | 10.6 | 10.4 | 10.55 | 10.4715 | 12 | 9.75 | 9.55 | 9.9 | 9.75 | 10 | 10.8 | 10.15 | 11.25 | 10.65
| Std Dev. |  | 0.17224 | 0 | 0.1444 | 0.40825 | 0.77889 | 0.20656 | 0.30332 | 0.098319 | 0.56156 | 0.13292 | 0.21602 | 0.42871 | 0.40825 | 0.30768 | 0.18619 | 0.18619 | 0.43551 | 0.5164 | 0.13663 | 0.13663 | 0.22509 | 0.4274
| Rel.Std.Dev. |  | 1.659% | 0% | 1.095% | 3.45% | 5.741% | 1.889% | 2.661% | 0.9469% | 5.182% | 1.276% | 2.044% | 4.143% | 3.45% | 3.118% | 1.953% | 1.874% | 4.497% | 4.997% | 1.269% | 1.348% | 2.004% | 4.007%
| PDM3 |  | -1.211% | 9.414% | 25.5% | 12.59% | 29.08% | 4.022% | 8.462% | -1.211% | 3.103% | -0.8934% | 0.5338% | -1.542% | 12.59% | -6.126% | -9.298% | -5.492% | -7.87% | -1.686% | 2.437% | -3.589% | 6.877% | 1.485%
| Z-Score (Absolute) |  | 0.29 | 1.45 | 4.08 | NA | 4.67 | 0.57 | 1.3 | 0.29 | 0.42 | 0.23 | 0 | 0.34 | NA | 1.09 | 1.61 | 0.99 | 1.37 | NA | 0.31 | 0.67 | 1.04 | 0.16
Results from laboratories 4, 16 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 236 (Certified Value 0.37 ppm) | [24 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.36 | 0.33 | 0.35 | 0.35 | 0.39 | 0.3 | 0.43869 | 0.35 | 0.3 | 0.3 | 0.38 | [12 empty] | 1
| 1 | 2 | 0.38 | 0.35 | 0.37 | 0.34 | 0.38 | 0.4 | 0.42114 | 0.35 | 0.4 | 0.3 | 0.39 | [12 empty] | 22
| 1 | 3 | 0.36 | 0.34 | 0.37 | 0.35 | 0.39 | 0.3 | 0.42908 | 0.36 | 0.4 | 0.3 | 0.39 | [12 empty] | 16
| 1 | 4 | 0.36 | 0.36 | 0.35 | 0.33 | 0.39 | 0.3 | 0.42842 | 0.34 | 0.3 | 0.3 | 0.38 | [12 empty] | 0.37163
| 1 | 5 | 0.38 | 0.34 | 0.34 | 0.37 | 0.37 | 0.3 | 0.39531 | 0.35 | 0.4 | 0.3 | 0.39 | [12 empty] | 40
| 1 | 6 | 0.38 | 0.34 | 0.37 | 0.35 | 0.4 | 0.3 | 0.43967 | 0.35 | 0.4 | 0.3 | 0.38
| Mean |  | 0.37 | 0.34333 | 0.35833 | 0.34833 | 0.38667 | 0.31667 | 0.42538 | 0.35 | 0.36667 | 0.3 | 0.385
| Median |  | 0.37 | 0.34 | 0.36 | 0.35 | 0.39 | 0.3 | 0.42875 | 0.35 | 0.4 | 0.3 | 0.385
| Std Dev. |  | 0.010954 | 0.010328 | 0.013292 | 0.013292 | 0.010328 | 0.040825 | 0.01629 | 0.0063246 | 0.05164 | 0 | 0.0054772
| Rel.Std.Dev. |  | 2.961% | 3.008% | 3.709% | 3.816% | 2.671% | 12.89% | 3.829% | 1.807% | 14.08% | 0% | 1.423%
| PDM3 |  | -0.4395% | -7.615% | -3.579% | -6.27% | 4.045% | -14.79% | 14.46% | -5.821% | -1.336% | -19.28% | 3.597%
| Z-Score (Absolute) |  | 0.21 | 0.77 | 0.21 | 0.58 | 0.83 | NA | 2.25 | 0.52 | NA | NA | 0.77
Results from laboratories 14, 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 236 (Certified Value 3.4 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.48 | 3.34 | 2.812 | 3 | 4.04 | 3.42 | 3.3037 | 3.35 | 3.55 | 3.729 | 3.3275 | 3.52 | 3.4772 | 3.51 | 3.37 | 3.3 | 3.36 | 3.4 | 3.35 | 3.3 | 3.48 | 3.25 |  | 1
| 1 | 2 | 3.33 | 3.37 | 2.855 | 3.15 | 3.87 | 3.46 | 3.2893 | 3.41 | 3.55 | 3.7199 | 3.3853 | 3.49 | 3.4751 | 3.51 | 3.41 | 3.32 | 3.44 | 3.5 | 3.31 | 3.36 | 3.54 | 3.37 |  | #N/A
| 1 | 3 | 3.26 | 3.46 | 2.867 | 3.08 | 4.01 | 3.46 | 3.2974 | 3.38 | 3.575 | 3.7285 | 3.3728 | 3.53 | 3.4929 | 3.52 | 3.3 | 3.34 | 3.42 | 3.41 | 3.33 | 3.42 | 3.45 | 3.44 |  | 16
| 1 | 4 | 3.3 | 3.43 | 2.8505 | 3.15 | 3.91 | 3.47 | 3.2984 | 3.42 | 3.577 | 3.7215 | 3.3953 | 3.56 | 3.4943 | 3.5 | 3.37 | 3.34 | 3.32 | 3.37 | 3.33 | 3.53 | 3.52 | 3.35 |  | 3.3955
| 1 | 5 | 3.27 | 3.35 | 2.857 | 3.18 | 3.89 | 3.46 | 3.322 | 3.5 | 3.557 | 3.7291 | 3.324 | 3.54 | 3.4907 | 3.43 | 3.37 | 3.37 | 3.03 | 3.44 | 3.25 | 3.42 | 3.52 | 3.35 |  | 41
| 1 | 6 | 3.28 | 3.4 | 2.874 | 3.08 | 3.95 | 3.42 | 3.3656 | 3.58 | 3.544 | 3.7256 | 3.3306 | 3.71 | 3.4706 | 3.5 | 3.46 | 3.36 | 3.37 | 3.4 | 3.31 | 3.37 | 3.51 | 3.33
| Mean |  | 3.32 | 3.3917 | 2.8526 | 3.1067 | 3.945 | 3.4483 | 3.3127 | 3.44 | 3.5588 | 3.7256 | 3.3559 | 3.5583 | 3.4835 | 3.495 | 3.38 | 3.3383 | 3.3233 | 3.42 | 3.3133 | 3.4 | 3.5033 | 3.3483
| Median |  | 3.29 | 3.385 | 2.856 | 3.115 | 3.93 | 3.46 | 3.3011 | 3.415 | 3.5535 | 3.727 | 3.3517 | 3.535 | 3.484 | 3.505 | 3.37 | 3.34 | 3.365 | 3.405 | 3.32 | 3.395 | 3.515 | 3.35
| Std Dev. |  | 0.082219 | 0.047081 | 0.021648 | 0.066232 | 0.068044 | 0.022286 | 0.028114 | 0.085088 | 0.013934 | 0.0040403 | 0.032145 | 0.077825 | 0.010315 | 0.032711 | 0.052915 | 0.025626 | 0.15002 | 0.045166 | 0.034448 | 0.077717 | 0.03266 | 0.061455
| Rel.Std.Dev. |  | 2.476% | 1.388% | 0.7589% | 2.132% | 1.725% | 0.6463% | 0.8487% | 2.473% | 0.3915% | 0.1084% | 0.9579% | 2.187% | 0.2961% | 0.9359% | 1.566% | 0.7676% | 4.514% | 1.321% | 1.04% | 2.286% | 0.9323% | 1.835%
| PDM3 |  | -2.225% | -0.1141% | -15.99% | -8.507% | 16.18% | 1.555% | -2.439% | 1.309% | 4.809% | 9.72% | -1.167% | 4.794% | 2.59% | 2.929% | -0.4577% | -1.685% | -2.127% | 0.7203% | -2.421% | 0.1313% | 3.174% | -1.39%
| Z-Score (Absolute) |  | 0.65 | 0.04 | 4.63 | 2.46 | 4.68 | 0.45 | 0.71 | 0.38 | 1.39 | 2.81 | 0.34 | 1.38 | 0.75 | 0.84 | 0.13 | 0.49 | 0.62 | 0.21 | 0.7 | 0.04 | 0.92 | 0.4
| Analytical results for Mn in OREAS 236 (Certified Value 0.133 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.134 | 0.133 | 0.1067 | 0.1419 | 0.14 | 0.1327 | 0.1358 | 0.1294 | 0.1498 | 0.15814 | 0.1324 | 0.128 | 0.13265 | 0.1227 | 0.131 | 0.127 | 0.1377 | 0.1208 | 0.1397 | 0.129 | 0.1299 | 0.127 |  | 1
| 1 | 2 | 0.128 | 0.134 | 0.10708 | 0.1398 | 0.135 | 0.1345 | 0.1366 | 0.1319 | 0.1521 | 0.15511 | 0.1354 | 0.126 | 0.13301 | 0.1261 | 0.131 | 0.127 | 0.142 | 0.1229 | 0.1417 | 0.131 | 0.1326 | 0.131 |  | 23
| 1 | 3 | 0.125 | 0.136 | 0.10699 | 0.1408 | 0.141 | 0.1359 | 0.1349 | 0.1307 | 0.1517 | 0.15667 | 0.1341 | 0.126 | 0.13273 | 0.126 | 0.125 | 0.129 | 0.144 | 0.1197 | 0.1419 | 0.132 | 0.1293 | 0.133 |  | 16
| 1 | 4 | 0.126 | 0.134 | 0.10743 | 0.1428 | 0.137 | 0.1353 | 0.134 | 0.1322 | 0.1481 | 0.15678 | 0.1345 | 0.128 | 0.13247 | 0.1231 | 0.133 | 0.128 | 0.1385 | 0.1197 | 0.1414 | 0.138 | 0.1318 | 0.13 |  | 0.13325
| 1 | 5 | 0.126 | 0.134 | 0.10718 | 0.1406 | 0.134 | 0.137 | 0.1354 | 0.1347 | 0.1509 | 0.15718 | 0.1319 | 0.13 | 0.13245 | 0.1228 | 0.132 | 0.129 | 0.1248 | 0.1229 | 0.1414 | 0.133 | 0.1314 | 0.129 |  | 42
| 1 | 6 | 0.127 | 0.137 | 0.10774 | 0.14 | 0.14 | 0.1377 | 0.1359 | 0.132 | 0.1477 | 0.15712 | 0.1327 | 0.142 | 0.13493 | 0.1233 | 0.135 | 0.128 | 0.1363 | 0.1197 | 0.1409 | 0.13 | 0.1315 | 0.13
| Mean |  | 0.12767 | 0.13467 | 0.10719 | 0.14098 | 0.13783 | 0.13552 | 0.13543 | 0.13182 | 0.15005 | 0.15684 | 0.1335 | 0.13 | 0.13304 | 0.124 | 0.13117 | 0.128 | 0.13722 | 0.12095 | 0.14117 | 0.13217 | 0.13108 | 0.13
| Median |  | 0.1265 | 0.134 | 0.10713 | 0.1407 | 0.1385 | 0.1356 | 0.1356 | 0.13195 | 0.15035 | 0.15695 | 0.1334 | 0.128 | 0.13269 | 0.1232 | 0.1315 | 0.128 | 0.1381 | 0.12025 | 0.1414 | 0.1315 | 0.13145 | 0.13
| Std Dev. |  | 0.003266 | 0.0015055 | 0.00036138 | 0.0011566 | 0.0029269 | 0.001796 | 0.00090037 | 0.0017657 | 0.0018458 | 0.00098895 | 0.0013697 | 0.0060663 | 0.0009469 | 0.0016025 | 0.0033714 | 0.00089443 | 0.0067199 | 0.0015694 | 0.00079415 | 0.0031885 | 0.0012384 | 0.002
| Rel.Std.Dev. |  | 2.558% | 1.118% | 0.3371% | 0.8204% | 2.123% | 1.325% | 0.6648% | 1.34% | 1.23% | 0.6306% | 1.026% | 4.666% | 0.7117% | 1.292% | 2.57% | 0.6988% | 4.897% | 1.298% | 0.5626% | 2.413% | 0.9448% | 1.538%
| PDM3 |  | -4.193% | 1.06% | -19.56% | 5.801% | 3.437% | 1.698% | 1.636% | -1.078% | 12.6% | 17.7% | 0.1848% | -2.442% | -0.16% | -6.944% | -1.566% | -3.943% | 2.974% | -9.233% | 5.938% | -0.8158% | -1.629% | -2.442%
| Z-Score (Absolute) |  | 0.89 | 0.37 | 4.56 | 1.5 | 0.94 | 0.52 | 0.51 | 0.14 | 3.13 | 4.35 | 0.16 | 0.47 | 0.08 | 1.54 | 0.26 | 0.83 | 0.83 | 2.09 | 1.54 | 0.08 | 0.27 | 0.47
| Analytical results for Mo in OREAS 236 (Certified Value 1.56 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.51 | 1.5 | < 2 | 1.61 | 1.72 | 1.7 | 1.4 | 1.51 | 0.8 | < 5 | 1.5 | 0.21 | 1.97 | 2 | 1.41 | 1.4 | 1.36 | < 2 | 1.09 | 1.47 | 1.73 | 1.46 |  | 1
| 1 | 2 | 1.48 | 1.5 | < 2 | 1.6 | 1.72 | 1.8 | 1.6 | 1.55 | 1.2 | < 5 | 1.5 | 0.13 | 2.01 | 1.8 | 1.33 | 1.4 | 1.41 | 2 | 1.17 | 1.51 | 1.74 | 1.46 |  | 24
| 1 | 3 | 1.47 | 1.5 | < 2 | 1.6 | 1.76 | 1.9 | 1.5 | 1.62 | 1.2 | < 5 | 1.5 | 0.15 | 1.88 | 1.8 | 1.4 | 1.5 | 1.44 | < 2 | 1.15 | 1.73 | 1.65 | 1.55 |  | 16
| 1 | 4 | 1.47 | 1.5 | < 2 | 1.61 | 1.76 | 1.8 | 1.5 | 1.66 | 0.9 | < 5 | 1.5 | 0.16 | 1.9 | 1.8 | 1.45 | 1.5 | 1.37 | < 2 | 1.11 | 1.56 | 1.7 | 1.48 |  | 1.5601
| 1 | 5 | 1.44 | 1.5 | < 2 | 1.71 | 1.75 | 1.7 | 1.5 | 1.56 | 1.1 | < 5 | 1.5 | 0.27 | 1.86 | 1.8 | 1.39 | 1.5 | 1.29 | < 2 | 1.23 | 1.61 | 1.66 | 1.5 |  | 43
| 1 | 6 | 1.43 | 1.5 | < 2 | 1.65 | 1.7 | 1.8 | 1.5 | 1.51 | 0.9 | < 5 | 1.5 | 1.32 | 1.8 | 1.8 | 1.42 | 1.5 | 1.35 | < 2 | 1.09 | 1.51 | 1.63 | 1.5
| Mean |  | 1.4667 | 1.5 |  | 1.63 | 1.735 | 1.7833 | 1.5 | 1.5683 | 1.0167 |  | 1.5 | 0.37333 | 1.9033 | 1.8333 | 1.4 | 1.4667 | 1.37 | 2 | 1.14 | 1.565 | 1.685 | 1.4917
| Median |  | 1.47 | 1.5 |  | 1.61 | 1.735 | 1.8 | 1.5 | 1.555 | 1 |  | 1.5 | 0.185 | 1.89 | 1.8 | 1.405 | 1.5 | 1.365 | 2 | 1.13 | 1.535 | 1.68 | 1.49
| Std Dev. |  | 0.028752 | 0 |  | 0.043359 | 0.0251 | 0.075277 | 0.063246 | 0.06047 | 0.17224 |  | 0 | 0.4665 | 0.07607 | 0.08165 | 0.04 | 0.05164 | 0.051769 |  | 0.054772 | 0.094181 | 0.045056 | 0.033714
| Rel.Std.Dev. |  | 1.96% | 0% |  | 2.66% | 1.447% | 4.221% | 4.216% | 3.856% | 16.94% |  | 0% | 125% | 3.997% | 4.454% | 2.857% | 3.521% | 3.779% |  | 4.805% | 6.018% | 2.674% | 2.26%
| PDM3 |  | -5.991% | -3.855% |  | 4.478% | 11.21% | 14.31% | -3.855% | 0.5253% | -34.83% |  | -3.855% | -76.07% | 22% | 17.51% | -10.26% | -5.991% | -12.19% | 28.19% | -26.93% | 0.3117% | 8.003% | -4.389%
| Z-Score (Absolute) |  | 0.17 | 0 | 2.59 | 0.67 | 1.22 | 1.47 | 0 | 0.35 | 2.51 | 5.19 | 0 | 5.84 | 2.09 | 1.73 | 0.52 | 0.17 | 0.67 | NA | 1.87 | 0.34 | 0.96 | 0.04
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 236 (Certified Value 2.15 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.21 | 2.07 | 1.905 | 2.29 | 2.55 | 2.22 | 2.0829 | 2.11 | 2.234 | 2.329 | 2.1192 | 2.12 | 2.2072 | 2.16 | 2.11 | 2.15 | 2.163 | 2.17 | 2.11 | 2.07 | 2.3 | 2.06 |  | 1
| 1 | 2 | 2.11 | 2.16 | 1.8782 | 2.35 | 2.41 | 2.27 | 2.1163 | 2.11 | 2.238 | 2.3431 | 2.1544 | 2.05 | 2.2344 | 2.21 | 2.14 | 2.15 | 2.156 | 2.21 | 2.16 | 2.11 | 2.35 | 2.12 |  | #N/A
| 1 | 3 | 2.06 | 2.2 | 1.936 | 2.22 | 2.53 | 2.23 | 2.0923 | 2.13 | 2.249 | 2.3354 | 2.1431 | 2.12 | 2.2482 | 2.23 | 2.04 | 2.17 | 2.283 | 2.15 | 2.16 | 2.15 | 2.28 | 2.17 |  | 16
| 1 | 4 | 2.08 | 2.09 | 1.89 | 2.32 | 2.46 | 2.24 | 2.0824 | 2.15 | 2.255 | 2.3386 | 2.1613 | 2.14 | 2.2511 | 2.16 | 2.1 | 2.16 | 2.12 | 2.15 | 2.14 | 2.2 | 2.33 | 2.11 |  | 2.155
| 1 | 5 | 2.06 | 2.16 | 1.8978 | 2.35 | 2.44 | 2.26 | 2.1359 | 2.18 | 2.254 | 2.3364 | 2.1273 | 2.15 | 2.2294 | 2.16 | 2.11 | 2.18 | 1.976 | 2.2 | 2.2 | 2.14 | 2.33 | 2.11 |  | 44
| 1 | 6 | 2.07 | 2.16 | 1.9263 | 2.32 | 2.51 | 2.24 | 2.1281 | 2.17 | 2.231 | 2.3408 | 2.1148 | 2.29 | 2.2205 | 2.17 | 2.17 | 2.17 | 2.148 | 2.15 | 2.14 | 2.12 | 2.33 | 2.11
| Mean |  | 2.0983 | 2.14 | 1.9056 | 2.3083 | 2.4833 | 2.2433 | 2.1063 | 2.1417 | 2.2435 | 2.3372 | 2.1367 | 2.145 | 2.2318 | 2.1817 | 2.1117 | 2.1633 | 2.141 | 2.1717 | 2.1517 | 2.1317 | 2.32 | 2.1133
| Median |  | 2.075 | 2.16 | 1.9014 | 2.32 | 2.485 | 2.24 | 2.1043 | 2.14 | 2.2435 | 2.3375 | 2.1352 | 2.13 | 2.2319 | 2.165 | 2.11 | 2.165 | 2.152 | 2.16 | 2.15 | 2.13 | 2.33 | 2.11
| Std Dev. |  | 0.057764 | 0.049396 | 0.021921 | 0.048751 | 0.055015 | 0.018619 | 0.023521 | 0.029944 | 0.010483 | 0.0049163 | 0.01915 | 0.079183 | 0.016654 | 0.030605 | 0.043551 | 0.012111 | 0.098578 | 0.027142 | 0.029944 | 0.043551 | 0.025298 | 0.035024
| Rel.Std.Dev. |  | 2.753% | 2.308% | 1.15% | 2.112% | 2.215% | 0.83% | 1.117% | 1.398% | 0.4673% | 0.2103% | 0.8963% | 3.692% | 0.7462% | 1.403% | 2.062% | 0.5598% | 4.604% | 1.25% | 1.392% | 2.043% | 1.09% | 1.657%
| PDM3 |  | -2.628% | -0.6942% | -11.57% | 7.117% | 15.24% | 4.101% | -2.257% | -0.6168% | 4.109% | 8.458% | -0.8481% | -0.4621% | 3.566% | 1.239% | -2.009% | 0.3886% | -0.6478% | 0.7753% | -0.1528% | -1.081% | 7.659% | -1.932%
| Z-Score (Absolute) |  | 0.94 | 0.16 | 4.57 | 3.01 | 6.3 | 1.79 | 0.79 | 0.13 | 1.79 | 3.55 | 0.22 | 0.06 | 1.57 | 0.63 | 0.69 | 0.28 | 0.14 | 0.44 | 0.06 | 0.31 | 3.23 | 0.66
| Analytical results for Nb in OREAS 236 (Certified Value 3.63 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 3.8 | 3.5 | < 5 | 4.3 | 4.6 | 3.6 | 3.67 | 3.64 | 2.6 | 3.4 | 0.5 | 4.1 | 3.1 | 3.5 | 3.23 | < 10 | 3.4 | 3.7 | 3.9 | 3.6 |  |  |  | 1
| 1 | 2 | 3.7 | 3.5 | < 5 | 4.3 | 4.9 | 3.9 | 3.71 | 3.5 | 3.2 | 3.5 | 0.2 | 3.9 | 3.1 | 3.6 | 3.22 | < 10 | 3.5 | 3.7 | 3.8 | 3.8 |  |  |  | 25
| 1 | 3 | 3.5 | 3.5 | < 5 | 4 | 4.5 | 3.9 | 3.69 | 3.56 | 3.4 | 3.5 | 0.3 | 3.9 | 3.2 | 3.6 | 3.31 | < 10 | 3.5 | 3.7 | 3.8 | 3.8 |  |  |  | 16
| 1 | 4 | 3.7 | 3.5 | < 5 | 4.3 | 4.8 | 3.8 | 3.65 | 3.62 | 2.4 | 3.7 | 0.4 | 4 | 3.1 | 3.5 | 3.13 | < 10 | 3.5 | 3.7 | 3.9 | 3.7 |  |  |  | 3.6281
| 1 | 5 | 3.6 | 3.5 | < 5 | 4.2 | 4.6 | 4 | 3.76 | 3.52 | 3.8 | 3.4 | 1 | 3.9 | 3.1 | 3.6 | 2.88 | < 10 | 3.7 | 3.7 | 4 | 3.8 |  |  |  | 45
| 1 | 6 | 3.7 | 3.5 | < 5 | 4.2 | 4.7 | 4 | 3.69 | 3.64 | 3.2 | 3.4 | 2.4 | 3.9 | 3.1 | 3.7 | 3.15 | < 10 | 3.4 | 3.7 | 3.8 | 3.6
| Mean |  | 3.6667 | 3.5 |  | 4.2167 | 4.6833 | 3.8667 | 3.695 | 3.58 | 3.1 | 3.4833 | 0.8 | 3.95 | 3.1167 | 3.5833 | 3.1533 |  | 3.5 | 3.7 | 3.8667 | 3.7167
| Median |  | 3.7 | 3.5 |  | 4.25 | 4.65 | 3.9 | 3.69 | 3.59 | 3.2 | 3.45 | 0.45 | 3.9 | 3.1 | 3.6 | 3.185 |  | 3.5 | 3.7 | 3.85 | 3.75
| Std Dev. |  | 0.10328 | 0 |  | 0.1169 | 0.1472 | 0.15055 | 0.037815 | 0.061968 | 0.51769 | 0.1169 | 0.83187 | 0.083666 | 0.040825 | 0.075277 | 0.14841 |  | 0.10954 | 4.8648e-16 | 0.08165 | 0.098319
| Rel.Std.Dev. |  | 2.817% | 0% |  | 2.772% | 3.143% | 3.894% | 1.023% | 1.731% | 16.7% | 3.356% | 104% | 2.118% | 1.31% | 2.101% | 4.707% |  | 3.13% | 1.315e-14% | 2.112% | 2.645%
| PDM3 |  | 1.063% | -3.531% |  | 16.22% | 29.08% | 6.575% | 1.844% | -1.326% | -14.56% | -3.99% | -77.95% | 8.872% | -14.1% | -1.234% | -13.09% |  | -3.531% | 1.982% | 6.575% | 2.441%
| Z-Score (Absolute) |  | 0.12 | 0.35 | 3.14 | 1.65 | 2.95 | 0.67 | 0.2 | 0.13 | 1.46 | 0.4 | 7.88 | 0.91 | 1.42 | 0.12 | 1.32 | 3.84 | 0.35 | 0.21 | 0.67 | 0.26
| Analytical results for Nd in OREAS 236 (Certified Value 9.01 ppm) | [24 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 9.55 | 9 | 8.85 | 9.67 | 8.7 | 8.9942 | 8.8 | 7.5 | 8.6 | 9.4 | [13 empty] | 1
| 1 | 2 | 9.5 | 9.4 | 9.16 | 9.38 | 8.6 | 8.7149 | 8.9 | 7.6 | 8.5 | 9.5 | [13 empty] | 26
| 1 | 3 | 9.55 | 9.4 | 9.01 | 9.64 | 8.6 | 8.9611 | 9.1 | 8.1 | 8.8 | 9.5 | [13 empty] | 16
| 1 | 4 | 9.15 | 9.1 | 8.87 | 9.16 | 8.8 | 8.803 | 9 | 7.8 | 8.2 | 9.5 | [13 empty] | 9.0119
| 1 | 5 | 9.6 | 9 | 9.31 | 9.2 | 8.9 | 8.9775 | 9 | 8.4 | 7.5 | 9.5 | [13 empty] | 46
| 1 | 6 | 9.6 | 9.5 | 9.3 | 9.62 | 9.2 | 8.9751 | 8.9 | 8.9 | 8.4 | 9.5
| Mean |  | 9.4917 | 9.2333 | 9.0833 | 9.445 | 8.8 | 8.9043 | 8.95 | 8.05 | 8.3333 | 9.4833
| Median |  | 9.55 | 9.25 | 9.085 | 9.5 | 8.75 | 8.9681 | 8.95 | 7.95 | 8.45 | 9.5
| Std Dev. |  | 0.17151 | 0.22509 | 0.20471 | 0.23011 | 0.22804 | 0.11645 | 0.10488 | 0.53198 | 0.45461 | 0.040825
| Rel.Std.Dev. |  | 1.807% | 2.438% | 2.254% | 2.436% | 2.591% | 1.308% | 1.172% | 6.608% | 5.455% | 0.4305%
| PDM3 |  | 5.323% | 2.457% | 0.7923% | 4.806% | -2.352% | -1.194% | -0.6872% | -10.67% | -7.53% | 5.231%
| Z-Score (Absolute) |  | 0.99 | 0.45 | 0.14 | 0.9 | 0.45 | 0.23 | 0.14 | 2.02 | 1.43 | 0.98
| Analytical results for Ni in OREAS 236 (Certified Value 68 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 72.2 | 74 | 48.365 | 61.5 | 65.8 | 73.2 | 69.8 | 72.7 | 76 | 60.19 | 66 | 66 | 68.4095 | 62 | 67.5 | 69.7 | 69 | 66.6 | 66 | 68 | 71.5 | 74 | 67.8 | 1
| 1 | 2 | 70 | 70 | 48.8558 | 62.9 | 67.3 | 74.3 | 70.5 | 73.7 | 77 | 58.93 | 68 | 64.2 | 68.4285 | 62 | 73.6 | 70 | 68 | 66.3 | 65 | 64 | 71.1 | 73.5 | 71.8 | 10
| 1 | 3 | 67.2 | 76 | 48.67 | 65.1 | 69.1 | 76.2 | 69.8 | 71.6 | 76 | 59.66 | 68 | 65.3 | 69.3785 | 63 | 66.6 | 69.4 | 69 | 67.9 | 64 | 67 | 73.8 | 72.3 | 73.1 | 16
| 1 | 4 | 68.5 | 68 | 48.48 | 65.5 | 65.6 | 75.1 | 69.2 | 73.2 | 75 | 59.45 | 68 | 65 | 70.0245 | 63 | 66.3 | 69.9 | 70 | 64.6 | 65 | 67 | 73.8 | 73.8 | 70.8 | 68.4509
| 1 | 5 | 67.7 | 70 | 48.39 | 65.2 | 64.2 | 76.7 | 71.2 | 70 | 76 | 60.11 | 66 | 68.8 | 68.643 | 62 | 66.1 | 71.1 | 69 | 60.2 | 65 | 67 | 73.2 | 73.3 | 72.8 | 47
| 1 | 6 | 68 | 72 | 48.61 | 62.3 | 68.8 | 74.2 | 70.9 | 73.3 | 75 | 60.09 | 67 | 72.4 | 67.6685 | 61 | 65.5 | 71.1 | 68 | 64.7 | 65 | 64 | 71.5 | 72.7 | 69.1
| Mean |  | 68.9333 | 71.6667 | 48.5618 | 63.75 | 66.8 | 74.95 | 70.2333 | 72.4167 | 75.8333 | 59.7383 | 67.1667 | 66.95 | 68.7588 | 62.1667 | 67.6 | 70.2 | 68.8333 | 65.05 | 65 | 66.1667 | 72.4833 | 73.2667 | 70.9
| Median |  | 68.25 | 71 | 48.545 | 64 | 66.55 | 74.7 | 70.15 | 72.95 | 76 | 59.875 | 67.5 | 65.65 | 68.5358 | 62 | 66.45 | 69.95 | 69 | 65.5 | 65 | 67 | 72.35 | 73.4 | 71.3
| Std Dev. |  | 1.8651 | 2.9439 | 0.18731 | 1.7248 | 1.936 | 1.3187 | 0.7607 | 1.3877 | 0.75277 | 0.49163 | 0.98319 | 3.1021 | 0.82696 | 0.75277 | 3.012 | 0.72664 | 0.75277 | 2.6809 | 0.63246 | 1.7224 | 1.2513 | 0.6532 | 2.1033
| Rel.Std.Dev. |  | 2.706% | 4.108% | 0.3857% | 2.706% | 2.898% | 1.759% | 1.083% | 1.916% | 0.9927% | 0.823% | 1.464% | 4.633% | 1.203% | 1.211% | 4.456% | 1.035% | 1.094% | 4.121% | 0.973% | 2.603% | 1.726% | 0.8915% | 2.967%
| PDM3 |  | 0.7047% | 4.698% | -29.06% | -6.868% | -2.412% | 9.494% | 2.604% | 5.794% | 10.78% | -12.73% | -1.876% | -2.193% | 0.4497% | -9.181% | -1.243% | 2.555% | 0.5587% | -4.968% | -5.041% | -3.337% | 5.891% | 7.035% | 3.578%
| Z-Score (Absolute) |  | 0.04 | 0.67 | 4.68 | 1.16 | 0.45 | 1.44 | 0.34 | 0.85 | 1.64 | 2.09 | 0.37 | 0.42 | 0 | 1.53 | 0.27 | 0.33 | 0.02 | 0.86 | 0.87 | 0.6 | 0.86 | 1.05 | 0.5
| Analytical results for P in OREAS 236 (Certified Value 0.045 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.046 | 0.05 | 0.043696 | 0.0396 | 0.054 | 0.0416 | 0.0463 | 0.046 | 0.0424 | 0.0442 | 0.042 | 0.048212 | 0.05 | 0.047 | 0.044 | 0.043 | 0.0457 | 0.047 | 0.046 | 0.0458 | 0.044 |  |  | 1
| 1 | 2 | 0.045 | 0.05 | 0.043723 | 0.0395 | 0.055 | 0.042 | 0.0457 | 0.046 | 0.0421 | 0.0453 | 0.042 | 0.049238 | 0.05 | 0.047 | 0.043 | 0.043 | 0.0465 | 0.0462 | 0.046 | 0.047 | 0.047 |  |  | #N/A
| 1 | 3 | 0.044 | 0.055 | 0.044353 | 0.0399 | 0.056 | 0.0417 | 0.046 | 0.047 | 0.0421 | 0.0454 | 0.043 | 0.048311 | 0.05 | 0.046 | 0.044 | 0.044 | 0.0456 | 0.047 | 0.048 | 0.0453 | 0.047 |  |  | 16
| 1 | 4 | 0.045 | 0.05 | 0.044124 | 0.0414 | 0.057 | 0.0416 | 0.047 | 0.046 | 0.0426 | 0.0451 | 0.043 | 0.04811 | 0.05 | 0.048 | 0.044 | 0.041 | 0.0456 | 0.0462 | 0.048 | 0.0463 | 0.046 |  |  | 0.045336
| 1 | 5 | 0.044 | 0.05 | 0.044152 | 0.043 | 0.055 | 0.0421 | 0.046 | 0.047 | 0.0423 | 0.044 | 0.044 | 0.048194 | 0.05 | 0.047 | 0.044 | 0.039 | 0.0466 | 0.0453 | 0.048 | 0.0464 | 0.045 |  |  | 48
| 1 | 6 | 0.044 | 0.05 | 0.044203 | 0.0422 | 0.055 | 0.0429 | 0.0459 | 0.045 | 0.0425 | 0.0449 | 0.046 | 0.047997 | 0.05 | 0.047 | 0.044 | 0.041 | 0.0457 | 0.0467 | 0.046 | 0.0468 | 0.045
| Mean |  | 0.044667 | 0.050833 | 0.044042 | 0.040933 | 0.055333 | 0.041983 | 0.04615 | 0.046167 | 0.042333 | 0.044817 | 0.043333 | 0.048344 | 0.05 | 0.047 | 0.043833 | 0.041833 | 0.04595 | 0.0464 | 0.047 | 0.046267 | 0.045667
| Median |  | 0.0445 | 0.05 | 0.044138 | 0.04065 | 0.055 | 0.04185 | 0.046 | 0.046 | 0.04235 | 0.045 | 0.043 | 0.048203 | 0.05 | 0.047 | 0.044 | 0.042 | 0.0457 | 0.04645 | 0.047 | 0.04635 | 0.0455
| Std Dev. |  | 0.0008165 | 0.0020412 | 0.00026944 | 0.0014828 | 0.0010328 | 0.00049565 | 0.00045935 | 0.00075277 | 0.00020656 | 0.00058452 | 0.0015055 | 0.00045074 | 7.6012e-18 | 0.00063246 | 0.00040825 | 0.0018348 | 0.00046797 | 0.00064807 | 0.0010954 | 0.0006314 | 0.0012111
| Rel.Std.Dev. |  | 1.828% | 4.016% | 0.6118% | 3.622% | 1.866% | 1.181% | 0.9953% | 1.631% | 0.4879% | 1.304% | 3.474% | 0.9324% | 1.52e-14% | 1.346% | 0.9314% | 4.386% | 1.018% | 1.397% | 2.331% | 1.365% | 2.652%
| PDM3 |  | -1.476% | 12.13% | -2.854% | -9.711% | 22.05% | -7.395% | 1.796% | 1.832% | -6.623% | -1.145% | -4.417% | 6.635% | 10.29% | 3.67% | -3.314% | -7.726% | 1.354% | 2.347% | 3.67% | 2.053% | 0.7295%
| Z-Score (Absolute) |  | 0.45 | 1.73 | 0.67 | 1.77 | 3.32 | 1.4 | 0.07 | 0.08 | 1.28 | 0.4 | 0.92 | 0.85 | 1.43 | 0.37 | 0.75 | 1.45 | 0 | 0.16 | 0.37 | 0.11 | 0.1
| Analytical results for Pb in OREAS 236 (Certified Value 30.5 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 30.9 | 30 | 27.65 | 31.5 | 30.1 | 32.7 | 31.8 | 32 | 29 | 28 | 32.1 | 26.7 | 36.2 | 27.8 | 35 | 29.95 | 31 | 28.6 | 31.4 | 30.4 | 29.3 |  |  | 1
| 1 | 2 | 30.2 | 30 | 27.9546 | 31.9 | 30.8 | 32.4 | 31.9 | 31.9 | 31 | 28 | 32 | 27.7 | 37.5 | 27.7 | 33 | 29.52 | 30 | 29.6 | 32.1 | 30.2 | 30.3 |  |  | 12
| 1 | 3 | 28.7 | 38 | 28.0785 | 32.9 | 31 | 32.1 | 31.8 | 32.1 | 30 | 28 | 32.2 | 26.8 | 33.4 | 25.8 | 32 | 30.72 | 29 | 29.7 | 31.7 | 29.9 | 31.1 |  |  | 16
| 1 | 4 | 28.9 | 30 | 27.6242 | 33.1 | 32.1 | 33.3 | 31.6 | 31.5 | 30 | 29 | 32 | 27.2 | 37.6 | 27.4 | 33 | 27.62 | 30 | 29 | 32.8 | 30.3 | 28.9 |  |  | 30.4861
| 1 | 5 | 28.6 | 30 | 27.95 | 32.7 | 32.4 | 34.6 | 32 | 30.8 | 30 | 28 | 32.7 | 28.8 | 33.3 | 28.3 | 33 | 25.79 | 28 | 29.9 | 32.8 | 31.3 | 29.5 |  |  | 49
| 1 | 6 | 29.7 | 30 | 27.8835 | 32.4 | 31.4 | 40.3 | 31.9 | 31.6 | 29 | 28 | 34.2 | 27.1 | 32.5 | 28.5 | 34 | 29.05 | 30 | 29.7 | 31.7 | 30.3 | 30.1
| Mean |  | 29.5 | 31.3333 | 27.8568 | 32.4167 | 31.3 | 34.2333 | 31.8333 | 31.65 | 29.8333 | 28.1667 | 32.5333 | 27.3833 | 35.0833 | 27.5833 | 33.3333 | 28.775 | 29.6667 | 29.4167 | 32.0833 | 30.4 | 29.8667
| Median |  | 29.3 | 30 | 27.9168 | 32.55 | 31.2 | 33 | 31.85 | 31.75 | 30 | 28 | 32.15 | 27.15 | 34.8 | 27.75 | 33 | 29.285 | 30 | 29.65 | 31.9 | 30.3 | 29.8
| Std Dev. |  | 0.92736 | 3.266 | 0.18164 | 0.61455 | 0.85323 | 3.1008 | 0.13663 | 0.47645 | 0.75277 | 0.40825 | 0.85713 | 0.77825 | 2.2851 | 0.96212 | 1.0328 | 1.791 | 1.0328 | 0.50365 | 0.59805 | 0.47329 | 0.79415
| Rel.Std.Dev. |  | 3.144% | 10.42% | 0.652% | 1.896% | 2.726% | 9.058% | 0.4292% | 1.505% | 2.523% | 1.449% | 2.635% | 2.842% | 6.513% | 3.488% | 3.098% | 6.224% | 3.481% | 1.712% | 1.864% | 1.557% | 2.659%
| PDM3 |  | -3.235% | 2.779% | -8.625% | 6.333% | 2.67% | 12.29% | 4.419% | 3.818% | -2.141% | -7.608% | 6.715% | -10.18% | 15.08% | -9.522% | 9.339% | -5.613% | -2.688% | -3.508% | 5.239% | -0.2825% | -2.032%
| Z-Score (Absolute) |  | 0.37 | 0.39 | 1.06 | 0.84 | 0.37 | 1.59 | 0.59 | 0.52 | 0.24 | 0.93 | 0.89 | 1.25 | 1.94 | 1.17 | 1.22 | 0.67 | 0.3 | 0.41 | 0.7 | 0 | 0.22
| Analytical results for Pr in OREAS 236 (Certified Value 1.94 ppm) | [24 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 1.88 | 1.91 | 1.93 | 2.06 | 1.9 | 2.2807 | 1.89 | 1.1 | 1.8 | 2.06 | [13 empty] | 1
| 1 | 2 | 1.96 | 2.03 | 1.91 | 1.9 | 2 | 2.2381 | 1.87 | 1.1 | 1.8 | 2.11 | [13 empty] | 29
| 1 | 3 | 1.92 | 1.99 | 1.94 | 2.06 | 1.9 | 2.24 | 1.89 | 1.2 | 1.8 | 2.11 | [13 empty] | 16
| 1 | 4 | 1.88 | 1.94 | 1.96 | 2 | 1.8 | 2.21 | 1.9 | 1.1 | 1.7 | 2.09 | [13 empty] | 1.9404
| 1 | 5 | 1.92 | 1.91 | 1.96 | 1.96 | 1.9 | 2.152 | 1.91 | 1.3 | 1.5 | 2.1 | [13 empty] | 50
| 1 | 6 | 1.98 | 1.97 | 1.99 | 2.11 | 2.1 | 2.1551 | 1.9 | 1.4 | 1.7 | 2.08
| Mean |  | 1.9233 | 1.9583 | 1.9483 | 2.015 | 1.9333 | 2.2126 | 1.8933 | 1.2 | 1.7167 | 2.0917
| Median |  | 1.92 | 1.955 | 1.95 | 2.03 | 1.9 | 2.2241 | 1.895 | 1.15 | 1.75 | 2.095
| Std Dev. |  | 0.040825 | 0.047504 | 0.027869 | 0.076877 | 0.10328 | 0.051054 | 0.013663 | 0.12649 | 0.1169 | 0.019408
| Rel.Std.Dev. |  | 2.123% | 2.426% | 1.43% | 3.815% | 5.342% | 2.307% | 0.7216% | 10.54% | 6.81% | 0.9279%
| PDM3 |  | -0.8804% | 0.9233% | 0.408% | 3.844% | -0.365% | 14.03% | -2.426% | -38.16% | -11.53% | 7.795%
| Z-Score (Absolute) |  | 0.19 | 0.19 | 0.08 | 0.82 | 0.08 | 3.01 | 0.53 | 8.21 | 2.48 | 1.67
| Analytical results for Rb in OREAS 236 (Certified Value 14.1 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 15 | 14.2 | 13.9 | 16.1 | 14.8 | 14.13 | 14.68 | 14.2 | 13.3 | 14 | 14.1414 | 12.9 | 13.6 | 12.8 | 13.4 | 14.3 | 14.4 | 15.6 | 13.1 |  |  |  |  | 1
| 1 | 2 | 14.8 | 14.8 | 14.2 | 16.8 | 15 | 13.85 | 13.7 | 14.4 | 13.4 | 13.5 | 13.9572 | 13.3 | 13.7 | 12.3 | 13.4 | 14.7 | 14.4 | 15.7 | 13.4 |  |  |  |  | 30
| 1 | 3 | 14.1 | 14.2 | 14.5 | 16.3 | 15.6 | 13.46 | 15.3 | 14.3 | 13.8 | 13.8 | 13.9892 | 12.8 | 13.1 | 12.1 | 14 | 14.5 | 14.9 | 15.2 | 13.9 |  |  |  |  | 16
| 1 | 4 | 14.5 | 14 | 14.7 | 16.3 | 14.8 | 13.65 | 14.98 | 14.2 | 14 | 13.8 | 13.7682 | 12.8 | 14.1 | 12.6 | 13.1 | 14.4 | 14.4 | 15.9 | 13.4 |  |  |  |  | 14.0762
| 1 | 5 | 14.3 | 14.2 | 14.4 | 16.3 | 15.5 | 13.96 | 15.11 | 14.4 | 13.3 | 14.1 | 14.4697 | 13.4 | 14 | 12.8 | 12 | 15.2 | 14.9 | 16.4 | 13.3 |  |  |  |  | 51
| 1 | 6 | 14.3 | 14.2 | 14.5 | 16.4 | 15.4 | 14.25 | 14.47 | 14.3 | 13.5 | 15.7 | 14.0602 | 12.7 | 14.2 | 12.8 | 13.1 | 14.6 | 14.1 | 15.6 | 13.6
| Mean |  | 14.5 | 14.2667 | 14.3667 | 16.3667 | 15.1833 | 13.8833 | 14.7067 | 14.3 | 13.55 | 14.15 | 14.0643 | 12.9833 | 13.7833 | 12.5667 | 13.1667 | 14.6167 | 14.5167 | 15.7333 | 13.45
| Median |  | 14.4 | 14.2 | 14.45 | 16.3 | 15.2 | 13.905 | 14.83 | 14.3 | 13.45 | 13.9 | 14.0247 | 12.85 | 13.85 | 12.7 | 13.25 | 14.55 | 14.4 | 15.65 | 13.4
| Std Dev. |  | 0.34059 | 0.27325 | 0.28048 | 0.23381 | 0.36009 | 0.29527 | 0.57611 | 0.089443 | 0.2881 | 0.78677 | 0.23456 | 0.29269 | 0.40702 | 0.30111 | 0.65929 | 0.31885 | 0.31885 | 0.39833 | 0.27386
| Rel.Std.Dev. |  | 2.349% | 1.915% | 1.952% | 1.429% | 2.372% | 2.127% | 3.917% | 0.6255% | 2.126% | 5.56% | 1.668% | 2.254% | 2.953% | 2.396% | 5.007% | 2.181% | 2.196% | 2.532% | 2.036%
| PDM3 |  | 3.011% | 1.353% | 2.064% | 16.27% | 7.866% | -1.37% | 4.479% | 1.59% | -3.738% | 0.5245% | -0.08415% | -7.764% | -2.08% | -10.72% | -6.461% | 3.84% | 3.129% | 11.77% | -4.448%
| Z-Score (Absolute) |  | 0.36 | 0 | 0.15 | 3.22 | 1.4 | 0.59 | 0.67 | 0.05 | 1.1 | 0.18 | 0.31 | 1.97 | 0.74 | 2.61 | 1.69 | 0.54 | 0.38 | 2.25 | 1.25
| Analytical results for Re in OREAS 236 (Certified Value 0.003 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.002 | < 0.1 | 0.005 | 0.002 | 0.006 | < 0.05 | < 0.05 | < 0.001 | 0.002 | < 0.05 | < 0.002 | < 5 | < 0.002 | 0.003 | 0.003 | 0.002 | [7 empty] | 1
| 1 | 2 | 0.002 | < 0.1 | 0.008 | < 0.002 | 0.005 | < 0.05 | < 0.05 | < 0.001 | 0.002 | < 0.05 | < 0.002 | < 5 | 0.009 | 0.004 | 0.003 | 0.002 | [7 empty] | 31
| 1 | 3 | 0.003 | < 0.1 | 0.005 | 0.002 | 0.015 | < 0.05 | < 0.05 | < 0.001 | 0.002 | < 0.05 | 0.002 | < 5 | < 0.002 | 0.004 | 0.003 | 0.003 | [7 empty] | 16
| 1 | 4 | 0.002 | < 0.1 | 0.006 | < 0.002 | 0.003 | < 0.05 | < 0.05 | < 0.001 | 0.003 | < 0.05 | < 0.002 | < 5 | 0.003 | 0.003 | 0.003 | 0.002 | [7 empty] | 0.0026667
| 1 | 5 | 0.002 | < 0.1 | 0.006 | 0.004 | 0.01 | < 0.05 | < 0.05 | 0.001 | 0.003 | < 0.05 | < 0.002 | < 5 | 0.007 | 0.004 | 0.003 | 0.002 | [7 empty] | 52
| 1 | 6 | 0.002 | < 0.1 | 0.005 | 0.004 | 0.005 | < 0.05 | < 0.05 | 0.003 | 0.002 | < 0.05 | 0.003 | < 5 | < 0.002 | 0.004 | 0.005 | 0.003
| Mean |  | 0.0021667 |  | 0.0058333 | 0.003 | 0.0073333 |  |  | 0.002 | 0.0023333 |  | 0.0025 |  | 0.0063333 | 0.0036667 | 0.0033333 | 0.0023333
| Median |  | 0.002 |  | 0.0055 | 0.003 | 0.0055 |  |  | 0.002 | 0.002 |  | 0.0025 |  | 0.007 | 0.004 | 0.003 | 0.002
| Std Dev. |  | 0.00040825 |  | 0.001169 | 0.0011547 | 0.0044121 |  |  | 0.0014142 | 0.0005164 |  | 0.00070711 |  | 0.0030551 | 0.0005164 | 0.0008165 | 0.0005164
| Rel.Std.Dev. |  | 18.84% |  | 20.04% | 38.49% | 60.17% |  |  | 70.71% | 22.13% |  | 28.28% |  | 48.24% | 14.08% | 24.49% | 22.13%
| PDM3 |  | -18.75% |  | 118.8% | 12.5% | 175% |  |  | -25% | -12.5% |  | -6.25% |  | 137.5% | 37.5% | 25% | -12.5%
| Z-Score (Absolute) |  | 0.47 | 14.42 | 0.67 | 0.41 | 1.14 | 6.64 | 6.64 | 0.83 | 0.41 | 6.64 | 0.67 | 776.91 | 0 | 0 | 0.1 | 0.41
| Analytical results for S in OREAS 236 (Certified Value 0.444 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.44 | 0.445 | 0.50712 | 0.44 | 0.49 | 0.44 | 0.47 | 0.44 | 0.454 | 0.4331 | 0.41 | 0.47 | 0.44 | 0.441 | 0.37 | 0.45 | 0.43 | 0.44 | 0.46 | 0.43 |  |  |  | 1
| 1 | 2 | 0.43 | 0.445 | 0.50976 | 0.45 | 0.49 | 0.45 | 0.43 | 0.44 | 0.458 | 0.443 | 0.41 | 0.48 | 0.45 | 0.433 | 0.37 | 0.45 | 0.42 | 0.44 | 0.46 | 0.44 |  |  |  | 16
| 1 | 3 | 0.42 | 0.465 | 0.50256 | 0.45 | 0.5 | 0.46 | 0.42 | 0.44 | 0.458 | 0.4437 | 0.42 | 0.47 | 0.45 | 0.444 | 0.38 | 0.44 | 0.41 | 0.45 | 0.46 | 0.45 |  |  |  | 16
| 1 | 4 | 0.43 | 0.45 | 0.50976 | 0.45 | 0.49 | 0.46 | 0.45 | 0.45 | 0.458 | 0.4459 | 0.41 | 0.47 | 0.44 | 0.441 | 0.37 | 0.45 | 0.44 | 0.46 | 0.46 | 0.44 |  |  |  | 0.44412
| 1 | 5 | 0.42 | 0.475 | 0.51552 | 0.45 | 0.48 | 0.45 | 0.42 | 0.44 | 0.461 | 0.4335 | 0.41 | 0.49 | 0.43 | 0.441 | 0.33 | 0.45 | 0.4 | 0.45 | 0.47 | 0.44 |  |  |  | 53
| 1 | 6 | 0.43 | 0.455 | 0.50832 | 0.45 | 0.5 | 0.45 | 0.42 | 0.44 | 0.455 | 0.4412 | 0.42 | 0.47 | 0.45 | 0.441 | 0.36 | 0.45 | 0.43 | 0.44 | 0.47 | 0.44
| Mean |  | 0.42833 | 0.45583 | 0.50884 | 0.44833 | 0.49167 | 0.45167 | 0.435 | 0.44167 | 0.45733 | 0.44007 | 0.41333 | 0.475 | 0.44333 | 0.44017 | 0.36333 | 0.44833 | 0.42167 | 0.44667 | 0.46333 | 0.44
| Median |  | 0.43 | 0.4525 | 0.50904 | 0.45 | 0.49 | 0.45 | 0.425 | 0.44 | 0.458 | 0.4421 | 0.41 | 0.47 | 0.445 | 0.441 | 0.37 | 0.45 | 0.425 | 0.445 | 0.46 | 0.44
| Std Dev. |  | 0.0075277 | 0.012007 | 0.0042186 | 0.0040825 | 0.0075277 | 0.0075277 | 0.020736 | 0.0040825 | 0.0025033 | 0.0054548 | 0.005164 | 0.0083666 | 0.008165 | 0.0037103 | 0.017512 | 0.0040825 | 0.01472 | 0.008165 | 0.005164 | 0.0063246
| Rel.Std.Dev. |  | 1.757% | 2.634% | 0.8291% | 0.9106% | 1.531% | 1.667% | 4.767% | 0.9243% | 0.5474% | 1.24% | 1.249% | 1.761% | 1.842% | 0.8429% | 4.82% | 0.9106% | 3.491% | 1.828% | 1.115% | 1.437%
| PDM3 |  | -3.555% | 2.637% | 14.57% | 0.9483% | 10.71% | 1.699% | -2.054% | -0.5528% | 2.975% | -0.913% | -6.932% | 6.953% | -0.1775% | -0.8905% | -18.19% | 0.9483% | -5.056% | 0.5731% | 4.326% | -0.928%
| Z-Score (Absolute) |  | 1.08 | 0.7 | 4.13 | 0.22 | 3.02 | 0.43 | 0.65 | 0.22 | 0.8 | 0.32 | 2.05 | 1.94 | 0.11 | 0.31 | 5.29 | 0.22 | 1.51 | 0.11 | 1.19 | 0.32
| Analytical results for Sb in OREAS 236 (Certified Value 1.76 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.84 | 1.8 | < 5 | 1.64 | 2.03 | 1.3 | 1.72 | 1.81 | 1.7 | < 10 | 1.6 | 0.2 | 1.9 | 2 | 1.51 | 1.7 | 1.6 | < 5 | 1.72 | 1.82 | 1.7 | 1.73 |  | 1
| 1 | 2 | 1.8 | 1.8 | < 5 | 1.63 | 2.14 | 1.4 | 1.62 | 1.81 | 1.8 | < 10 | 1.6 | 0.1 | 1.86 | 2 | 1.6 | 1.6 | 1.62 | < 5 | 1.75 | 1.89 | 1.6 | 1.78 |  | 32
| 1 | 3 | 1.76 | 1.8 | < 5 | 1.65 | 2.2 | 1.3 | 1.65 | 1.79 | 1.9 | < 10 | 1.6 | 0.2 | 1.8 | 2 | 1.66 | 1.6 | 1.68 | < 5 | 1.7 | 1.97 | 1.7 | 1.83 |  | 16
| 1 | 4 | 1.79 | 1.8 | < 5 | 1.65 | 2.06 | 1.4 | 1.62 | 1.85 | 1.7 | < 10 | 1.6 | 0.1 | 1.9 | 2 | 1.56 | 1.6 | 1.54 | < 5 | 1.71 | 1.9 | 1.7 | 1.79 |  | 1.7628
| 1 | 5 | 1.73 | 1.8 | < 5 | 1.68 | 2.01 | 1.3 | 1.73 | 1.87 | 1.9 | < 10 | 1.6 | 0.2 | 1.92 | 1.9 | 1.52 | 1.7 | 1.47 | < 5 | 1.83 | 1.91 | 1.7 | 1.74 |  | 54
| 1 | 6 | 1.84 | 1.8 | < 5 | 1.72 | 2.18 | 1.4 | 1.66 | 1.72 | 1.8 | < 10 | 1.6 | 0.3 | 1.8 | 2.3 | 1.6 | 1.7 | 1.59 | < 5 | 1.8 | 1.84 | 1.7 | 1.76
| Mean |  | 1.7933 | 1.8 |  | 1.6617 | 2.1033 | 1.35 | 1.6667 | 1.8083 | 1.8 |  | 1.6 | 0.18333 | 1.8633 | 2.0333 | 1.575 | 1.65 | 1.5833 |  | 1.7517 | 1.8883 | 1.6833 | 1.7717
| Median |  | 1.795 | 1.8 |  | 1.65 | 2.1 | 1.35 | 1.655 | 1.81 | 1.8 |  | 1.6 | 0.2 | 1.88 | 2 | 1.58 | 1.65 | 1.595 |  | 1.735 | 1.895 | 1.7 | 1.77
| Std Dev. |  | 0.043665 | 0 |  | 0.033116 | 0.080664 | 0.054772 | 0.048028 | 0.052313 | 0.089443 |  | 2.4324e-16 | 0.075277 | 0.052789 | 0.13663 | 0.05648 | 0.054772 | 0.07174 |  | 0.052694 | 0.053448 | 0.040825 | 0.03656
| Rel.Std.Dev. |  | 2.435% | 0% |  | 1.993% | 3.835% | 4.057% | 2.882% | 2.893% | 4.969% |  | 1.52e-14% | 41.06% | 2.833% | 6.719% | 3.586% | 3.32% | 4.531% |  | 3.008% | 2.83% | 2.425% | 2.064%
| PDM3 |  | 1.73% | 2.108% |  | -5.739% | 19.31% | -23.42% | -5.456% | 2.581% | 2.108% |  | -9.238% | -89.6% | 5.7% | 15.34% | -10.66% | -6.401% | -10.18% |  | -0.634% | 7.119% | -4.51% | 0.5005%
| Z-Score (Absolute) |  | 0.06 | 0.09 | 3.82 | 0.64 | 1.71 | 2.3 | 0.62 | 0.14 | 0.09 | 17.13 | 0.97 | 8.51 | 0.43 | 1.34 | 1.1 | 0.71 | 1.06 | 3.82 | 0.16 | 0.56 | 0.53 | 0.06
| Analytical results for Sc in OREAS 236 (Certified Value 38.5 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 41.8 | 39 | 31.25 | 36 | 40.8 | 37.8 | 36.2 | 44 | 51 | 38 | 37 | 38.9075 | 45.7 | 28.8 | 35.9 | 36 | 37.4 | 37 | 40 | 37 | 40.6 | 36.2 |  | 1
| 1 | 2 | 39.9 | 39 | 31.5712 | 36.8 | 41 | 39.5 | 36.9 | 42 | 53 | 39 | 37 | 38.8164 | 46.5 | 26.7 | 35.5 | 36 | 37.8 | 37 | 39 | 38.5 | 41.4 | 39.1 |  | 33
| 1 | 3 | 38.1 | 41 | 31.1615 | 35.8 | 41.3 | 39.1 | 36.6 | 42.1 | 51 | 39 | 38 | 38.3456 | 45.2 | 31.2 | 36.5 | 36 | 38.9 | 36 | 40 | 39.5 | 40.2 | 40 |  | 16
| 1 | 4 | 39.6 | 39 | 31.6243 | 37.2 | 40.1 | 39.7 | 35.9 | 41.5 | 51 | 39 | 38 | 38.2354 | 46.2 | 30.6 | 37.1 | 36 | 36.9 | 36 | 40 | 39.2 | 41 | 37.7 |  | 38.5064
| 1 | 5 | 39.4 | 39 | 31.05 | 36.7 | 40.1 | 39.9 | 36.1 | 43.5 | 53 | 38 | 37 | 37.7024 | 43.2 | 34.8 | 36.2 | 36 | 34.3 | 37 | 39 | 39.1 | 41 | 38.5 |  | 55
| 1 | 6 | 39.6 | 39 | 31.2375 | 36.9 | 41.4 | 39.4 | 37 | 43.6 | 51 | 38 | 38 | 38.3019 | 45.9 | 37.6 | 37.4 | 36 | 37.2 | 37 | 40 | 38.2 | 41.1 | 37.1
| Mean |  | 39.7333 | 39.3333 | 31.3158 | 36.5667 | 40.7833 | 39.2333 | 36.45 | 42.7833 | 51.6667 | 38.5 | 37.5 | 38.3849 | 45.45 | 31.6167 | 36.4333 | 36 | 37.0833 | 36.6667 | 39.6667 | 38.5833 | 40.8833 | 38.1
| Median |  | 39.6 | 39 | 31.2438 | 36.75 | 40.9 | 39.45 | 36.4 | 42.8 | 51 | 38.5 | 37.5 | 38.3237 | 45.8 | 30.9 | 36.35 | 36 | 37.3 | 37 | 40 | 38.8 | 41 | 38.1
| Std Dev. |  | 1.1928 | 0.8165 | 0.23035 | 0.5465 | 0.57067 | 0.75277 | 0.45056 | 1.0381 | 1.0328 | 0.54772 | 0.54772 | 0.43723 | 1.1879 | 3.9812 | 0.72019 | 0 | 1.5303 | 0.5164 | 0.5164 | 0.91086 | 0.42151 | 1.3813
| Rel.Std.Dev. |  | 3.002% | 2.076% | 0.7356% | 1.495% | 1.399% | 1.919% | 1.236% | 2.426% | 1.999% | 1.423% | 1.461% | 1.139% | 2.614% | 12.59% | 1.977% | 0% | 4.127% | 1.408% | 1.302% | 2.361% | 1.031% | 3.625%
| PDM3 |  | 3.186% | 2.148% | -18.67% | -5.037% | 5.913% | 1.888% | -5.34% | 11.11% | 34.18% | -0.01657% | -2.614% | -0.3156% | 18.03% | -17.89% | -5.384% | -6.509% | -3.696% | -4.778% | 3.013% | 0.1998% | 6.173% | -1.055%
| Z-Score (Absolute) |  | 0.48 | 0.33 | 2.63 | 0.69 | 0.86 | 0.29 | 0.74 | 1.6 | 4.88 | 0.02 | 0.35 | 0.02 | 2.59 | 2.52 | 0.74 | 0.9 | 0.5 | 0.66 | 0.45 | 0.05 | 0.9 | 0.13
| Analytical results for Se in OREAS 236 (Certified Value < 5 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 1 | < 5 | < 5 | < 2 | < 0.5 | < 0.5 | 30 | < 10 | < 1 | < 0.1 | < 2 | 0.5 | 1 | 1 | 0.3 | < 5 | 1 | 1 | < 1 | 1 |  |  |  | 1
| 1 | 2 | 1 | < 5 | < 5 | < 2 | 0.6 | 0.7 | 31 | < 10 | < 1 | 0.2 | < 2 | < 0.1 | 1 | 1 | < 0.3 | < 5 | 2 | 1 | < 1 | 1 |  |  |  | 34
| 1 | 3 | 1 | < 5 | < 5 | < 2 | < 0.5 | < 0.5 | 25 | < 10 | < 1 | 0.2 | < 2 | 0.6 | 1 | 1 | < 0.3 | < 5 | 1 | 1 | < 1 | 1 |  |  |  | 16
| 1 | 4 | 1 | < 5 | < 5 | < 2 | 0.5 | 1.2 | 27 | < 10 | < 1 | 0.2 | < 2 | 0.5 | 1 | 1 | 0.4 | < 5 | 2 | 1 | < 1 | 1 |  |  |  | < 5
| 1 | 5 | 1 | < 5 | < 5 | < 2 | 0.5 | < 0.5 | 26 | < 10 | < 1 | 0.3 | < 2 | < 0.1 | 1 | 1 | < 0.3 | < 5 | 1 | 1 | < 1 | 1 |  |  |  | 56
| 1 | 6 | < 1 | < 5 | < 5 | < 2 | < 0.5 | 3 | 30 | < 10 | < 1 | 0.9 | < 2 | < 0.1 | 1 | 1 | < 0.3 | < 5 | 2 | < 1 | < 1 | 1
| Mean |  | 1 |  |  |  | 0.53333 | 1.6333 | 28.1667 |  |  | 0.36 |  | 0.53333 | 1 | 1 | 0.35 |  | 1.5 | 1 |  | 1
| Median |  | 1 |  |  |  | 0.5 | 1.2 | 28.5 |  |  | 0.2 |  | 0.5 | 1 | 1 | 0.35 |  | 1.5 | 1 |  | 1
| Std Dev. |  | 0 |  |  |  | 0.057735 | 1.2097 | 2.4833 |  |  | 0.30496 |  | 0.057735 | 0 | 0 | 0.070711 |  | 0.54772 | 0 |  | 0
| Rel.Std.Dev. |  | 0% |  |  |  | 10.83% | 74.06% | 8.816% |  |  | 84.71% |  | 10.83% | 0% | 0% | 20.2% |  | 36.51% | 0% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.11 | 2.02 | 2.02 | 0 | 0.82 | 0.08 | 36.64 | 5.39 | 0.67 | 0.93 | 0 | 0.96 | 0 | 0 | 1.06 | 2.02 | 0.67 | 0.11 | 0.67 | 0
| Analytical results for Sm in OREAS 236 (Certified Value 2.71 ppm) | [24 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 2.7 | 2.7 | 2.68 | 3.05 | 2.6 | 2.5929 | 2.6 | 2 | 2.5 | 2.79 | [13 empty] | 1
| 1 | 2 | 2.85 | 2.8 | 2.61 | 2.89 | 2.7 | 2.5949 | 2.7 | 2 | 2.5 | 2.83 | [13 empty] | 18
| 1 | 3 | 2.8 | 2.8 | 2.69 | 3.02 | 3 | 2.5867 | 2.7 | 2.1 | 2.5 | 2.79 | [13 empty] | 16
| 1 | 4 | 2.75 | 2.7 | 2.69 | 3.15 | 2.7 | 2.5848 | 2.7 | 2 | 2.4 | 2.83 | [13 empty] | 2.7077
| 1 | 5 | 2.7 | 2.6 | 2.76 | 2.81 | 2.4 | 2.4418 | 2.7 | 2.2 | 2.2 | 2.8 | [13 empty] | 57
| 1 | 6 | 2.8 | 2.8 | 2.82 | 2.95 | 2.7 | 2.6029 | 2.7 | 2.4 | 2.4 | 2.8
| Mean |  | 2.7667 | 2.7333 | 2.7083 | 2.9783 | 2.6833 | 2.5674 | 2.6833 | 2.1167 | 2.4167 | 2.8067
| Median |  | 2.775 | 2.75 | 2.69 | 2.985 | 2.7 | 2.5898 | 2.7 | 2.05 | 2.45 | 2.8
| Std Dev. |  | 0.060553 | 0.08165 | 0.072503 | 0.12106 | 0.19408 | 0.061826 | 0.040825 | 0.16021 | 0.1169 | 0.018619
| Rel.Std.Dev. |  | 2.189% | 2.987% | 2.677% | 4.065% | 7.233% | 2.408% | 1.521% | 7.569% | 4.837% | 0.6634%
| PDM3 |  | 2.178% | 0.9474% | 0.02409% | 9.996% | -0.8992% | -5.182% | -0.8992% | -21.83% | -10.75% | 3.656%
| Z-Score (Absolute) |  | 0.53 | 0.28 | 0.09 | 2.1 | 0.09 | 0.95 | 0.09 | 4.3 | 2.07 | 0.82
| Analytical results for Sn in OREAS 236 (Certified Value 1.08 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.1 | 1 | < 10 | 1.1 | 1.7 | 1.6 | 1.1 | 1.2 | 0.9 | < 10 | 1 | < 1 | 1.1 | 0.5 | 0.9 | 1.3 | 0.9 | < 10 | 1.1 | 1.1 | 1.2 | 0.9 |  | 1
| 1 | 2 | 1 | 1 | < 10 | 1.1 | 1.7 | 1.5 | 1.1 | 1.1 | 1 | < 10 | 1 | < 1 | 1.2 | 0.4 | 0.9 | 1.3 | 1 | < 10 | 1.2 | 1.1 | 1.2 | 0.9 |  | 7
| 1 | 3 | 1 | 1 | < 10 | 1.1 | 1.7 | 1.6 | 1.1 | 1.3 | 1.1 | < 10 | 1 | < 1 | 1.2 | 0.4 | 1 | 1.3 | 1 | < 10 | 1 | 1.1 | 1.2 | 1 |  | 16
| 1 | 4 | 1 | 1 | < 10 | 1.1 | 1.6 | 1.8 | 1 | 1.3 | 0.9 | < 10 | 1.1 | < 1 | 1.2 | 0.4 | 0.9 | 1.2 | 0.9 | < 10 | 1 | 1.1 | 1.3 | 0.9 |  | 1.0822
| 1 | 5 | 1.1 | 1 | < 10 | 1.3 | 1.6 | 1.7 | 1.1 | 1.3 | 1 | < 10 | 1 | < 1 | 1.2 | 0.4 | 0.9 | 1.4 | 0.9 | < 10 | 1.1 | 1.1 | 1.3 | 0.9 |  | 58
| 1 | 6 | 1.1 | < 1 | < 10 | 1.1 | 1.8 | 1.9 | 1.2 | 1.4 | 1.1 | < 10 | 1 | 1 | 1.2 | 0.4 | 0.9 | 1.3 | 0.9 | < 10 | 1.1 | 1 | 1.2 | 1
| Mean |  | 1.05 | 1 |  | 1.1333 | 1.6833 | 1.6833 | 1.1 | 1.2667 | 1 |  | 1.0167 | 1 | 1.1833 | 0.41667 | 0.91667 | 1.3 | 0.93333 |  | 1.0833 | 1.0833 | 1.2333 | 0.93333
| Median |  | 1.05 | 1 |  | 1.1 | 1.7 | 1.65 | 1.1 | 1.3 | 1 |  | 1 | 1 | 1.2 | 0.4 | 0.9 | 1.3 | 0.9 |  | 1.1 | 1.1 | 1.2 | 0.9
| Std Dev. |  | 0.054772 | 0 |  | 0.08165 | 0.075277 | 0.1472 | 0.063246 | 0.10328 | 0.089443 |  | 0.040825 |  | 0.040825 | 0.040825 | 0.040825 | 0.063246 | 0.05164 |  | 0.075277 | 0.040825 | 0.05164 | 0.05164
| Rel.Std.Dev. |  | 5.216% | 0% |  | 7.204% | 4.472% | 8.744% | 5.75% | 8.154% | 8.944% |  | 4.016% |  | 3.45% | 9.798% | 4.454% | 4.865% | 5.533% |  | 6.949% | 3.768% | 4.187% | 5.533%
| PDM3 |  | -2.977% | -7.598% |  | 4.723% | 55.54% | 55.54% | 1.643% | 17.04% | -7.598% |  | -6.057% | -7.598% | 9.343% | -61.5% | -15.3% | 20.12% | -13.76% |  | 0.1027% | 0.1027% | 13.96% | -13.76%
| Z-Score (Absolute) |  | 0.2 | NA | 15.78 | 0.13 | 2.36 | 2.36 | 0 | 0.67 | 0.4 | 15.78 | 0.34 | 2.09 | 0.34 | 2.76 | 0.74 | 0.81 | 0.67 | 15.78 | 0.07 | 0.07 | 0.54 | 0.67
Results from laboratory 2 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 236 (Certified Value 95 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 101.5 | 96.5 | 116.343 | 91.8 | 105 | 89 | 93.8 | 103.5 | 98.3 | 80 | 88.9 | 92.4 | 125.9785 | 94.2 | 94.8 | 96 | 85.2 | 94 | 89 | 102 | 94.5 | 99 | 92.4 | 1
| 1 | 2 | 97.6 | 95.5 | 117.091 | 93.6 | 99.9 | 89 | 93.66 | 102.7 | 98.1 | 82 | 89.7 | 91.6 | 125.086 | 98.5 | 96.1 | 97.1 | 86.3 | 94 | 89 | 102 | 95.6 | 101.1 | 95.3 | 8
| 1 | 3 | 94.9 | 96.5 | 117.695 | 95.7 | 103 | 93 | 92.89 | 106.4 | 98.4 | 81 | 91.3 | 91.2 | 127.687 | 95.1 | 94.2 | 88.7 | 85.3 | 96 | 87 | 102 | 96.8 | 98 | 99.3 | 16
| 1 | 4 | 96.1 | 95.5 | 118.722 | 97.2 | 100 | 91 | 93.61 | 103.3 | 98 | 81 | 93.4 | 91.6 | 126.3695 | 95 | 94.3 | 97.3 | 86.7 | 91 | 88 | 102 | 97.6 | 100 | 95 | 95.1012
| 1 | 5 | 95.1 | 96 | 115.2145 | 95.3 | 99.9 | 91 | 94.25 | 100.8 | 97.6 | 82 | 90.8 | 94.9 | 125.1495 | 95 | 93.4 | 96.7 | 84.2 | 83 | 90 | 105 | 98.2 | 99.8 | 94.4 | 59
| 1 | 6 | 96.4 | 96.5 | 117.7825 | 96.9 | 102 | 90 | 93.44 | 102.6 | 97.5 | 80 | 90 | 107 | 121.9325 | 95.1 | 93.5 | 98.2 | 84.4 | 91 | 87 | 103 | 94.7 | 99.7 | 92.8
| Mean |  | 96.9333 | 96.0833 | 117.1413 | 95.0833 | 101.6333 | 90.5 | 93.6083 | 103.2167 | 97.9833 | 81 | 90.6833 | 94.7833 | 125.3672 | 95.4833 | 94.3833 | 95.6667 | 85.35 | 91.5 | 88.3333 | 102.6667 | 96.2333 | 99.6 | 94.8667
| Median |  | 96.25 | 96.25 | 117.393 | 95.5 | 101 | 90.5 | 93.635 | 103 | 98.05 | 81 | 90.4 | 92 | 125.564 | 95.05 | 94.25 | 96.9 | 85.25 | 92.5 | 88.5 | 102 | 96.2 | 99.75 | 94.7
| Std Dev. |  | 2.4402 | 0.4916 | 1.2303 | 2.0586 | 2.0983 | 1.5166 | 0.4462 | 1.828 | 0.3656 | 0.89443 | 1.5741 | 6.1327 | 1.933 | 1.5171 | 0.99079 | 3.4886 | 0.9975 | 4.5935 | 1.2111 | 1.2111 | 1.5371 | 1.0373 | 2.4671
| Rel.Std.Dev. |  | 2.517% | 0.5116% | 1.05% | 2.165% | 2.065% | 1.676% | 0.4767% | 1.771% | 0.3731% | 1.104% | 1.736% | 6.47% | 1.542% | 1.589% | 1.05% | 3.647% | 1.169% | 5.02% | 1.371% | 1.18% | 1.597% | 1.041% | 2.601%
| PDM3 |  | 1.926% | 1.033% | 23.18% | -0.01884% | 6.869% | -4.838% | -1.57% | 8.533% | 3.031% | -14.83% | -4.645% | -0.3343% | 31.82% | 0.4018% | -0.7549% | 0.5945% | -10.25% | -3.787% | -7.117% | 7.955% | 1.19% | 4.73% | -0.2467%
| Z-Score (Absolute) |  | 0.25 | 0.1 | 3.67 | 0.07 | 1.04 | 0.84 | 0.32 | 1.31 | 0.42 | 2.45 | 0.81 | 0.12 | 5.06 | 0 | 0.19 | 0.03 | 1.72 | 0.67 | 1.21 | 1.22 | 0.13 | 0.7 | 0.1
| Analytical results for Ta in OREAS 236 (Certified Value 0.25 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.25 | 0.2 | 7.62 | 1.43 | 0.4 | 0.23 | 0.26 | 0.1 | 0.22 | < 0.1 | 0.26 | 0.21 | 0.22 | 0.2 | 0.31 | 0.26 | 0.29 | 0.21 | [5 empty] | 1
| 1 | 2 | 0.25 | 0.2 | 7.6 | 1.36 | 0.4 | 0.25 | 0.24 | 0.11 | 0.24 | < 0.1 | 0.24 | 0.21 | 0.21 | 0.2 | 0.33 | 0.26 | 0.28 | 0.22 | [5 empty] | 21
| 1 | 3 | 0.23 | 0.2 | 7.46 | 1.3 | 0.4 | 0.24 | 0.25 | 0.13 | 0.24 | < 0.1 | 0.24 | 0.22 | 0.21 | 0.2 | 0.31 | 0.28 | 0.29 | 0.23 | [5 empty] | 16
| 1 | 4 | 0.24 | 0.2 | 7.8908 | 1.15 | 0.4 | 0.25 | 0.24 | 0.12 | 0.24 | < 0.1 | 0.25 | 0.22 | 0.2 | 0.2 | 0.31 | 0.27 | 0.28 | 0.22 | [5 empty] | 0.24791
| 1 | 5 | 0.24 | 0.2 | 7.7238 | 1.53 | 0.4 | 0.25 | 0.23 | 0.1 | 0.24 | < 0.1 | 0.24 | 0.22 | 0.21 | 0.2 | 0.37 | 0.27 | 0.28 | 0.23 | [5 empty] | 60
| 1 | 6 | 0.24 | 0.2 | 7.7968 | 1.16 | 0.4 | 0.26 | 0.25 | 0.12 | 0.25 | 0.1 | 0.24 | 0.22 | 0.22 | 0.2 | 0.3 | 0.26 | 0.27 | 0.22
| Mean |  | 0.24167 | 0.2 | 7.6819 | 1.3217 | 0.4 | 0.24667 | 0.245 | 0.11333 | 0.23833 | 0.1 | 0.245 | 0.21667 | 0.21167 | 0.2 | 0.32167 | 0.26667 | 0.28167 | 0.22167
| Median |  | 0.24 | 0.2 | 7.6719 | 1.33 | 0.4 | 0.25 | 0.245 | 0.115 | 0.24 | 0.1 | 0.24 | 0.22 | 0.21 | 0.2 | 0.31 | 0.265 | 0.28 | 0.22
| Std Dev. |  | 0.0075277 | 3.0405e-17 | 0.15388 | 0.15012 | 6.0809e-17 | 0.010328 | 0.010488 | 0.012111 | 0.0098319 |  | 0.0083666 | 0.005164 | 0.0075277 | 3.0405e-17 | 0.025626 | 0.008165 | 0.0075277 | 0.0075277
| Rel.Std.Dev. |  | 3.115% | 1.52e-14% | 2.003% | 11.36% | 1.52e-14% | 4.187% | 4.281% | 10.69% | 4.125% |  | 3.415% | 2.383% | 3.556% | 1.52e-14% | 7.966% | 3.062% | 2.673% | 3.396%
| PDM3 |  | -2.518% | -19.33% | 2999% | 433.1% | 61.35% | -0.5012% | -1.173% | -54.28% | -3.863% | -59.66% | -1.173% | -12.6% | -14.62% | -19.33% | 29.75% | 7.566% | 13.62% | -10.59%
| Z-Score (Absolute) |  | 0.09 | NA | 194.12 | 28.1 | NA | 0.04 | 0 | 3.44 | 0.17 | NA | 0 | 0.74 | 0.87 | NA | 2 | 0.57 | 0.96 | 0.61
Results from laboratories 2, 7, 14 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 236 (Certified Value 0.62 ppm) | [24 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.58 | 0.63 | 0.62 | 0.58 | 0.64 | 0.6 | 0.56767 | 0.59 | 0.4 | 0.5 | 0.65 | [12 empty] | 1
| 1 | 2 | 0.6 | 0.67 | 0.65 | 0.63 | 0.62 | 0.6 | 0.55322 | 0.6 | 0.4 | 0.5 | 0.65 | [12 empty] | 22
| 1 | 3 | 0.6 | 0.66 | 0.65 | 0.6 | 0.64 | 0.6 | 0.55684 | 0.6 | 0.4 | 0.5 | 0.66 | [12 empty] | 16
| 1 | 4 | 0.6 | 0.68 | 0.63 | 0.6 | 0.63 | 0.6 | 0.56208 | 0.61 | 0.4 | 0.5 | 0.66 | [12 empty] | 0.61604
| 1 | 5 | 0.6 | 0.65 | 0.62 | 0.61 | 0.62 | 0.6 | 0.52265 | 0.58 | 0.4 | 0.5 | 0.64 | [12 empty] | 61
| 1 | 6 | 0.6 | 0.72 | 0.66 | 0.61 | 0.64 | 0.6 | 0.58345 | 0.59 | 0.4 | 0.5 | 0.65
| Mean |  | 0.59667 | 0.66833 | 0.63833 | 0.605 | 0.63167 | 0.6 | 0.55765 | 0.595 | 0.4 | 0.5 | 0.65167
| Median |  | 0.6 | 0.665 | 0.64 | 0.605 | 0.635 | 0.6 | 0.55946 | 0.595 | 0.4 | 0.5 | 0.65
| Std Dev. |  | 0.008165 | 0.030605 | 0.017224 | 0.016432 | 0.0098319 | 0 | 0.020153 | 0.010488 | 6.0809e-17 | 0 | 0.0075277
| Rel.Std.Dev. |  | 1.368% | 4.579% | 2.698% | 2.716% | 1.557% | 0% | 3.614% | 1.763% | 1.52e-14% | 0% | 1.155%
| PDM3 |  | -3.144% | 8.489% | 3.62% | -1.791% | 2.537% | -2.603% | -9.477% | -3.415% | -35.07% | -18.84% | 5.784%
| Z-Score (Absolute) |  | 0.21 | 1.6 | 0.84 | 0 | 0.67 | 0.13 | 1.2 | 0.25 | NA | NA | 1.18
Results from laboratories 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 236 (Certified Value 0.11 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 0.12 | < 0.2 | 10.23 | 0.13 | < 0.1 | < 0.2 | < 0.01 | < 0.1 | 0.1 | < 0.1 | 0.11 | 0.11 | 0.1 | 0.1 | < 5 | 0.13 | 0.11 | 0.1 | 0.11 |  |  |  |  | 1
| 1 | 2 | 0.11 | < 0.2 | 10.55 | 0.12 | 0.1 | < 0.2 | 0.16 | 0.1 | < 0.1 | < 0.1 | 0.12 | 0.11 | 0.1 | 0.08 | < 5 | 0.15 | 0.1 | 0.1 | 0.12 |  |  |  |  | 23
| 1 | 3 | 0.1 | < 0.2 | 10.355 | 0.12 | < 0.1 | < 0.2 | 0.14 | 0.1 | 0.1 | < 0.1 | 0.1 | 0.1 | 0.1 | 0.11 | < 5 | 0.15 | 0.14 | 0.08 | 0.1 |  |  |  |  | 16
| 1 | 4 | 0.1 | < 0.2 | 10.4878 | 0.1 | < 0.1 | < 0.2 | < 0.01 | < 0.1 | < 0.1 | < 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | < 5 | 0.12 | 0.13 | 0.12 | 0.11 |  |  |  |  | 0.11278
| 1 | 5 | 0.12 | < 0.2 | 10.5396 | 0.11 | < 0.1 | < 0.2 | 0.26 | 0.1 | < 0.1 | < 0.1 | 0.12 | 0.1 | 0.1 | 0.08 | < 5 | 0.14 | 0.15 | 0.12 | 0.12 |  |  |  |  | 62
| 1 | 6 | 0.14 | < 0.2 | 10.4625 | 0.1 | 0.1 | < 0.2 | 0.23 | 0.1 | 0.1 | < 0.1 | 0.09 | 0.11 | 0.1 | 0.08 | < 5 | 0.13 | 0.14 | 0.1 | 0.12
| Mean |  | 0.115 |  | 10.4375 | 0.11333 | 0.1 |  | 0.1975 | 0.1 | 0.1 |  | 0.10833 | 0.105 | 0.1 | 0.091667 |  | 0.13667 | 0.12833 | 0.10333 | 0.11333
| Median |  | 0.115 |  | 10.4752 | 0.115 | 0.1 |  | 0.195 | 0.1 | 0.1 |  | 0.11 | 0.105 | 0.1 | 0.09 |  | 0.135 | 0.135 | 0.1 | 0.115
| Std Dev. |  | 0.015166 |  | 0.12337 | 0.012111 | 0 |  | 0.056789 | 0 | 1.6997e-17 |  | 0.01169 | 0.0054772 | 1.5202e-17 | 0.013292 |  | 0.012111 | 0.019408 | 0.015055 | 0.008165
| Rel.Std.Dev. |  | 13.19% |  | 1.182% | 10.69% | 0% |  | 28.75% | 0% | 1.7e-14% |  | 10.79% | 5.216% | 1.52e-14% | 14.5% |  | 8.861% | 15.12% | 14.57% | 7.204%
| PDM3 |  | 1.97% |  | 9155% | 0.4926% | -11.33% |  | 75.12% | -11.33% | -11.33% |  | -3.941% | -6.897% | -11.33% | -18.72% |  | 21.18% | 13.79% | -8.374% | 0.4926%
| Z-Score (Absolute) |  | 0.08 | 0.67 | 522.13 | 0 | NA | 0.67 | 1.01 | NA | NA | 3.2 | 0.25 | 0.42 | NA | 1.1 | 120.7 | 1.18 | 0.76 | 0.51 | 0
| Analytical results for Th in OREAS 236 (Certified Value 1.1 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 1.09 | 1.1 | 1.2 | 1.1 | 1.1 | 1.12 | 1.2 | 1.01 | 0.99 | 1.1 | 0.8954 | 1.1 | 1.3 | 0.97 | 1.09 | 1 | 1.25 | 1.11 | 1.2 | 0.99 |  |  |  | 1
| 1 | 2 | 1.07 | 1.2 | 1.3 | 1.1 | 1.1 | 1.11 | 1.1 | 0.96 | 0.97 | 1.1 | 0.85617 | 1.1 | 1.4 | 0.97 | 1.1 | 1 | 1.26 | 1.09 | 1.3 | 1.02 |  |  |  | 24
| 1 | 3 | 1.02 | 1.1 | 1.2 | 1.1 | 1.1 | 1.11 | 1.1 | 0.98 | 1.02 | 1.1 | 0.94295 | 1.1 | 1.3 | 0.93 | 1.14 | 1 | 1.2 | 1.13 | 1.2 | 1.08 |  |  |  | 16
| 1 | 4 | 1.04 | 1.2 | 1.3 | 1.1 | 1.1 | 1.1 | 1.1 | 0.97 | 1.05 | 1.1 | 0.96305 | 1.1 | 1.3 | 1.01 | 1.07 | 0.9 | 1.24 | 1.16 | 1.3 | 1.02 |  |  |  | 1.1001
| 1 | 5 | 1.02 | 1.2 | 1.2 | 1 | 1.1 | 1.1 | 1.1 | 1.02 | 1 | 1.1 | 0.95082 | 1.2 | 1.3 | 0.98 | 1.06 | 0.8 | 1.22 | 1.15 | 1.3 | 1.02 |  |  |  | 63
| 1 | 6 | 1.04 | 1.2 | 1.6 | 1.1 | 1.1 | 1.15 | 1.1 | 1.04 | 0.99 | 1.2 | 0.90105 | 1.1 | 1.4 | 1 | 1.13 | 0.9 | 1.21 | 1.08 | 1.3 | 1.05
| Mean |  | 1.0467 | 1.1667 | 1.3 | 1.0833 | 1.1 | 1.115 | 1.1167 | 0.99667 | 1.0033 | 1.1167 | 0.91824 | 1.1167 | 1.3333 | 0.97667 | 1.0983 | 0.93333 | 1.23 | 1.12 | 1.2667 | 1.03
| Median |  | 1.04 | 1.2 | 1.25 | 1.1 | 1.1 | 1.11 | 1.1 | 0.995 | 0.995 | 1.1 | 0.922 | 1.1 | 1.3 | 0.975 | 1.095 | 0.95 | 1.23 | 1.12 | 1.3 | 1.02
| Std Dev. |  | 0.028048 | 0.05164 | 0.15492 | 0.040825 | 2.4324e-16 | 0.018708 | 0.040825 | 0.031411 | 0.028048 | 0.040825 | 0.040866 | 0.040825 | 0.05164 | 0.028048 | 0.031885 | 0.08165 | 0.023664 | 0.032249 | 0.05164 | 0.030984
| Rel.Std.Dev. |  | 2.68% | 4.426% | 11.92% | 3.768% | 2.211e-14% | 1.678% | 3.656% | 3.152% | 2.795% | 3.656% | 4.45% | 3.656% | 3.873% | 2.872% | 2.903% | 8.748% | 1.924% | 2.879% | 4.077% | 3.008%
| PDM3 |  | -4.854% | 6.055% | 18.18% | -1.521% | -0.005634% | 1.358% | 1.509% | -9.399% | -8.793% | 1.509% | -16.53% | 1.509% | 21.21% | -11.22% | -0.1571% | -15.16% | 11.81% | 1.812% | 15.15% | -6.369%
| Z-Score (Absolute) |  | 0.59 | 0.58 | 1.88 | 0.24 | 0.07 | 0.07 | 0.09 | 1.08 | 1.02 | 0.09 | 1.85 | 0.09 | 2.2 | 1.28 | 0.09 | 1.7 | 1.19 | 0.12 | 1.55 | 0.76
| Analytical results for Ti in OREAS 236 (Certified Value 0.639 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.644 | 0.67 | 0.44 | 0.61 | 0.69 | 0.655 | 0.638 | 0.59 | 0.6552 | 0.6397 | 0.266 | 0.68194 | 0.63 | 0.634 | 0.599 | 0.63 | 0.62 | 0.63 | 0.629 | 0.66 | 0.607 |  |  | 1
| 1 | 2 | 0.615 | 0.68 | 0.445 | 0.64 | 0.66 | 0.655 | 0.6465 | 0.58 | 0.6622 | 0.6515 | 0.244 | 0.67627 | 0.63 | 0.634 | 0.596 | 0.656 | 0.62 | 0.62 | 0.629 | 0.67 | 0.627 |  |  | 25
| 1 | 3 | 0.605 | 0.69 | 0.44 | 0.62 | 0.69 | 0.644 | 0.631 | 0.59 | 0.6642 | 0.6504 | 0.346 | 0.68967 | 0.63 | 0.606 | 0.606 | 0.65 | 0.61 | 0.63 | 0.637 | 0.65 | 0.633 |  |  | 16
| 1 | 4 | 0.606 | 0.68 | 0.445 | 0.64 | 0.67 | 0.655 | 0.6346 | 0.59 | 0.6564 | 0.6528 | 0.362 | 0.68775 | 0.63 | 0.627 | 0.607 | 0.634 | 0.61 | 0.63 | 0.661 | 0.67 | 0.612 |  |  | 0.63857
| 1 | 5 | 0.601 | 0.67 | 0.445 | 0.64 | 0.66 | 0.646 | 0.6517 | 0.61 | 0.6662 | 0.6426 | 0.394 | 0.68129 | 0.62 | 0.632 | 0.603 | 0.57 | 0.63 | 0.63 | 0.635 | 0.66 | 0.614 |  |  | 64
| 1 | 6 | 0.609 | 0.68 | 0.445 | 0.66 | 0.68 | 0.648 | 0.6493 | 0.6 | 0.6628 | 0.64 | 0.579 | 0.67914 | 0.63 | 0.648 | 0.601 | 0.623 | 0.61 | 0.63 | 0.632 | 0.67 | 0.609
| Mean |  | 0.61333 | 0.67833 | 0.44333 | 0.635 | 0.675 | 0.6505 | 0.64185 | 0.59333 | 0.66117 | 0.64617 | 0.36517 | 0.68268 | 0.62833 | 0.63017 | 0.602 | 0.62717 | 0.61667 | 0.62833 | 0.63717 | 0.66333 | 0.617
| Median |  | 0.6075 | 0.68 | 0.445 | 0.64 | 0.675 | 0.6515 | 0.64225 | 0.59 | 0.6625 | 0.6465 | 0.354 | 0.68161 | 0.63 | 0.633 | 0.602 | 0.632 | 0.615 | 0.63 | 0.6335 | 0.665 | 0.613
| Std Dev. |  | 0.015731 | 0.0075277 | 0.002582 | 0.017607 | 0.013784 | 0.0050892 | 0.0084765 | 0.010328 | 0.0043953 | 0.0060487 | 0.11954 | 0.005112 | 0.0040825 | 0.013747 | 0.0041952 | 0.030636 | 0.008165 | 0.0040825 | 0.012106 | 0.008165 | 0.010526
| Rel.Std.Dev. |  | 2.565% | 1.11% | 0.5824% | 2.773% | 2.042% | 0.7824% | 1.321% | 1.741% | 0.6648% | 0.9361% | 32.74% | 0.7488% | 0.6497% | 2.181% | 0.6969% | 4.885% | 1.324% | 0.6497% | 1.9% | 1.231% | 1.706%
| PDM3 |  | -3.953% | 6.226% | -30.57% | -0.5596% | 5.704% | 1.868% | 0.5131% | -7.085% | 3.538% | 1.189% | -42.82% | 6.907% | -1.604% | -1.316% | -5.727% | -1.786% | -3.431% | -1.604% | -0.2203% | 3.877% | -3.378%
| Z-Score (Absolute) |  | 0.67 | 1.93 | 7.48 | 0.19 | 1.8 | 0.81 | 0.47 | 1.48 | 1.24 | 0.64 | 10.62 | 2.1 | 0.07 | 0 | 1.13 | 0.12 | 0.54 | 0.07 | 0.28 | 1.33 | 0.53
| Analytical results for Tl in OREAS 236 (Certified Value 0.24 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.26 | 0.3 | 0.22 | 0.2 | 0.3 | 0.26 | 0.26 | 0.24 | 0.25 | 0.23 | 0.26957 | 0.25 | 0.1 | 0.22 | 0.22 | 0.23 | < 5 | 0.13 | 0.23 | 0.25 | 0.23 |  |  | 1
| 1 | 2 | 0.26 | 0.3 | 0.24 | 0.3 | 0.3 | 0.26 | 0.27 | 0.26 | 0.24 | 0.23 | 0.27761 | 0.25 | 0.1 | 0.21 | 0.23 | 0.23 | < 5 | 0.12 | 0.25 | 0.24 | 0.24 |  |  | 26
| 1 | 3 | 0.22 | 0.3 | 0.23 | 0.2 | 0.3 | 0.26 | 0.26 | 0.26 | 0.26 | 0.23 | 0.25671 | 0.24 | 0.1 | 0.22 | 0.2 | 0.24 | < 5 | 0.08 | 0.24 | 0.24 | 0.24 |  |  | 16
| 1 | 4 | 0.24 | 0.3 | 0.23 | 0.3 | 0.3 | 0.26 | 0.27 | 0.25 | 0.25 | 0.22 | 0.278 | 0.25 | 0.1 | 0.22 | 0.23 | 0.22 | < 5 | 0.06 | 0.25 | 0.24 | 0.23 |  |  | 0.24219
| 1 | 5 | 0.22 | 0.3 | 0.23 | 0.3 | 0.3 | 0.26 | 0.26 | 0.25 | 0.25 | 0.23 | 0.2531 | 0.27 | 0.1 | 0.22 | 0.2 | 0.2 | < 5 | 0.16 | 0.23 | 0.26 | 0.24 |  |  | 65
| 1 | 6 | 0.24 | 0.3 | 0.24 | 0.3 | 0.3 | 0.27 | 0.28 | 0.24 | 0.25 | 0.26 | 0.27439 | 0.25 | 0.1 | 0.22 | 0.22 | 0.23 | < 5 | 0.12 | 0.23 | 0.24 | 0.24
| Mean |  | 0.24 | 0.3 | 0.23167 | 0.26667 | 0.3 | 0.26167 | 0.26667 | 0.25 | 0.25 | 0.23333 | 0.26823 | 0.25167 | 0.1 | 0.21833 | 0.21667 | 0.225 |  | 0.11167 | 0.23833 | 0.245 | 0.23667
| Median |  | 0.24 | 0.3 | 0.23 | 0.3 | 0.3 | 0.26 | 0.265 | 0.25 | 0.25 | 0.23 | 0.27198 | 0.25 | 0.1 | 0.22 | 0.22 | 0.23 |  | 0.12 | 0.235 | 0.24 | 0.24
| Std Dev. |  | 0.017889 | 0 | 0.0075277 | 0.05164 | 0 | 0.0040825 | 0.008165 | 0.0089443 | 0.0063246 | 0.013663 | 0.010816 | 0.0098319 | 1.5202e-17 | 0.0040825 | 0.013663 | 0.013784 |  | 0.036009 | 0.0098319 | 0.0083666 | 0.005164
| Rel.Std.Dev. |  | 7.454% | 0% | 3.249% | 19.36% | 0% | 1.56% | 3.062% | 3.578% | 2.53% | 5.855% | 4.032% | 3.907% | 1.52e-14% | 1.87% | 6.306% | 6.126% |  | 32.25% | 4.125% | 3.415% | 2.182%
| PDM3 |  | -0.9055% | 23.87% | -4.346% | 10.1% | 23.87% | 8.041% | 10.1% | 3.223% | 3.223% | -3.658% | 10.75% | 3.912% | -58.71% | -9.852% | -10.54% | -7.099% |  | -53.89% | -1.594% | 1.159% | -2.282%
| Z-Score (Absolute) |  | 0.12 | NA | 0.52 | NA | 2.74 | 0.91 | 1.15 | 0.36 | 0.36 | 0.44 | 1.22 | 0.44 | NA | 1.15 | 1.23 | 0.83 | 107.45 | 6.23 | 0.2 | 0.12 | 0.28
Results from laboratories 2, 5 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 236 (Certified Value 0.36 ppm) | [24 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.02 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.36 | 0.36 | 0.34 | 0.39 | 0.3 | 0.32109 | 0.35 | 0.3 | 0.3 | 0.37 | [13 empty] | 1
| 1 | 2 | 0.34 | 0.37 | 0.35 | 0.38 | 0.3 | 0.3154 | 0.35 | 0.3 | 0.3 | 0.38 | [13 empty] | 27
| 1 | 3 | 0.36 | 0.36 | 0.35 | 0.39 | 0.3 | 0.32018 | 0.35 | 0.3 | 0.3 | 0.39 | [13 empty] | 16
| 1 | 4 | 0.36 | 0.36 | 0.36 | 0.37 | 0.3 | 0.32388 | 0.35 | 0.3 | 0.3 | 0.39 | [13 empty] | 0.35807
| 1 | 5 | 0.34 | 0.35 | 0.35 | 0.38 | 0.3 | 0.29059 | 0.35 | 0.3 | 0.3 | 0.4 | [13 empty] | 66
| 1 | 6 | 0.36 | 0.36 | 0.37 | 0.4 | 0.4 | 0.31871 | 0.36 | 0.4 | 0.3 | 0.37
| Mean |  | 0.35333 | 0.36 | 0.35333 | 0.385 | 0.31667 | 0.31498 | 0.35167 | 0.31667 | 0.3 | 0.38333
| Median |  | 0.36 | 0.36 | 0.35 | 0.385 | 0.3 | 0.31945 | 0.35 | 0.3 | 0.3 | 0.385
| Std Dev. |  | 0.010328 | 0.0063246 | 0.010328 | 0.010488 | 0.040825 | 0.012268 | 0.0040825 | 0.040825 | 0 | 0.012111
| Rel.Std.Dev. |  | 2.923% | 1.757% | 2.923% | 2.724% | 12.89% | 3.895% | 1.161% | 12.89% | 0% | 3.159%
| PDM3 |  | -1.324% | 0.5378% | -1.324% | 7.52% | -11.56% | -12.04% | -1.789% | -11.56% | -16.22% | 7.054%
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 3.2 | NA | 3.88 | 0.17 | NA | NA | 3.03
Results from laboratories 14, 17 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 236 (Certified Value 0.35 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.3 | 0.3 | 0.4 | 0.36 | 0.4 | 0.35 | 0.39 | 0.31 | 0.33 | 0.4 | 0.40474 | 0.35 | 0.4 | 0.3 | 0.35 | 0.3 | < 10 | 0.4 | 0.3 | 0.4 | 0.3 |  |  | 1
| 1 | 2 | 0.3 | 0.3 | 0.4 | 0.38 | 0.4 | 0.37 | 0.35 | 0.3 | 0.33 | 0.3 | 0.40602 | 0.37 | 0.4 | 0.3 | 0.32 | 0.3 | < 10 | 0.4 | 0.3 | 0.4 | 0.3 |  |  | 28
| 1 | 3 | 0.3 | 0.3 | 0.4 | 0.37 | 0.4 | 0.35 | 0.38 | 0.32 | 0.36 | 0.3 | 0.41756 | 0.36 | 0.3 | 0.3 | 0.35 | 0.3 | < 10 | 0.4 | 0.3 | 0.3 | 0.4 |  |  | 16
| 1 | 4 | 0.3 | 0.3 | 0.4 | 0.37 | 0.4 | 0.37 | 0.38 | 0.3 | 0.34 | 0.3 | 0.4173 | 0.37 | 0.3 | 0.3 | 0.34 | 0.3 | < 10 | 0.4 | 0.3 | 0.4 | 0.3 |  |  | 0.3461
| 1 | 5 | 0.3 | 0.3 | 0.4 | 0.37 | 0.4 | 0.36 | 0.35 | 0.32 | 0.32 | 0.4 | 0.39654 | 0.38 | 0.3 | 0.3 | 0.35 | 0.3 | < 10 | 0.4 | 0.3 | 0.4 | 0.3 |  |  | 67
| 1 | 6 | 0.3 | 0.3 | 0.3 | 0.38 | 0.4 | 0.36 | 0.37 | 0.32 | 0.34 | 0.5 | 0.42935 | 0.36 | 0.3 | 0.3 | 0.32 | 0.3 | < 10 | 0.4 | 0.3 | 0.4 | 0.3
| Mean |  | 0.3 | 0.3 | 0.38333 | 0.37167 | 0.4 | 0.36 | 0.37 | 0.31167 | 0.33667 | 0.36667 | 0.41192 | 0.365 | 0.33333 | 0.3 | 0.33833 | 0.3 |  | 0.4 | 0.3 | 0.38333 | 0.31667
| Median |  | 0.3 | 0.3 | 0.4 | 0.37 | 0.4 | 0.36 | 0.375 | 0.315 | 0.335 | 0.35 | 0.41166 | 0.365 | 0.3 | 0.3 | 0.345 | 0.3 |  | 0.4 | 0.3 | 0.4 | 0.3
| Std Dev. |  | 0 | 0 | 0.040825 | 0.0075277 | 6.0809e-17 | 0.0089443 | 0.016733 | 0.0098319 | 0.013663 | 0.08165 | 0.011725 | 0.010488 | 0.05164 | 0 | 0.01472 | 0 |  | 6.0809e-17 | 0 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 0% | 0% | 10.65% | 2.025% | 1.52e-14% | 2.485% | 4.522% | 3.155% | 4.058% | 22.27% | 2.846% | 2.873% | 15.49% | 0% | 4.351% | 0% |  | 1.52e-14% | 0% | 10.65% | 12.89%
| PDM3 |  | -13.32% | -13.32% | 10.76% | 7.388% | 15.57% | 4.017% | 6.907% | -9.948% | -2.724% | 5.944% | 19.02% | 5.462% | -3.688% | -13.32% | -2.243% | -13.32% |  | 15.57% | -13.32% | 10.76% | -8.503%
| Z-Score (Absolute) |  | 1.01 | 1.01 | 0.39 | 0.2 | 0.67 | 0 | 0.17 | 0.81 | 0.39 | 0.11 | 0.88 | 0.08 | 0.45 | 1.01 | 0.37 | 1.01 | 78.22 | 0.67 | 1.01 | 0.39 | 0.73
| Analytical results for V in OREAS 236 (Certified Value 279 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 277 | 290 | 228.2573 | 262 | 315 | 305 | 270 | 292 | 283 | 257 | 280 | 138 | 269.1093 | 269 | 274 | 267 | 288 | 276 | 277 | 256 | 287 | 263 |  | 1
| 1 | 2 | 266 | 300 | 228.54 | 267 | 303 | 308 | 280 | 285 | 299 | 258 | 286 | 141 | 281.7117 | 266 | 273 | 264 | 290 | 272 | 280 | 260 | 291 | 272 |  | 29
| 1 | 3 | 260 | 290 | 228.087 | 262 | 320 | 310 | 270 | 292 | 299 | 254 | 287 | 186 | 287.9328 | 272 | 256 | 267 | 295 | 267 | 280 | 265 | 284 | 273 |  | 16
| 1 | 4 | 262 | 285 | 230.187 | 268 | 312 | 311 | 268 | 293 | 293 | 254 | 286 | 198 | 287.4077 | 273 | 277 | 266 | 283 | 273 | 277 | 272 | 290 | 268 |  | 279.1667
| 1 | 5 | 259 | 290 | 230.383 | 273 | 302 | 312 | 276 | 296 | 299 | 259 | 279 | 211 | 275.3677 | 267 | 270 | 263 | 256 | 277 | 278 | 262 | 289 | 267 |  | 68
| 1 | 6 | 263 | 275 | 229.032 | 265 | 311 | 314 | 275 | 294 | 284 | 260 | 282 | 282 | 283.2781 | 267 | 280 | 265 | 281 | 270 | 278 | 264 | 289 | 267
| Mean |  | 264.5 | 288.3333 | 229.0811 | 266.1667 | 310.5 | 310 | 273.1667 | 292 | 292.8333 | 257 | 283.3333 | 192.6667 | 280.8012 | 269 | 271.6667 | 265.3333 | 282.1667 | 272.5 | 278.3333 | 263.1667 | 288.3333 | 268.3333
| Median |  | 262.5 | 290 | 228.786 | 266 | 311.5 | 310.5 | 272.5 | 292.5 | 296 | 257.5 | 284 | 192 | 282.4949 | 268 | 273.5 | 265.5 | 285.5 | 272.5 | 278 | 263 | 289 | 267.5
| Std Dev. |  | 6.5955 | 8.165 | 0.98793 | 4.1673 | 6.9498 | 3.1623 | 4.5789 | 3.7417 | 7.6004 | 2.5298 | 3.4448 | 53.0044 | 7.317 | 2.8983 | 8.4063 | 1.633 | 13.7611 | 3.7283 | 1.3663 | 5.3821 | 2.5033 | 3.6697
| Rel.Std.Dev. |  | 2.494% | 2.832% | 0.4313% | 1.566% | 2.238% | 1.02% | 1.676% | 1.281% | 2.595% | 0.9844% | 1.216% | 27.51% | 2.606% | 1.077% | 3.094% | 0.6154% | 4.877% | 1.368% | 0.4909% | 2.045% | 0.8682% | 1.368%
| PDM3 |  | -5.254% | 3.284% | -17.94% | -4.657% | 11.22% | 11.04% | -2.149% | 4.597% | 4.895% | -7.94% | 1.493% | -30.99% | 0.5855% | -3.642% | -2.687% | -4.955% | 1.075% | -2.388% | -0.2985% | -5.731% | 3.284% | -3.881%
| Z-Score (Absolute) |  | 0.59 | 1.1 | 3.11 | 0.47 | 2.67 | 2.64 | 0.02 | 1.36 | 1.42 | 1.12 | 0.75 | 5.69 | 0.57 | 0.27 | 0.08 | 0.53 | 0.66 | 0.02 | 0.39 | 0.69 | 1.1 | 0.32
| Analytical results for W in OREAS 236 (Certified Value 30.5 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 32 | 32 | 37.9601 | 28.9 | 28.8 | 31.9 | 32.67 | 20.6 | 28.5 | 1 | 30.3 | 30.3 | 28.1 | 28.6 | 25 | 30.6 | 31.1 | 31.4 | 29.4 |  |  |  |  | 1
| 1 | 2 | 31.3 | 32.5 | 37.1894 | 30.1 | 28.7 | 31.8 | 32 | 22.2 | 29.4 | 0.8 | 31.3 | 29.9 | 29.4 | 28.2 | 23 | 31.7 | 32.4 | 31.8 | 31.1 |  |  |  |  | 30
| 1 | 3 | 29.4 | 31.5 | 38.1577 | 29.5 | 28.4 | 31.6 | 33.01 | 20.9 | 29.4 | 1.5 | 30.9 | 29.2 | 29.1 | 28.9 | 25 | 30.6 | 33.3 | 31.2 | 31.4 |  |  |  |  | 16
| 1 | 4 | 30.3 | 32 | 38.4538 | 30.4 | 28.5 | 31.1 | 31.79 | 22.8 | 29.8 | 1.2 | 31.2 | 30.2 | 29.1 | 26.8 | 22 | 30.9 | 32.7 | 31.9 | 30 |  |  |  |  | 30.5099
| 1 | 5 | 29.6 | 32.5 | 37.7641 | 29.7 | 28.9 | 32 | 31.7 | 22.4 | 29.6 | 2.1 | 33.3 | 30.1 | 28.9 | 24.9 | 29 | 34.7 | 33.1 | 32 | 30.5 |  |  |  |  | 69
| 1 | 6 | 29.5 | 32.5 | 37.27 | 26.3 | 29.2 | 32 | 31.94 | 22.9 | 29.1 | 8.9 | 31.1 | 30.7 | 29.5 | 27.5 | 21 | 30.2 | 31.7 | 31.7 | 30.9
| Mean |  | 30.35 | 32.1667 | 37.7992 | 29.15 | 28.75 | 31.7333 | 32.185 | 21.9667 | 29.3 | 2.5833 | 31.35 | 30.0667 | 29.0167 | 27.4833 | 24.1667 | 31.45 | 32.3833 | 31.6667 | 30.55
| Median |  | 29.95 | 32.25 | 37.8621 | 29.6 | 28.75 | 31.85 | 31.97 | 22.3 | 29.4 | 1.35 | 31.15 | 30.15 | 29.1 | 27.85 | 24 | 30.75 | 32.55 | 31.75 | 30.7
| Std Dev. |  | 1.0784 | 0.40825 | 0.49721 | 1.4883 | 0.2881 | 0.34448 | 0.52941 | 0.98116 | 0.45607 | 3.1276 | 1.0193 | 0.50067 | 0.49967 | 1.477 | 2.8577 | 1.6694 | 0.84479 | 0.30768 | 0.74498
| Rel.Std.Dev. |  | 3.553% | 1.269% | 1.315% | 5.106% | 1.002% | 1.086% | 1.645% | 4.467% | 1.557% | 121.1% | 3.251% | 1.665% | 1.722% | 5.374% | 11.83% | 5.308% | 2.609% | 0.9716% | 2.439%
| PDM3 |  | -0.5241% | 5.43% | 23.89% | -4.457% | -5.768% | 4.01% | 5.49% | -28% | -3.966% | -91.53% | 2.754% | -1.453% | -4.894% | -9.92% | -20.79% | 3.081% | 6.14% | 3.791% | 0.1315%
| Z-Score (Absolute) |  | 0 | 0.89 | 3.63 | 0.58 | 0.78 | 0.67 | 0.89 | 4.09 | 0.51 | 13.53 | 0.49 | 0.14 | 0.65 | 1.4 | 3.01 | 0.54 | 0.99 | 0.64 | 0.1
| Analytical results for Y in OREAS 236 (Certified Value 21.7 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 23.6 | 21.3 | 16.645 | 19.9 | 23.4 | 22.1 | 20.66 | 22.95 | 21.1 | 17.88 | 20.9 | 21.5 | 21.5254 | 21.5 | 16.3 | 20.7 | 21 | 21.1 | 22 | 22.7 | 21.1 | 22.2 | 20.9 | 1
| 1 | 2 | 22.2 | 21.3 | 16.715 | 20.5 | 24.5 | 21.9 | 20.49 | 22.12 | 21.1 | 17.99 | 21 | 21.2 | 21.3971 | 20.9 | 15.1 | 20.9 | 21.8 | 20.5 | 22 | 22.9 | 21.5 | 22 | 21.7 | 31
| 1 | 3 | 21.7 | 21.6 | 16.755 | 21 | 23.6 | 22.6 | 20.66 | 23.13 | 20.7 | 17.66 | 21.1 | 21.1 | 21.2891 | 21.3 | 17.4 | 21.3 | 22 | 21.4 | 22 | 22.9 | 21.8 | 21.6 | 22 | 16
| 1 | 4 | 23.7 | 21.6 | 16.77 | 21.2 | 23.4 | 22.3 | 20.91 | 22.18 | 20.7 | 17.93 | 21.8 | 21.3 | 21.0623 | 21.5 | 16.8 | 21.2 | 20.8 | 20.5 | 22 | 22.9 | 21.9 | 23.2 | 24.1 | 21.654
| 1 | 5 | 21.7 | 21.4 | 16.595 | 20.8 | 23.4 | 22.3 | 20.76 | 22.27 | 20.9 | 17.76 | 21.1 | 21.3 | 21.3283 | 21 | 19 | 21.3 | 21.9 | 17.7 | 22 | 22.8 | 22.5 | 22.5 | 21.3 | 70
| 1 | 6 | 23.1 | 21.3 | 16.81 | 21.2 | 23.7 | 22.6 | 20.63 | 22.86 | 20.8 | 17.97 | 21 | 22.2 | 21.5136 | 21.1 | 20.4 | 21.2 | 21.5 | 20.1 | 22 | 22.8 | 21.3 | 22 | 20.7
| Mean |  | 22.6667 | 21.4167 | 16.715 | 20.7667 | 23.6667 | 22.3 | 20.685 | 22.585 | 20.8833 | 17.865 | 21.15 | 21.4333 | 21.3526 | 21.2167 | 17.5 | 21.1 | 21.5 | 20.2167 | 22 | 22.8333 | 21.6833 | 22.25 | 21.7833
| Median |  | 22.65 | 21.35 | 16.735 | 20.9 | 23.5 | 22.3 | 20.66 | 22.565 | 20.85 | 17.905 | 21.05 | 21.3 | 21.3627 | 21.2 | 17.1 | 21.2 | 21.65 | 20.5 | 22 | 22.85 | 21.65 | 22.1 | 21.5
| Std Dev. |  | 0.91797 | 0.1472 | 0.081179 | 0.50067 | 0.4274 | 0.27568 | 0.14039 | 0.44393 | 0.18348 | 0.12973 | 0.32711 | 0.39833 | 0.17126 | 0.25626 | 1.9162 | 0.24495 | 0.498 | 1.3182 | 0 | 0.08165 | 0.49967 | 0.55045 | 1.2336
| Rel.Std.Dev. |  | 4.05% | 0.6873% | 0.4857% | 2.411% | 1.806% | 1.236% | 0.6787% | 1.966% | 0.8786% | 0.7262% | 1.547% | 1.858% | 0.802% | 1.208% | 10.95% | 1.161% | 2.316% | 6.52% | 0% | 0.3576% | 2.304% | 2.474% | 5.663%
| PDM3 |  | 4.677% | -1.096% | -22.81% | -4.098% | 9.295% | 2.983% | -4.475% | 4.3% | -3.559% | -17.5% | -2.327% | -1.019% | -1.392% | -2.019% | -19.18% | -2.558% | -0.711% | -6.638% | 1.598% | 5.446% | 0.1356% | 2.753% | 0.5974%
| Z-Score (Absolute) |  | 1.3 | 0 | 4.88 | 0.67 | 2.33 | 0.92 | 0.76 | 1.21 | 0.55 | 3.68 | 0.28 | 0.02 | 0.07 | 0.21 | 4.06 | 0.33 | 0.09 | 1.24 | 0.61 | 1.47 | 0.28 | 0.86 | 0.38
| Analytical results for Yb in OREAS 236 (Certified Value 2.32 ppm) | [24 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.24 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 2.35 | 2.2 | 2.3 | 2.29 | 2.66 | 2.4 | 2.179 | 2.3 | 1.3 | 2.1 | 2.52 | [12 empty] | 1
| 1 | 2 | 2.4 | 2.3 | 2.4 | 2.29 | 2.61 | 2.3 | 2.1371 | 2.2 | 1.3 | 2.1 | 2.54 | [12 empty] | 32
| 1 | 3 | 2.35 | 2.3 | 2.4 | 2.28 | 2.57 | 2.3 | 2.1505 | 2.4 | 1.5 | 2.1 | 2.54 | [12 empty] | 16
| 1 | 4 | 2.35 | 2.3 | 2.3 | 2.29 | 2.54 | 2.3 | 2.1472 | 2.3 | 1.4 | 2 | 2.53 | [12 empty] | 2.3228
| 1 | 5 | 2.35 | 2.3 | 2.4 | 2.28 | 2.54 | 2.3 | 1.9995 | 2.2 | 1.4 | 1.8 | 2.54 | [12 empty] | 71
| 1 | 6 | 2.35 | 2.3 | 2.4 | 2.26 | 2.58 | 2.6 | 2.1689 | 2.3 | 1.5 | 2 | 2.54
| Mean |  | 2.3583 | 2.2833 | 2.3667 | 2.2817 | 2.5833 | 2.3667 | 2.1304 | 2.2833 | 1.4 | 2.0167 | 2.535
| Median |  | 2.35 | 2.3 | 2.4 | 2.285 | 2.575 | 2.3 | 2.1488 | 2.3 | 1.4 | 2.05 | 2.54
| Std Dev. |  | 0.020412 | 0.040825 | 0.05164 | 0.01169 | 0.045898 | 0.12111 | 0.065892 | 0.075277 | 0.089443 | 0.1169 | 0.0083666
| Rel.Std.Dev. |  | 0.8655% | 1.788% | 2.182% | 0.5124% | 1.777% | 5.117% | 3.093% | 3.297% | 6.389% | 5.797% | 0.33%
| PDM3 |  | 1.529% | -1.7% | 1.888% | -1.772% | 11.22% | 1.888% | -8.285% | -1.7% | -39.73% | -13.18% | 9.135%
| Z-Score (Absolute) |  | 0.61 | 0 | 0.67 | 0.01 | 2.43 | 0.67 | 1.24 | 0 | 7.15 | 2.16 | 2.04
| Analytical results for Zn in OREAS 236 (Certified Value 144 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 156 | 142 | 143.4145 | 138 | 167 | 145 | 140 | 147 | 125 | 137 | 136 | 139.24 | 130 | 157.4 | 147 | 145 | 133.2 | 140 | 146 | 150 | 143 | 145 |  | 1
| 1 | 2 | 151 | 142 | 144.9256 | 143 | 169 | 144 | 140 | 148 | 121 | 140 | 135 | 139.79 | 137 | 166 | 150 | 143 | 129.9 | 139 | 147 | 151 | 148 | 150 |  | 11
| 1 | 3 | 148 | 146 | 143.7117 | 146 | 170 | 144 | 143 | 148 | 126 | 139 | 134 | 139.48 | 131 | 151.6 | 146 | 147 | 135.4 | 137 | 148 | 154 | 142 | 153 |  | 16
| 1 | 4 | 149 | 140 | 143.6512 | 146 | 175 | 143 | 146 | 149 | 123 | 140 | 136 | 140.15 | 132 | 160.5 | 147 | 146 | 127.6 | 138 | 147 | 158 | 145 | 149 |  | 143.6533
| 1 | 5 | 148 | 144 | 143.7365 | 143 | 168 | 147 | 142 | 149 | 119 | 140 | 138 | 138.27 | 132 | 164.2 | 147 | 141 | 120.2 | 140 | 146 | 155 | 143 | 149 |  | 72
| 1 | 6 | 148 | 140 | 144.6154 | 144 | 171 | 145 | 145 | 147 | 124 | 138 | 144 | 138.69 | 134 | 154.1 | 152 | 140 | 129.2 | 141 | 146 | 152 | 144 | 151
| Mean |  | 150 | 142.3333 | 144.0091 | 143.3333 | 170 | 144.6667 | 142.6667 | 148 | 123 | 139 | 137.1667 | 139.27 | 132.6667 | 158.9667 | 148.1667 | 143.6667 | 129.25 | 139.1667 | 146.6667 | 153.3333 | 144.1667 | 149.5
| Median |  | 148.5 | 142 | 143.7241 | 143.5 | 169.5 | 144.5 | 142.5 | 148 | 123.5 | 139.5 | 136 | 139.36 | 132 | 158.95 | 147 | 144 | 129.55 | 139.5 | 146.5 | 153 | 143.5 | 149.5
| Std Dev. |  | 3.1623 | 2.3381 | 0.60858 | 2.9439 | 2.8284 | 1.3663 | 2.5033 | 0.89443 | 2.6077 | 1.2649 | 3.6009 | 0.69659 | 2.5033 | 5.65 | 2.3166 | 2.8048 | 5.2626 | 1.472 | 0.8165 | 2.9439 | 2.137 | 2.6646
| Rel.Std.Dev. |  | 2.108% | 1.643% | 0.4226% | 2.054% | 1.664% | 0.9444% | 1.755% | 0.6043% | 2.12% | 0.91% | 2.625% | 0.5002% | 1.887% | 3.554% | 1.564% | 1.952% | 4.072% | 1.058% | 0.5567% | 1.92% | 1.482% | 1.782%
| PDM3 |  | 4.418% | -0.9188% | 0.2477% | -0.2227% | 18.34% | 0.7054% | -0.6868% | 3.026% | -14.38% | -3.239% | -4.515% | -3.051% | -7.648% | 10.66% | 3.142% | 0.009312% | -10.03% | -3.123% | 2.098% | 6.738% | 0.3574% | 4.07%
| Z-Score (Absolute) |  | 0.9 | 0.22 | 0.02 | 0.07 | 3.82 | 0.12 | 0.17 | 0.61 | 3.04 | 0.71 | 0.97 | 0.67 | 1.63 | 2.21 | 0.63 | 0.02 | 2.13 | 0.68 | 0.41 | 1.39 | 0.05 | 0.83
| Analytical results for Zr in OREAS 236 (Certified Value 55 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 59.9 | 53 | 52.4835 | 74.6 | 73.9 | 54.1 | 62 | 48.5 | 51 | 52.7 | 27 | 71.41 | 54.1 | 50.6 | 52.5 | 50.9 | 55 | 69 | 56.1 | 58.7 | 53.4 |  |  | 1
| 1 | 2 | 58.1 | 56 | 52.347 | 78.4 | 77.8 | 55 | 61.7 | 51.2 | 55 | 61.8 | 27 | 71.59 | 52.6 | 51.6 | 53.4 | 50.3 | 56 | 69 | 58.2 | 57.5 | 55.3 |  |  | 12
| 1 | 3 | 64.2 | 53 | 52.0756 | 78.4 | 72.9 | 53.9 | 61.2 | 51 | 52 | 55.7 | 46 | 71.73 | 54.3 | 52.7 | 53.7 | 52.3 | 56 | 68 | 56.4 | 57.7 | 55.4 |  |  | 16
| 1 | 4 | 59.5 | 54 | 50.7505 | 76.9 | 75.6 | 54.4 | 61 | 48.7 | 55 | 54.1 | 46 | 71.36 | 55.1 | 52.2 | 54 | 49.8 | 54 | 68 | 56.9 | 59.4 | 56.1 |  |  | 54.7689
| 1 | 5 | 58.4 | 56 | 51.7666 | 73.5 | 74.7 | 54.6 | 65.6 | 50.5 | 56 | 66.2 | 52 | 64.65 | 54.2 | 52 | 54.4 | 44.3 | 50 | 72 | 58.4 | 60.1 | 56.8 |  |  | 73
| 1 | 6 | 58.8 | 57 | 52.506 | 76.1 | 75.6 | 53.9 | 63.1 | 49.9 | 54 | 52.7 | 68 | 70.6 | 53.4 | 52.1 | 54.2 | 49 | 51 | 67 | 54.9 | 61 | 57
| Mean |  | 59.8167 | 54.8333 | 51.9882 | 76.3167 | 75.0833 | 54.3167 | 62.4333 | 49.9667 | 53.8333 | 57.2 | 44.3333 | 70.2233 | 53.95 | 51.8667 | 53.7 | 49.4333 | 53.6667 | 68.8333 | 56.8167 | 59.0667 | 55.6667
| Median |  | 59.15 | 55 | 52.2113 | 76.5 | 75.15 | 54.25 | 61.85 | 50.2 | 54.5 | 54.9 | 46 | 71.385 | 54.15 | 52.05 | 53.85 | 50.05 | 54.5 | 68.5 | 56.65 | 59.05 | 55.75
| Std Dev. |  | 2.2498 | 1.7224 | 0.66796 | 1.9974 | 1.6869 | 0.43551 | 1.7189 | 1.1518 | 1.9408 | 5.5591 | 15.6546 | 2.7583 | 0.85499 | 0.71461 | 0.68702 | 2.7493 | 2.582 | 1.7224 | 1.3258 | 1.3692 | 1.311
| Rel.Std.Dev. |  | 3.761% | 3.141% | 1.285% | 2.617% | 2.247% | 0.8018% | 2.753% | 2.305% | 3.605% | 9.719% | 35.31% | 3.928% | 1.585% | 1.378% | 1.279% | 5.562% | 4.811% | 2.502% | 2.333% | 2.318% | 2.355%
| PDM3 |  | 9.216% | 0.1176% | -5.077% | 39.34% | 37.09% | -0.8257% | 13.99% | -8.768% | -1.708% | 4.439% | -19.05% | 28.22% | -1.495% | -5.299% | -1.952% | -9.742% | -2.013% | 25.68% | 3.739% | 7.847% | 1.639%
| Z-Score (Absolute) |  | 1.13 | 0 | 0.65 | 4.88 | 4.6 | 0.12 | 1.73 | 1.11 | 0.23 | 0.54 | 2.39 | 3.5 | 0.2 | 0.67 | 0.26 | 1.23 | 0.27 | 3.18 | 0.45 | 0.96 | 0.19
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C240: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C258: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C276: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C294: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C350: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C460: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C478: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C497: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C516: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C535: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C553: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C571: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C590: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C608: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C626: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C644: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C680: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C698: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C716: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C734: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C752: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C770: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 236 (Certified Value 0.488 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.45 | 0.45 | 0.6 | 0.44 | 0.48 | 0.5 | 0.5 | 0.55 | 0.51 | < 5 | 0.56 | 0.479 |  | 0.5 | 0.51 | 0.4 | 0.508 | < 1 | 0.9 | < 0.5 | 0.47 | 0.46 | 0.47 |  | 1
| 1 | 2 | 0.47 | 0.45 | 0.58398 | 0.46 | 0.54 | 0.5 | 0.49 | 0.5 | 0.52 | < 5 | 0.5 | 0.467 |  | 0.54 | 0.51 | 0.4 | 0.533 | < 1 | 1 | < 0.5 | 0.46 | 0.48 | 0.47 |  | 18
| 1 | 3 | 0.48 | 0.45 | 0.59917 | 0.47 | 0.46 | 0.5 | 0.5 | 0.54 | 0.59 | < 5 | 0.51 | 0.481 |  | 0.52 | 0.5 | 0.4 | 0.528 | < 1 | 0.9 | < 0.5 | 0.47 | 0.47 | 0.48 |  | 16
| 1 | 4 | 0.47 | 0.45 | 0.59269 | 0.49 | 0.45 | 0.5 | 0.47 | 0.51 | 0.58 | < 5 | 0.48 | 0.451 |  | 0.48 | 0.52 | 0.4 | 0.516 | < 1 | 0.8 | < 0.5 | 0.46 | 0.47 | 0.48 |  | 0.48763
| 1 | 5 | 0.48 | 0.45 | 0.5835 | 0.48 | 0.47 | 0.5 | 0.48 | 0.51 | 0.52 | < 5 | 0.45 | 0.476 |  | 0.54 | 0.51 | 0.39 | 0.502 | < 1 | 0.8 | < 0.5 | 0.48 | 0.47 | 0.48 |  | 75
| 1 | 6 | 0.46 | 0.5 | 0.5935 | 0.46 | 0.47 | 0.5 | 0.49 | 0.51 | 0.52 | < 5 | 0.46 | 0.47 |  | 0.54 | 0.5 | 0.39 | 0.52 | < 1 | 0.8 | < 0.5 | 0.47 | 0.48 | 0.47
| Mean |  | 0.46833 | 0.45833 | 0.59214 | 0.46667 | 0.47833 | 0.5 | 0.48833 | 0.52 | 0.54 |  | 0.49333 | 0.47067 |  | 0.52 | 0.50833 | 0.39667 | 0.51783 |  | 0.86667 |  | 0.46833 | 0.47167 | 0.475
| Median |  | 0.47 | 0.45 | 0.5931 | 0.465 | 0.47 | 0.5 | 0.49 | 0.51 | 0.52 |  | 0.49 | 0.473 |  | 0.53 | 0.51 | 0.4 | 0.518 |  | 0.85 |  | 0.47 | 0.47 | 0.475
| Std Dev. |  | 0.01169 | 0.020412 | 0.0071358 | 0.017512 | 0.031885 | 0 | 0.01169 | 0.02 | 0.035214 |  | 0.039833 | 0.011003 |  | 0.025298 | 0.0075277 | 0.005164 | 0.011737 |  | 0.08165 |  | 0.0075277 | 0.0075277 | 0.0054772
| Rel.Std.Dev. |  | 2.496% | 4.454% | 1.205% | 3.753% | 6.666% | 0% | 2.394% | 3.846% | 6.521% |  | 8.074% | 2.338% |  | 4.865% | 1.481% | 1.302% | 2.267% |  | 9.421% |  | 1.607% | 1.596% | 1.153%
| PDM3 |  | -3.958% | -6.009% | 21.43% | -4.3% | -1.907% | 2.536% | 0.1436% | 6.638% | 10.74% |  | 1.169% | -3.479% |  | 6.638% | 4.245% | -18.65% | 6.193% |  | 77.73% |  | -3.958% | -3.274% | -2.591%
| Z-Score (Absolute) |  | 0.49 | 0.81 | 3.52 | 0.54 | 0.16 | NA | 0.16 | 1.19 | 1.83 | 65.27 | 0.32 | 0.41 | NA | 1.19 | 0.81 | 2.81 | 1.12 | 0.54 | NA | 7.55 | 0.49 | 0.38 | 0.27
Results from laboratories 7 and 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 236 (Certified Value 3.16 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 3.13 | 3.03 | 2.9008 | 2.99 | 3.02 | 3.22 | 3.0267 | 3 | 3.319 | 3.3814 | 2.34 | 2.91 | 3.0124 | 3.34 | 3.266 | 3.41 | 3.26 | 3.23 | 3.0742 | 3.26 | 3.24 | 3.35 | 2.79 | 3.08 | 1
| 1 | 2 | 3.18 | 3.15 | 2.9142 | 2.98 | 2.93 | 3.24 | 3.1622 | 2.99 | 3.345 | 3.3298 | 2.29 | 2.9 | 3.0428 | 3.29 | 3.4211 | 3.34 | 3.23 | 3.3 | 3.2313 | 3.25 | 3.2 | 3.36 | 2.87 | 3.14 | #N/A
| 1 | 3 | 3.21 | 3.18 | 2.8775 | 3.01 | 3 | 3.23 | 3.1263 | 2.95 | 3.344 | 3.3582 | 2.29 | 2.86 | 3.0546 | 3.33 | 3.4432 | 3.23 | 3.23 | 3.25 | 3.1158 | 3.19 | 3.18 | 3.3 | 2.78 | 3.04 | 16
| 1 | 4 | 3.22 | 3.25 | 2.9133 | 2.96 | 3.06 | 3.19 | 3.1292 | 2.95 | 3.343 | 3.3821 | 2.35 | 3.02 | 3.0069 | 3.25 | 3.3237 | 3.32 | 3.28 | 3.34 | 3.1807 | 3.28 | 3.18 | 3.32 | 2.79 | 3.06 | 3.155
| 1 | 5 | 3.15 | 3.2 | 2.9525 | 3.01 | 3.1 | 3.21 | 3.1793 | 3.07 | 3.317 | 3.3296 | 2.32 | 2.94 | 3.0658 | 3.32 | 3.3815 | 3.29 | 3.32 | 3.33 | 3.2178 | 3.34 | 3.12 | 3.34 | 2.77 | 3.06 | 76
| 1 | 6 | 3.22 | 3.03 | 2.9425 | 2.99 | 2.81 | 3.29 | 3.047 | 2.85 | 3.337 | 3.3589 | 2.36 | 2.85 | 3.1202 | 3.29 | 3.2957 | 3.34 | 3.37 | 3.27 | 3.1611 | 3.3 | 3.2 | 3.32 | 2.91 | 3.11
| Mean |  | 3.185 | 3.14 | 2.9168 | 2.99 | 2.9867 | 3.23 | 3.1118 | 2.9683 | 3.3342 | 3.3567 | 2.325 | 2.9133 | 3.0505 | 3.3033 | 3.3552 | 3.3217 | 3.2817 | 3.2867 | 3.1635 | 3.27 | 3.1867 | 3.3317 | 2.8183 | 3.0817
| Median |  | 3.195 | 3.165 | 2.9138 | 2.99 | 3.01 | 3.225 | 3.1277 | 2.97 | 3.34 | 3.3585 | 2.33 | 2.905 | 3.0487 | 3.305 | 3.3526 | 3.33 | 3.27 | 3.285 | 3.1709 | 3.27 | 3.19 | 3.33 | 2.79 | 3.07
| Std Dev. |  | 0.038341 | 0.091214 | 0.027398 | 0.018974 | 0.10386 | 0.034059 | 0.061726 | 0.072778 | 0.012844 | 0.023324 | 0.030166 | 0.06186 | 0.041282 | 0.033267 | 0.071087 | 0.059805 | 0.054924 | 0.044121 | 0.060145 | 0.050596 | 0.039328 | 0.022286 | 0.057417 | 0.037103
| Rel.Std.Dev. |  | 1.204% | 2.905% | 0.9393% | 0.6346% | 3.477% | 1.054% | 1.984% | 2.452% | 0.3852% | 0.6949% | 1.297% | 2.123% | 1.353% | 1.007% | 2.119% | 1.8% | 1.674% | 1.342% | 1.901% | 1.547% | 1.234% | 0.6689% | 2.037% | 1.204%
| PDM3 |  | 0.9506% | -0.4757% | -7.55% | -5.23% | -5.336% | 2.377% | -1.37% | -5.917% | 5.678% | 6.391% | -26.31% | -7.66% | -3.314% | 4.701% | 6.345% | 5.282% | 4.014% | 4.173% | 0.2687% | 3.645% | 1.003% | 5.599% | -10.67% | -2.325%
| Z-Score (Absolute) |  | 0.05 | 0.17 | 1.26 | 0.9 | 0.91 | 0.27 | 0.3 | 1 | 0.78 | 0.89 | 4.14 | 1.27 | 0.6 | 0.63 | 0.88 | 0.72 | 0.52 | 0.55 | 0.05 | 0.47 | 0.06 | 0.77 | 1.74 | 0.45
| Analytical results for As in OREAS 236 (Certified Value 67 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 0
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 1
| 1 | 1 | 66.4 | 63.6 | 55.4455 | 65 | 65.4 | 64.8 | 61 | 70.8 | 68 | 59.142 | 70 | 67.4 | 69.1588 | 60 | 18.7458 | 71.5 | 68 | 71.8 | 56.7289 | 66 | 66 | 68.3 | 62.4 | 64.9 | 1
| 1 | 2 | 73.4 | 64.8 | 55.6254 | 66 | 65.9 | 64.8 | 63 | 68 | 69 | 61.084 | 69 | 67.4 | 67.524 | 63 | 19.043 | 71.4 | 68 | 75.2 | 52.3718 | 65 | 63 | 66.3 | 63.6 | 64.8 | 6
| 1 | 3 | 70.6 | 64.6 | 56.2445 | 65 | 67.1 | 64.5 | 65 | 70.6 | 69 | 59.779 | 70 | 67.6 | 65.7288 | 62 | 15.8467 | 67.7 | 68 | 74.3 | 52.5927 | 65 | 64 | 70 | 63.2 | 65.2 | 16
| 1 | 4 | 75.1 | 65.4 | 56.644 | 66 | 66.5 | 63.7 | 63 | 67.6 | 67 | 60.457 | 70 | 67.9 | 65.8818 | 62 | 16.6452 | 71.1 | 69 | 73.8 | 53.6181 | 66 | 64 | 66.1 | 63.8 | 64.8 | 66.6126
| 1 | 5 | 75.3 | 65.4 | 55.7175 | 67 | 66.7 | 64 | 64 | 69.4 | 66 | 61.043 | 67 | 67.4 | 68.0238 | 64 | 18.275 | 68.4 | 67 | 72.1 | 53.3966 | 68 | 66 | 66 | 62.7 | 65.3 | 77
| 1 | 6 | 73.7 | 67.8 | 55.4455 | 65 | 65.4 | 64.2 | 63 | 69.6 | 66 | 59.874 | 70 | 66.1 | 68.4326 | 63 | 16.37 | 70.7 | 69 | 71.6 | 53.6373 | 66 | 64 | 65 | 64.3 | 65.5
| Mean |  | 72.4167 | 65.2667 | 55.8537 | 65.6667 | 66.1667 | 64.3333 | 63.1667 | 69.3333 | 67.5 | 60.2298 | 69.3333 | 67.3 | 67.4583 | 62.3333 | 17.4876 | 70.1333 | 68.1667 | 73.1333 | 53.7242 | 66 | 64.5 | 66.95 | 63.3333 | 65.0833
| Median |  | 73.55 | 65.1 | 55.6715 | 65.5 | 66.2 | 64.35 | 63 | 69.5 | 67.5 | 60.1655 | 70 | 67.4 | 67.7739 | 62.5 | 17.4601 | 70.9 | 68 | 72.95 | 53.5073 | 66 | 64 | 66.2 | 63.4 | 65.05
| Std Dev. |  | 3.3949 | 1.4067 | 0.48605 | 0.8165 | 0.70899 | 0.44572 | 1.3292 | 1.3125 | 1.3784 | 0.76886 | 1.2111 | 0.61968 | 1.3886 | 1.3663 | 1.3619 | 1.6525 | 0.75277 | 1.5016 | 1.5661 | 1.0954 | 1.2247 | 1.8426 | 0.70899 | 0.29269
| Rel.Std.Dev. |  | 4.688% | 2.155% | 0.8702% | 1.243% | 1.072% | 0.6928% | 2.104% | 1.893% | 2.042% | 1.277% | 1.747% | 0.9208% | 2.059% | 2.192% | 7.788% | 2.356% | 1.104% | 2.053% | 2.915% | 1.66% | 1.899% | 2.752% | 1.119% | 0.4497%
| PDM3 |  | 8.713% | -2.021% | -16.15% | -1.42% | -0.6695% | -3.422% | -5.173% | 4.084% | 1.332% | -9.582% | 4.084% | 1.032% | 1.27% | -6.424% | -73.75% | 5.285% | 2.333% | 9.789% | -19.35% | -0.9197% | -3.171% | 0.5065% | -4.923% | -2.296%
| Z-Score (Absolute) |  | 1.84 | 0.16 | 2.78 | 0.05 | 0.09 | 0.42 | 0.74 | 0.98 | 0.47 | 1.56 | 0.98 | 0.41 | 0.45 | 0.98 | 13.49 | 1.2 | 0.65 | 2.04 | 3.38 | 0.05 | 0.37 | 0.31 | 0.7 | 0.21
| Analytical results for B in OREAS 236 (Certified Value 90 ppm) | [25 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.08 | 1.11 | 1.13 | 1.14 | 1.18 | 1.20 | 1.25 | 1.26 | 1.27 | [12 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 0
|  |  | 0.5g | 0.5g | 0.25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 15g | 0.5g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 90 | 89.0367 | 96 | 81 | 100 | 82 | 97 | 90 | 83 | 80 | 91 | 80 | [12 empty] | 1
| 1 | 2 | 90 | 89.3233 | 93 | 82 | 99 | 80 | 99 | 90 | 102 | 80 | 95 | 80 | [12 empty] | #N/A
| 1 | 3 | 90 | 88.6767 | 96 | 83 | 100 | 80 | 97 | 80 | 85 | 80 | 92 | 80 | [12 empty] | 16
| 1 | 4 | 90 | 88.3233 | 93 | 83 | 98 | 82 | 100 | 90 | 85 | 80 | 95 | 80 | [12 empty] | 90.2605
| 1 | 5 | 90 | 89.1067 | 95 | 83 | 98 | 78 | 98 | 90 | 87 | 80 | 90 | 80 | [12 empty] | 78
| 1 | 6 | 90 | 89.2367 | 97 | 82 | 99 | 82 | 97 | 90 | 84 | 80 | 97 | 80
| Mean |  | 90 | 88.9506 | 95 | 82.3333 | 99 | 80.6667 | 98 | 88.3333 | 87.6667 | 80 | 93.3333 | 80
| Median |  | 90 | 89.0717 | 95.5 | 82.5 | 99 | 81 | 97.5 | 90 | 85 | 80 | 93.5 | 80
| Std Dev. |  | 0 | 0.37975 | 1.6733 | 0.8165 | 0.89443 | 1.633 | 1.2649 | 4.0825 | 7.1461 | 0 | 2.7325 | 0
| Rel.Std.Dev. |  | 0% | 0.4269% | 1.761% | 0.9917% | 0.9035% | 2.024% | 1.291% | 4.622% | 8.151% | 0% | 2.928% | 0%
| PDM3 |  | -0.2886% | -1.451% | 5.251% | -8.783% | 9.683% | -10.63% | 8.575% | -2.135% | -2.874% | -11.37% | 3.404% | -11.37%
| Z-Score (Absolute) |  | NA | 0.28 | 0.49 | 1.11 | 0.99 | 1.32 | 0.86 | NA | 0.44 | NA | 0.28 | NA
Results from laboratories 1, 18, 25 and 27 were removed due to their 10 ppm reading resolution.
| Analytical results for Ba in OREAS 236 (Certified Value 35.7 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 40 | 33 | 65.9088 | 39 | 36.9 | 33 | 35 | 37 | 37 | 32.927 | 32 | 39.6 | 47.4752 | 32 | 35.4265 | 40 | 33 | 37.8 | 37.5621 | 35 | 37 | 40 | 36 | 30 | 1
| 1 | 2 | 40 | 33 | 65.5326 | 38 | 36.3 | 33 | 36 | 35 | 38 | 33.525 | 31 | 43.9 | 47.7159 | 30 | 37.5412 | 40 | 33 | 40 | 36.5005 | 36 | 36 | 40 | 37 | 40 | 20
| 1 | 3 | 40 | 33 | 65.4189 | 37 | 36.6 | 34 | 37 | 38 | 37 | 33.648 | 31 | 41.3 | 47.7197 | 30 | 40.0407 | 40 | 32 | 41.5 | 34.7902 | 35 | 36 | 40 | 36 | 30 | 16
| 1 | 4 | 40 | 34 | 65.3435 | 37 | 36.5 | 34 | 35 | 37 | 37 | 32.916 | 31 | 43 | 46.5124 | 30 | 35.7021 | 40 | 34 | 39.6 | 35.5209 | 36 | 36 | 40 | 36 | 30 | 35.6851
| 1 | 5 | 40 | 34 | 65.8058 | 39 | 36.6 | 34 | 36 | 36 | 37 | 33.566 | 30 | 41.7 | 48.8394 | 30 | 39.4902 | 40 | 33 | 39.5 | 36.3443 | 37 | 35 | 40 | 36 | 30 | 79
| 1 | 6 | 40 | 34 | 65.4964 | 39 | 35.9 | 34 | 35 | 34 | 38 | 33.647 | 32 | 43.4 | 48.1212 | 30 | 37.9007 | 40 | 33 | 36.9 | 35.9421 | 36 | 36 | 40 | 38 | 40
| Mean |  | 40 | 33.5 | 65.5843 | 38.1667 | 36.4667 | 33.6667 | 35.6667 | 36.1667 | 37.3333 | 33.3715 | 31.1667 | 42.15 | 47.7306 | 30.3333 | 37.6836 | 40 | 33 | 39.2167 | 36.11 | 35.8333 | 36 | 40 | 36.5 | 33.3333
| Median |  | 40 | 33.5 | 65.5145 | 38.5 | 36.55 | 34 | 35.5 | 36.5 | 37 | 33.5455 | 31 | 42.35 | 47.7178 | 30 | 37.721 | 40 | 33 | 39.55 | 36.1432 | 36 | 36 | 40 | 36 | 30
| Std Dev. |  | 0 | 0.54772 | 0.22365 | 0.98319 | 0.33862 | 0.5164 | 0.8165 | 1.472 | 0.5164 | 0.3518 | 0.75277 | 1.5984 | 0.76568 | 0.8165 | 1.8925 | 0 | 0.63246 | 1.6388 | 0.94159 | 0.75277 | 0.63246 | 0 | 0.83666 | 5.164
| Rel.Std.Dev. |  | 0% | 1.635% | 0.341% | 2.576% | 0.9286% | 1.534% | 2.289% | 4.07% | 1.383% | 1.054% | 2.415% | 3.792% | 1.604% | 2.692% | 5.022% | 0% | 1.917% | 4.179% | 2.608% | 2.101% | 1.757% | 0% | 2.292% | 15.49%
| PDM3 |  | 12.09% | -6.123% | 83.79% | 6.954% | 2.19% | -5.656% | -0.05164% | 1.35% | 4.619% | -6.483% | -12.66% | 18.12% | 33.76% | -15% | 5.6% | 12.09% | -7.524% | 9.896% | 1.191% | 0.4154% | 0.8825% | 12.09% | 2.284% | -6.59%
| Z-Score (Absolute) |  | NA | 0.79 | 8.82 | 0.61 | 0.1 | 0.74 | 0.14 | 0.01 | 0.36 | 0.83 | 1.49 | 1.8 | 3.47 | 1.74 | 0.46 | NA | 0.94 | 0.92 | 0.01 | 0.09 | 0.04 | NA | 0.11 | NA
Results from laboratories 1, 18, 25 and 27 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 236 (Certified Value 0.26 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | [5 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.15g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.23 | < 5 | 0.3 | 0.3 | < 1 | 0.27 | 0.29 | < 0.5 | < 5 | 0.26 | 0.2 | 0.2283 | 0.33 | < 0.5 | < 1 | < 0.5 | 0.3 | 0.28 | 0.24 | [5 empty] | 1
| 1 | 2 | 0.27 | < 5 | 0.3 | 0.3 | < 1 | 0.24 | 0.26 | < 0.5 | < 5 | 0.26 | 0.2 | 0.25036 | 0.33 | < 0.5 | < 1 | < 0.5 | 0.29 | 0.24 | 0.24 | [5 empty] | 21
| 1 | 3 | 0.24 | < 5 | 0.3 | 0.3 | < 1 | 0.27 | 0.3 | < 0.5 | < 5 | 0.24 | 0.2 | 0.25361 | 0.28 | < 0.5 | < 1 | < 0.5 | 0.29 | 0.25 | 0.26 | [5 empty] | 16
| 1 | 4 | 0.29 | < 5 | 0.3 | 0.3 | < 1 | 0.24 | 0.3 | < 0.5 | < 5 | 0.26 | 0.2 | 0.21672 | 0.31 | < 0.5 | < 1 | < 0.5 | 0.29 | 0.26 | 0.25 | [5 empty] | 0.26335
| 1 | 5 | 0.26 | < 5 | 0.3 | 0.3 | < 1 | 0.27 | 0.31 | < 0.5 | < 5 | 0.26 | 0.3 | 0.24882 | 0.31 | < 0.5 | < 1 | < 0.5 | 0.27 | 0.26 | 0.26 | [5 empty] | 80
| 1 | 6 | 0.27 | < 5 | 0.3 | 0.3 | < 1 | 0.26 | 0.28 | < 0.5 | < 5 | 0.26 | 0.2 | 0.25284 | 0.3 | < 0.5 | < 1 | < 0.5 | 0.29 | 0.26 | 0.24
| Mean |  | 0.26 |  | 0.3 | 0.3 |  | 0.25833 | 0.29 |  |  | 0.25667 | 0.21667 | 0.24177 | 0.31 |  |  |  | 0.28833 | 0.25833 | 0.24833
| Median |  | 0.265 |  | 0.3 | 0.3 |  | 0.265 | 0.295 |  |  | 0.26 | 0.2 | 0.24959 | 0.31 |  |  |  | 0.29 | 0.26 | 0.245
| Std Dev. |  | 0.021909 |  | 0 | 0 |  | 0.01472 | 0.017889 |  |  | 0.008165 | 0.040825 | 0.01546 | 0.018974 |  |  |  | 0.0098319 | 0.013292 | 0.0098319
| Rel.Std.Dev. |  | 8.427% |  | 0% | 0% |  | 5.698% | 6.168% |  |  | 3.181% | 18.84% | 6.394% | 6.121% |  |  |  | 3.41% | 5.145% | 3.959%
| PDM3 |  | -1.271% |  | 13.92% | 13.92% |  | -1.904% | 10.12% |  |  | -2.537% | -17.73% | -8.192% | 17.71% |  |  |  | 9.488% | -1.904% | -5.701%
| Z-Score (Absolute) |  | 0.04 | 107.06 | NA | NA | 11.51 | 0.04 | 1.47 | 0.44 | 107.06 | 0.12 | NA | 0.83 | 2.43 | 0.44 | 11.51 | 0.44 | 1.39 | 0.04 | 0.52
Results from laboratories 4, 5 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 236 (Certified Value 0.062 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.06 | 0.06 | < 5 | 0.07 | < 0.1 | 0.06 | 0.07 | < 0.05 | < 5 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 8 | < 5 | 0.05 | 0.06 | 0.06 |  |  |  |  | 1
| 1 | 2 | 0.07 | 0.06 | < 5 | 0.07 | < 0.1 | 0.07 | 0.07 | < 0.05 | < 5 | 0.07 | 0.07 | 0.05 | 0.06 | 0.06 | 0.06 | 8 | < 5 | 0.05 | 0.06 | 0.06 |  |  |  |  | 22
| 1 | 3 | 0.06 | 0.06 | < 5 | 0.06 | < 0.1 | 0.06 | 0.07 | < 0.05 | < 5 | 0.07 | 0.06 | 0.05 | 0.06 | 0.06 | 0.06 | 8 | < 5 | 0.06 | 0.06 | 0.06 |  |  |  |  | 16
| 1 | 4 | 0.06 | 0.06 | < 5 | 0.07 | < 0.1 | 0.07 | 0.06 | < 0.05 | < 5 | 0.07 | 0.07 | 0.05 | 0.06 | 0.06 | 0.09 | 9 | < 5 | 0.06 | 0.06 | 0.06 |  |  |  |  | 0.061833
| 1 | 5 | 0.06 | 0.06 | < 5 | 0.07 | < 0.1 | 0.06 | 0.07 | < 0.05 | < 5 | 0.07 | 0.06 | 0.05 | 0.06 | 0.06 | 0.07 | 7 | < 5 | 0.05 | 0.06 | 0.06 |  |  |  |  | 81
| 1 | 6 | 0.06 | 0.06 | < 5 | 0.07 | < 0.1 | 0.06 | 0.07 | < 0.05 | < 5 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 9 | < 5 | 0.05 | 0.06 | 0.06
| Mean |  | 0.061667 | 0.06 |  | 0.068333 |  | 0.063333 | 0.068333 |  |  | 0.068333 | 0.065 | 0.053333 | 0.06 | 0.06 | 0.068333 | 8.1667 |  | 0.053333 | 0.06 | 0.06
| Median |  | 0.06 | 0.06 |  | 0.07 |  | 0.06 | 0.07 |  |  | 0.07 | 0.065 | 0.05 | 0.06 | 0.06 | 0.065 | 8 |  | 0.05 | 0.06 | 0.06
| Std Dev. |  | 0.0040825 | 0 |  | 0.0040825 |  | 0.005164 | 0.0040825 |  |  | 0.0040825 | 0.0054772 | 0.005164 | 0 | 0 | 0.01169 | 0.75277 |  | 0.005164 | 0 | 0
| Rel.Std.Dev. |  | 6.62% | 0% |  | 5.974% |  | 8.154% | 5.974% |  |  | 5.974% | 8.427% | 9.682% | 0% | 0% | 17.11% | 9.218% |  | 9.682% | 0% | 0%
| PDM3 |  | -0.2695% | -2.965% |  | 10.51% |  | 2.426% | 10.51% |  |  | 10.51% | 5.121% | -13.75% | -2.965% | -2.965% | 10.51% | 13108% |  | -13.75% | -2.965% | -2.965%
| Z-Score (Absolute) |  | 0 | 0.17 | 246.63 | 0.67 | 1.18 | 0.17 | 0.67 | 3.71 | 246.63 | 0.67 | 0.34 | 0.84 | 0.17 | 0.17 | 0.67 | NA | 246.63 | 0.84 | 0.17 | 0.17
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 236 (Certified Value 2.64 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.63 | 2.08 | 2.4542 | 2.21 | 2.75 | 2.42 | 2.66 | 2.68 | 2.7778 | 3.17 | 3.03 | 2.3383 | 2.87 | 2.8602 | 2.81 | 2.78 | 2.56 | 2.8244 | 2.53 | 2.38 | 2.77 | 2.55 | 2.41 |  | 1
| 1 | 2 | 2.62 | 2.18 | 2.4628 | 2.21 | 2.79 | 2.51 | 2.74 | 2.67 | 2.7279 | 3.04 | 3.14 | 2.3971 | 2.72 | 3.0436 | 2.77 | 2.72 | 2.59 | 2.8392 | 2.5 | 2.38 | 2.78 | 2.6 | 2.51 |  | #N/A
| 1 | 3 | 2.63 | 2.24 | 2.446 | 2.22 | 2.79 | 2.52 | 2.71 | 2.66 | 2.7916 | 3.05 | 3.13 | 2.4036 | 2.66 | 3.063 | 2.65 | 2.8 | 2.57 | 2.8068 | 2.45 | 2.36 | 2.72 | 2.49 | 2.37 |  | 16
| 1 | 4 | 2.64 | 2.33 | 2.4657 | 2.17 | 2.76 | 2.52 | 2.7 | 2.69 | 2.7782 | 3.1 | 3.25 | 2.3682 | 2.67 | 2.9395 | 2.73 | 2.78 | 2.63 | 2.8281 | 2.53 | 2.36 | 2.71 | 2.49 | 2.44 |  | 2.6373
| 1 | 5 | 2.59 | 2.29 | 2.4973 | 2.22 | 2.73 | 2.54 | 2.73 | 2.6 | 2.7278 | 3.04 | 3.08 | 2.3933 | 2.68 | 3.0006 | 2.71 | 2.8 | 2.66 | 2.8025 | 2.61 | 2.35 | 2.75 | 2.49 | 2.38 |  | 82
| 1 | 6 | 2.7 | 1.96 | 2.5007 | 2.21 | 2.75 | 2.4 | 2.3 | 2.6 | 2.7914 | 3.12 | 3.08 | 2.4499 | 2.71 | 2.9197 | 2.74 | 2.75 | 2.59 | 2.9058 | 2.6 | 2.35 | 2.73 | 2.67 | 2.47
| Mean |  | 2.635 | 2.18 | 2.4711 | 2.2067 | 2.7617 | 2.485 | 2.64 | 2.65 | 2.7657 | 3.0867 | 3.1183 | 2.3917 | 2.7183 | 2.9711 | 2.735 | 2.7717 | 2.6 | 2.8345 | 2.5367 | 2.3633 | 2.7433 | 2.5483 | 2.43
| Median |  | 2.63 | 2.21 | 2.4642 | 2.21 | 2.755 | 2.515 | 2.705 | 2.665 | 2.778 | 3.075 | 3.105 | 2.3952 | 2.695 | 2.97 | 2.735 | 2.78 | 2.59 | 2.8263 | 2.53 | 2.36 | 2.74 | 2.52 | 2.425
| Std Dev. |  | 0.036194 | 0.139 | 0.022695 | 0.018619 | 0.024014 | 0.059245 | 0.16888 | 0.04 | 0.029993 | 0.052789 | 0.075741 | 0.037325 | 0.077825 | 0.078109 | 0.054314 | 0.031252 | 0.037947 | 0.037508 | 0.060553 | 0.013663 | 0.028048 | 0.074409 | 0.054037
| Rel.Std.Dev. |  | 1.374% | 6.376% | 0.9184% | 0.8438% | 0.8695% | 2.384% | 6.397% | 1.509% | 1.084% | 1.71% | 2.429% | 1.561% | 2.863% | 2.629% | 1.986% | 1.128% | 1.46% | 1.323% | 2.387% | 0.5781% | 1.022% | 2.92% | 2.224%
| PDM3 |  | -0.08749% | -17.34% | -6.302% | -16.33% | 4.715% | -5.775% | 0.1021% | 0.4813% | 4.87% | 17.04% | 18.24% | -9.312% | 3.072% | 12.66% | 3.704% | 5.095% | -1.415% | 7.476% | -3.816% | -10.39% | 4.02% | -3.374% | -7.861%
| Z-Score (Absolute) |  | 0.03 | 2.36 | 0.86 | 2.22 | 0.62 | 0.79 | 0 | 0.05 | 0.64 | 2.29 | 2.45 | 1.27 | 0.4 | 1.7 | 0.49 | 0.67 | 0.2 | 1 | 0.53 | 1.42 | 0.53 | 0.47 | 1.08
| Analytical results for Cd in OREAS 236 (Certified Value 0.61 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.58 | 0.6 | 0.744 | 0.65 | 0.63 | 0.8 | 0.58 | 0.68 | 0.7 | < 5 | 0.63 | 0.69 | 0.58 | < 3 | 0.59 | 0.45 | 0.68 | 7.4008 | < 1 | < 1 | 0.61 | 0.57 | 0.61 |  | 1
| 1 | 2 | 0.64 | 0.6 | 0.75 | 0.67 | 0.63 | 0.7 | 0.62 | 0.67 | 0.63 | < 5 | 0.59 | 0.58 | 0.61 | < 3 | 0.61 | 0.45 | 0.72 | 7.2989 | < 1 | < 1 | 0.6 | 0.57 | 0.63 |  | 23
| 1 | 3 | 0.6 | 0.6 | 0.74033 | 0.65 | 0.63 | 0.8 | 0.6 | 0.65 | 0.65 | < 5 | 0.62 | 0.58 | 0.58 | < 3 | 0.62 | 0.44 | 0.74 | 6.561 | < 1 | < 1 | 0.63 | 0.58 | 0.59 |  | 16
| 1 | 4 | 0.62 | 0.6 | 0.74547 | 0.65 | 0.64 | 0.7 | 0.6 | 0.64 | 0.61 | < 5 | 0.62 | 0.55 | 0.58 | < 3 | 0.6 | 0.45 | 0.69 | 6.6881 | < 1 | < 1 | 0.6 | 0.58 | 0.59 |  | 0.61448
| 1 | 5 | 0.6 | 0.6 | 0.732 | 0.67 | 0.64 | 0.7 | 0.61 | 0.66 | 0.66 | < 5 | 0.59 | 0.63 | 0.62 | < 3 | 0.59 | 0.45 | 0.71 | 6.3516 | < 1 | < 1 | 0.61 | 0.57 | 0.6 |  | 83
| 1 | 6 | 0.62 | 0.65 | 0.734 | 0.65 | 0.62 | 0.7 | 0.61 | 0.65 | 0.7 | < 5 | 0.62 | 0.59 | 0.62 | < 3 | 0.6 | 0.47 | 0.69 | 7.0026 | < 1 | < 1 | 0.63 | 0.57 | 0.58
| Mean |  | 0.61 | 0.60833 | 0.74097 | 0.65667 | 0.63167 | 0.73333 | 0.60333 | 0.65833 | 0.65833 |  | 0.61167 | 0.60333 | 0.59833 |  | 0.60167 | 0.45167 | 0.705 | 6.8838 |  |  | 0.61333 | 0.57333 | 0.6
| Median |  | 0.61 | 0.6 | 0.74217 | 0.65 | 0.63 | 0.7 | 0.605 | 0.655 | 0.655 |  | 0.62 | 0.585 | 0.595 |  | 0.6 | 0.45 | 0.7 | 6.8453 |  |  | 0.61 | 0.57 | 0.595
| Std Dev. |  | 0.020976 | 0.020412 | 0.0069345 | 0.010328 | 0.0075277 | 0.05164 | 0.013663 | 0.01472 | 0.03656 |  | 0.017224 | 0.049666 | 0.020412 |  | 0.01169 | 0.0098319 | 0.022583 | 0.4194 |  |  | 0.013663 | 0.005164 | 0.017889
| Rel.Std.Dev. |  | 3.439% | 3.355% | 0.9359% | 1.573% | 1.192% | 7.042% | 2.265% | 2.236% | 5.553% |  | 2.816% | 8.232% | 3.412% |  | 1.943% | 2.177% | 3.203% | 6.093% |  |  | 2.228% | 0.9007% | 2.981%
| PDM3 |  | -0.7285% | -0.9997% | 20.59% | 6.866% | 2.798% | 19.34% | -1.813% | 7.137% | 7.137% |  | -0.4572% | -1.813% | -2.627% |  | -2.085% | -26.5% | 14.73% | 1020% |  |  | -0.186% | -6.696% | -2.356%
| Z-Score (Absolute) |  | 0.01 | 0.04 | 2.11 | 0.74 | 0.34 | NA | 0.12 | 0.77 | 0.77 | 30.57 | 0.01 | 0.12 | 0.2 | 14.39 | 0.15 | 2.58 | 1.52 | 101.52 | 1.79 | 1.79 | 0.04 | 0.61 | 0.18
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 236 (Certified Value 10.6 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 9.92 | 10 | 11.34 | 10.7 | 10 | 10.411 | 10.52 | 11.38 | 22.062 | 7.63 | 10.8 | 10.3442 | 10.14 | 11.05 | 10.5 | 11 | < 5 | 11.15 | 10.12 | 9.97 |  |  |  |  | 1
| 1 | 2 | 10.85 | 10.3 | 10.96 | 10.6 | 10 | 10.491 | 10.35 | 11.37 | 21.656 | 7.54 | 11.1 | 10.6345 | 10.11 | 11.05 | 10.7 | 10 | < 5 | 10.65 | 10.11 | 10.5 |  |  |  |  | 24
| 1 | 3 | 10.35 | 10.4 | 10.95 | 11.1 | 10 | 10.85 | 10.65 | 11.36 | 21.25 | 7.46 | 10.9 | 10.605 | 10.12 | 10.65 | 10.9 | < 10 | < 5 | 11.05 | 10 | 10.2 |  |  |  |  | 16
| 1 | 4 | 10.55 | 10.6 | 11.11 | 10.9 | 10 | 10.573 | 10.56 | 11.14 | 22.146 | 7.64 | 11.2 | 10.3841 | 10.16 | 10.9 | 10.7 | 11 | < 5 | 10.85 | 9.95 | 10.35 |  |  |  |  | 10.6291
| 1 | 5 | 10.75 | 10.5 | 11.08 | 10.9 | 10 | 10.773 | 10.25 | 11 | 21.657 | 7.5 | 11 | 10.8992 | 10.22 | 10.65 | 10.7 | 11 | < 5 | 10.8 | 10.02 | 10 |  |  |  |  | 84
| 1 | 6 | 10.35 | 10.7 | 11.22 | 10.7 | 10 | 10.231 | 9.89 | 10.94 | 21.21 | 7.61 | 10.9 | 10.8857 | 10.2 | 11.35 | 10.5 | 12 | < 5 | 10.75 | 10.28 | 10.05
| Mean |  | 10.4617 | 10.4167 | 11.11 | 10.8167 | 10 | 10.5548 | 10.37 | 11.1983 | 21.6635 | 7.5633 | 10.9833 | 10.6254 | 10.1583 | 10.9417 | 10.6667 | 11 |  | 10.875 | 10.08 | 10.1783
| Median |  | 10.45 | 10.45 | 11.095 | 10.8 | 10 | 10.532 | 10.435 | 11.25 | 21.6565 | 7.575 | 10.95 | 10.6198 | 10.15 | 10.975 | 10.7 | 11 |  | 10.825 | 10.065 | 10.125
| Std Dev. |  | 0.33469 | 0.24833 | 0.151 | 0.18348 | 0 | 0.23008 | 0.27648 | 0.19904 | 0.39201 | 0.074476 | 0.1472 | 0.23685 | 0.044008 | 0.2691 | 0.15055 | 0.70711 |  | 0.18908 | 0.11781 | 0.21217
| Rel.Std.Dev. |  | 3.199% | 2.384% | 1.359% | 1.696% | 0% | 2.18% | 2.666% | 1.777% | 1.81% | 0.9847% | 1.34% | 2.229% | 0.4332% | 2.459% | 1.411% | 6.428% |  | 1.739% | 1.169% | 2.085%
| PDM3 |  | -1.576% | -1.999% | 4.524% | 1.764% | -5.919% | -0.699% | -2.438% | 5.355% | 103.8% | -28.84% | 3.332% | -0.03469% | -4.429% | 2.94% | 0.3532% | 3.489% |  | 2.313% | -5.166% | -4.241%
| Z-Score (Absolute) |  | 0.23 | 0.31 | 0.94 | 0.41 | NA | 0.06 | 0.4 | 1.1 | 20.05 | 5.48 | 0.71 | 0.06 | 0.78 | 0.64 | 0.14 | NA | 14.65 | 0.52 | 0.92 | 0.75
Results from laboratories 7 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 236 (Certified Value 30.8 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 27.3 | 30 | 28.0712 | 31.3 | 33.9 | 32.7 | 28.6 | 30.1 | 31.8 | 27.862 | 31.6 | 31.4 | 34.4269 | 28.7 | 40.9983 | 32.9 | 35.6 | 31.9 | 40.5053 | 27 | 28 | 31 | 30.6 | 29.2 | 1
| 1 | 2 | 29.7 | 31 | 28.4805 | 31.3 | 33.2 | 32.2 | 29.7 | 29.2 | 31.8 | 28.183 | 30.5 | 31.2 | 34.9654 | 30.6 | 41.1052 | 33.2 | 35.9 | 33.5 | 40.7775 | 27 | 28 | 31.6 | 30.7 | 30 | 10
| 1 | 3 | 28.5 | 31 | 27.937 | 31 | 32.5 | 32.6 | 30.1 | 30 | 32 | 27.541 | 31.3 | 30.1 | 34.9746 | 29.9 | 38.5643 | 31.5 | 36.9 | 34.1 | 38.7037 | 27 | 28 | 30.7 | 30.4 | 29.5 | 16
| 1 | 4 | 31.5 | 32 | 28.1113 | 30.3 | 33.5 | 32.2 | 29.3 | 29 | 31.2 | 27.367 | 31.5 | 30.9 | 33.0988 | 29.5 | 38.0232 | 32.4 | 35.4 | 33.1 | 39.0289 | 27 | 28 | 30.9 | 30.7 | 29.1 | 30.8067
| 1 | 5 | 31.4 | 31 | 28.4343 | 31.4 | 33 | 32 | 29.6 | 30.2 | 30.7 | 28.138 | 30.5 | 30.4 | 34.0132 | 30.8 | 37.6368 | 32.3 | 35.7 | 33.5 | 38.191 | 27 | 28 | 31.4 | 30.5 | 29.3 | 85
| 1 | 6 | 31.7 | 31 | 28.8037 | 31.5 | 32.8 | 32.1 | 28.8 | 28.8 | 30.9 | 27.443 | 31.1 | 30.2 | 33.832 | 30.3 | 40.2494 | 32.6 | 36.1 | 31.4 | 39.7443 | 27 | 28 | 31.6 | 31.2 | 29.5
| Mean |  | 30.0167 | 31 | 28.3063 | 31.1333 | 33.15 | 32.3 | 29.35 | 29.55 | 31.4 | 27.7557 | 31.0833 | 30.7 | 34.2185 | 29.9667 | 39.4295 | 32.4833 | 35.9333 | 32.9167 | 39.4918 | 27 | 28 | 31.2 | 30.6833 | 29.4333
| Median |  | 30.55 | 31 | 28.2728 | 31.3 | 33.1 | 32.2 | 29.45 | 29.6 | 31.5 | 27.7015 | 31.2 | 30.65 | 34.22 | 30.1 | 39.4069 | 32.5 | 35.8 | 33.3 | 39.3866 | 27 | 28 | 31.2 | 30.65 | 29.4
| Std Dev. |  | 1.8291 | 0.63246 | 0.32366 | 0.44121 | 0.501 | 0.28284 | 0.56833 | 0.61887 | 0.54037 | 0.35637 | 0.48339 | 0.54406 | 0.72363 | 0.77889 | 1.5415 | 0.58452 | 0.53166 | 1.0439 | 1.0268 | 0 | 0 | 0.38471 | 0.27869 | 0.32042
| Rel.Std.Dev. |  | 6.094% | 2.04% | 1.143% | 1.417% | 1.511% | 0.8757% | 1.936% | 2.094% | 1.721% | 1.284% | 1.555% | 1.772% | 2.115% | 2.599% | 3.91% | 1.799% | 1.48% | 3.171% | 2.6% | 0% | 0% | 1.233% | 0.9083% | 1.089%
| PDM3 |  | -2.564% | 0.6275% | -8.116% | 1.06% | 7.607% | 4.847% | -4.728% | -4.079% | 1.926% | -9.904% | 0.898% | -0.3463% | 11.07% | -2.727% | 27.99% | 5.442% | 16.64% | 6.849% | 28.19% | -12.36% | -9.111% | 1.277% | -0.4004% | -4.458%
| Z-Score (Absolute) |  | 0.45 | 0.02 | 1.19 | 0.04 | 0.92 | 0.55 | 0.74 | 0.65 | 0.16 | 1.43 | 0.02 | 0.15 | 1.38 | 0.47 | 3.65 | 0.63 | 2.13 | 0.82 | 3.68 | 1.76 | 1.32 | 0.07 | 0.16 | 0.7
| Analytical results for Cr in OREAS 236 (Certified Value 32.8 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 31 | 40 | 34.3193 | 34 | 31.5 | 35 | 30 | 35 | 35 | 31.34 | 30 | 32 | 32.63 | 33 | 33.39 | 33 | 39 | 36.7 | 34.8008 | 32 | 34 | 33 | 34 | 30 | 1
| 1 | 2 | 32 | 40 | 34.7913 | 33 | 30.4 | 35 | 31 | 33 | 35 | 31.015 | 29 | 32 | 33.18 | 33 | 34.8349 | 32 | 39 | 37.6 | 35.0562 | 33 | 34 | 33 | 35 | 30 | 26
| 1 | 3 | 32 | 40 | 34.3643 | 34 | 29.8 | 35 | 31 | 33 | 35 | 31.364 | 29 | 32 | 32.92 | 34 | 34.6757 | 31 | 39 | 36.3 | 33.4872 | 31 | 33 | 32 | 35 | 30 | 16
| 1 | 4 | 32 | 40 | 34.092 | 33 | 31 | 35 | 31 | 33 | 35 | 31.521 | 29 | 33 | 32.57 | 33 | 34.788 | 32 | 39 | 35.7 | 33.4406 | 31 | 33 | 32 | 34 | 30 | 32.8153
| 1 | 5 | 32 | 40 | 34.49 | 34 | 30.9 | 35 | 32 | 34 | 34 | 31.217 | 29 | 32 | 33.13 | 33 | 34.2288 | 31 | 39 | 36.3 | 33.117 | 32 | 33 | 33 | 34 | 30 | 86
| 1 | 6 | 32 | 35 | 34.545 | 33 | 30.2 | 35 | 30 | 32 | 34 | 31.463 | 29 | 32 | 33.49 | 34 | 33.7592 | 32 | 40 | 34.1 | 34.097 | 32 | 33 | 32 | 36 | 31
| Mean |  | 31.8333 | 39.1667 | 34.4337 | 33.5 | 30.6333 | 35 | 30.8333 | 33.3333 | 34.6667 | 31.32 | 29.1667 | 32.1667 | 32.9867 | 33.3333 | 34.2794 | 31.8333 | 39.1667 | 36.1167 | 33.9998 | 31.8333 | 33.3333 | 32.5 | 34.6667 | 30.1667
| Median |  | 32 | 40 | 34.4272 | 33.5 | 30.65 | 35 | 31 | 33 | 35 | 31.352 | 29 | 32 | 33.025 | 33 | 34.4522 | 32 | 39 | 36.3 | 33.7921 | 32 | 33 | 32.5 | 34.5 | 30
| Std Dev. |  | 0.40825 | 2.0412 | 0.23584 | 0.54772 | 0.61536 | 0 | 0.75277 | 1.0328 | 0.5164 | 0.18266 | 0.40825 | 0.40825 | 0.35115 | 0.5164 | 0.59808 | 0.75277 | 0.40825 | 1.1703 | 0.79016 | 0.75277 | 0.5164 | 0.54772 | 0.8165 | 0.40825
| Rel.Std.Dev. |  | 1.282% | 5.212% | 0.6849% | 1.635% | 2.009% | 0% | 2.441% | 3.098% | 1.49% | 0.5832% | 1.4% | 1.269% | 1.065% | 1.549% | 1.745% | 2.365% | 1.042% | 3.24% | 2.324% | 2.365% | 1.549% | 1.685% | 2.355% | 1.353%
| PDM3 |  | -2.992% | 19.35% | 4.932% | 2.087% | -6.649% | 6.658% | -6.04% | 1.579% | 5.642% | -4.557% | -11.12% | -1.977% | 0.5223% | 1.579% | 4.462% | -2.992% | 19.35% | 10.06% | 3.61% | -2.992% | 1.579% | -0.9608% | 5.642% | -8.071%
| Z-Score (Absolute) |  | 0.71 | 2.78 | 0.52 | 0.08 | 1.29 | 0.79 | 1.19 | 0 | 0.63 | 0.96 | 1.98 | 0.56 | 0.17 | 0 | 0.45 | 0.71 | 2.78 | 1.32 | 0.32 | 0.71 | 0 | 0.4 | 0.63 | 1.51
| Analytical results for Cs in OREAS 236 (Certified Value 0.73 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.18 | 1.19 | 1.25 | 1.26 | 1.27 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 0.71 | 0.7 | 0.87 | 0.82 | 0.7 | 0.72 | 0.76 | 0.74 | 0.53 | 0.61 | 0.75116 | 0.73 | 0.67 | 0.7 | 0.75 | 0.73 | [8 empty] | 1
| 1 | 2 | 0.78 | 0.72 | 0.86 | 0.79 | 0.7 | 0.74 | 0.74 | 0.74 | 0.53 | 0.67 | 0.76158 | 0.73 | 0.67 | 0.68 | 0.77 | 0.73 | [8 empty] | 27
| 1 | 3 | 0.73 | 0.72 | 0.84 | 0.81 | 0.7 | 0.74 | 0.75 | 0.76 | 0.54 | 0.66 | 0.72197 | 0.7 | 0.7 | 0.69 | 0.75 | 0.73 | [8 empty] | 16
| 1 | 4 | 0.72 | 0.72 | 0.87 | 0.79 | 0.7 | 0.74 | 0.73 | 0.75 | 0.55 | 0.68 | 0.71676 | 0.72 | 0.68 | 0.7 | 0.75 | 0.72 | [8 empty] | 0.72621
| 1 | 5 | 0.73 | 0.74 | 0.87 | 0.8 | 0.7 | 0.73 | 0.75 | 0.73 | 0.53 | 0.66 | 0.70703 | 0.71 | 0.69 | 0.7 | 0.75 | 0.73 | [8 empty] | 87
| 1 | 6 | 0.75 | 0.74 | 0.87 | 0.78 | 0.7 | 0.72 | 0.72 | 0.72 | 0.54 | 0.64 | 0.72718 | 0.73 | 0.7 | 0.71 | 0.78 | 0.73
| Mean |  | 0.73667 | 0.72333 | 0.86333 | 0.79833 | 0.7 | 0.73167 | 0.74167 | 0.74 | 0.53667 | 0.65333 | 0.73095 | 0.72 | 0.685 | 0.69667 | 0.75833 | 0.72833
| Median |  | 0.73 | 0.72 | 0.87 | 0.795 | 0.7 | 0.735 | 0.745 | 0.74 | 0.535 | 0.66 | 0.72458 | 0.725 | 0.685 | 0.7 | 0.75 | 0.73
| Std Dev. |  | 0.025033 | 0.015055 | 0.012111 | 0.01472 | 1.2162e-16 | 0.0098319 | 0.01472 | 0.014142 | 0.008165 | 0.025033 | 0.021047 | 0.012649 | 0.013784 | 0.010328 | 0.013292 | 0.0040825
| Rel.Std.Dev. |  | 3.398% | 2.081% | 1.403% | 1.844% | 1.737e-14% | 1.344% | 1.985% | 1.911% | 1.521% | 3.832% | 2.879% | 1.757% | 2.012% | 1.482% | 1.753% | 0.5605%
| PDM3 |  | 1.44% | -0.3958% | 18.88% | 9.932% | -3.609% | 0.7518% | 2.129% | 1.899% | -26.1% | -10.03% | 0.6527% | -0.8548% | -5.674% | -4.068% | 4.424% | 0.2928%
| Z-Score (Absolute) |  | 0.35 | 0.47 | 8.15 | 4.15 | NA | 0.04 | 0.66 | 0.56 | 11.97 | 4.78 | 0 | 0.67 | 2.83 | 2.11 | 1.69 | 0.16
Results from laboratory 7 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 236 (Certified Value 169 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 0
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 0
| 1 | 1 | 165 | 168 | 161.798 | 190.8 | 180 | 150 | 162.8 | 175.9 | 163 | 162.536 | 178.7 | 147 | 167.307 | 180 | 189.5891 | 182 | 161 | 165.97 | 166.6298 | 171 | 160 | 174.5 | 177.9 | 166.5 | 1
| 1 | 2 | 174.5 | 172 | 163.814 | 192.3 | 171 | 149 | 166.7 | 169.9 | 165 | 161.932 | 177.5 | 147 | 169.848 | 173 | 185.691 | 180.5 | 161 | 176.2 | 163.6388 | 176 | 162 | 173 | 173.5 | 167 | 12
| 1 | 3 | 173.5 | 173 | 164.864 | 188.1 | 170 | 151 | 167.7 | 173.1 | 166 | 158.853 | 178.9 | 143 | 169.7115 | 167 | 174.236 | 174.5 | 162 | 169.66 | 163.1226 | 172 | 159 | 172 | 174.2 | 167.5 | 16
| 1 | 4 | 175 | 173 | 164.164 | 188.4 | 173 | 152 | 163.3 | 172.7 | 165 | 162.522 | 179.5 | 143 | 169.701 | 169 | 167.5442 | 179.5 | 164 | 167.07 | 163.6328 | 174 | 161 | 172 | 178.8 | 165 | 169.4907
| 1 | 5 | 171.5 | 171 | 167.146 | 191.8 | 170 | 147 | 169.5 | 175.7 | 161 | 161.996 | 172.8 | 145 | 171.9585 | 168 | 225.5924 | 177 | 159 | 160.84 | 163.0088 | 175 | 159 | 174 | 176.1 | 166.5 | 88
| 1 | 6 | 173.5 | 165 | 168.154 | 191.3 | 170 | 153 | 164 | 172.6 | 161 | 158.954 | 176.5 | 141 | 174.3315 | 170 | 180.5649 | 179 | 160 | 162.69 | 166.6252 | 174 | 159 | 173 | 182 | 166.5
| Mean |  | 172.1667 | 170.3333 | 164.99 | 190.45 | 172.3333 | 150.3333 | 165.6667 | 173.3167 | 163.5 | 161.1322 | 177.3167 | 144.3333 | 170.4763 | 171.1667 | 187.2029 | 178.75 | 161.1667 | 167.0717 | 164.443 | 173.6667 | 160 | 173.0833 | 177.0833 | 166.5
| Median |  | 173.5 | 171.5 | 164.514 | 191.05 | 170.5 | 150.5 | 165.35 | 172.9 | 164 | 161.964 | 178.1 | 144 | 169.7797 | 169.5 | 183.1279 | 179.25 | 161 | 166.52 | 163.6358 | 174 | 159.5 | 173 | 177 | 166.5
| Std Dev. |  | 3.7103 | 3.2042 | 2.3207 | 1.7785 | 3.9328 | 2.1602 | 2.7016 | 2.2347 | 2.1679 | 1.7451 | 2.4612 | 2.4221 | 2.3952 | 4.7924 | 20.398 | 2.6599 | 1.7224 | 5.4639 | 1.7116 | 1.8619 | 1.2649 | 1.0206 | 3.1594 | 0.83666
| Rel.Std.Dev. |  | 2.155% | 1.881% | 1.407% | 0.9338% | 2.282% | 1.437% | 1.631% | 1.289% | 1.326% | 1.083% | 1.388% | 1.678% | 1.405% | 2.8% | 10.9% | 1.488% | 1.069% | 3.27% | 1.041% | 1.072% | 0.7906% | 0.5897% | 1.784% | 0.5025%
| PDM3 |  | 1.579% | 0.4971% | -2.655% | 12.37% | 1.677% | -11.3% | -2.256% | 2.257% | -3.535% | -4.932% | 4.617% | -14.84% | 0.5815% | 0.9888% | 10.45% | 5.463% | -4.911% | -1.427% | -2.978% | 2.464% | -5.6% | 2.12% | 4.48% | -1.765%
| Z-Score (Absolute) |  | 0.21 | 0.01 | 0.64 | 2.38 | 0.23 | 2.38 | 0.56 | 0.35 | 0.82 | 1.1 | 0.82 | 3.09 | 0.01 | 0.09 | 1.99 | 0.99 | 1.1 | 0.4 | 0.71 | 0.39 | 1.23 | 0.32 | 0.79 | 0.46
| Analytical results for Dy in OREAS 236 (Certified Value 2.58 ppm) | [25 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 3
| 1 | 1 | 2.23 | 2.413 | 2.69 | 2.8 | 2.5125 | [19 empty] | 1
| 1 | 2 | 2.34 | 2.41 | 2.65 | 2.9 | 2.6722 | [19 empty] | 29
| 1 | 3 | 2.33 | 2.43 | 2.87 | 2.8 | 2.5605 | [19 empty] | 16
| 1 | 4 | 2.5 | 2.447 | 2.78 | 2.9 | 2.5356 | [19 empty] | 2.5829
| 1 | 5 | 2.46 | 2.441 | 2.59 | 2.8 | 2.6416 | [19 empty] | 89
| 1 | 6 | 2.34 | 2.46 | 2.33 | 3 | 2.6539
| Mean |  | 2.3667 | 2.4335 | 2.6517 | 2.8667 | 2.5961
| Median |  | 2.34 | 2.4355 | 2.67 | 2.85 | 2.601
| Std Dev. |  | 0.097912 | 0.019624 | 0.18595 | 0.08165 | 0.067968
| Rel.Std.Dev. |  | 4.137% | 0.8064% | 7.012% | 2.848% | 2.618%
| PDM3 |  | -8.372% | -5.785% | 2.662% | 10.99% | 0.5088%
| Z-Score (Absolute) |  | 0.95 | 0.67 | 0.23 | 1.12 | 0
| Analytical results for Er in OREAS 236 (Certified Value 1.52 ppm) | [25 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 2
| 1 | 1 | 1.22 | 1.449 | 1.64 | 1.6 | 1.5468 | [19 empty] | 1
| 1 | 2 | 1.26 | 1.409 | 1.67 | 1.6 | 1.6195 | [19 empty] | 30
| 1 | 3 | 1.25 | 1.452 | 1.7 | 1.6 | 1.5815 | [19 empty] | 16
| 1 | 4 | 1.36 | 1.476 | 1.73 | 1.7 | 1.5391 | [19 empty] | 1.5241
| 1 | 5 | 1.3 | 1.444 | 1.6 | 1.8 | 1.623 | [19 empty] | 90
| 1 | 6 | 1.24 | 1.36 | 1.41 | 1.7 | 1.584
| Mean |  | 1.2717 | 1.4317 | 1.625 | 1.6667 | 1.5823
| Median |  | 1.255 | 1.4465 | 1.655 | 1.65 | 1.5827
| Std Dev. |  | 0.050761 | 0.041186 | 0.11467 | 0.08165 | 0.035125
| Rel.Std.Dev. |  | 3.992% | 2.877% | 7.057% | 4.899% | 2.22%
| PDM3 |  | -16.56% | -6.063% | 6.623% | 9.357% | 3.822%
| Z-Score (Absolute) |  | 2.48 | 1.2 | 0.34 | 0.67 | 0
| Analytical results for Eu in OREAS 236 (Certified Value 0.58 ppm) | [25 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.424 | 0.544 | 0.63 | 0.6 | 0.62458 | [19 empty] | 1
| 1 | 2 | 0.44 | 0.541 | 0.62 | 0.6 | 0.65777 | [19 empty] | 31
| 1 | 3 | 0.448 | 0.552 | 0.66 | 0.6 | 0.64069 | [19 empty] | 16
| 1 | 4 | 0.472 | 0.547 | 0.64 | 0.7 | 0.61143 | [19 empty] | 0.57988
| 1 | 5 | 0.472 | 0.563 | 0.6 | 0.6 | 0.64808 | [19 empty] | 91
| 1 | 6 | 0.44 | 0.535 | 0.51 | 0.7 | 0.65573
| Mean |  | 0.44933 | 0.547 | 0.61 | 0.63333 | 0.63971
| Median |  | 0.444 | 0.5455 | 0.625 | 0.6 | 0.64438
| Std Dev. |  | 0.019211 | 0.0096954 | 0.052915 | 0.05164 | 0.018345
| Rel.Std.Dev. |  | 4.275% | 1.772% | 8.675% | 8.154% | 2.868%
| PDM3 |  | -22.51% | -5.669% | 5.195% | 9.219% | 10.32%
| Z-Score (Absolute) |  | 3.65 | 1.43 | 0 | 0.53 | 0.67
| Analytical results for Fe in OREAS 236 (Certified Value 5.95 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 5.85 | 5.95 | 5.4717 | 5.88 | 5.78 | 5.17 | 5.73 | 6.33 | 6.15 | 6.119 | 6.12 | 6.16 | 5.0609 | 6.06 | 5.9466 | 6.21 | 5.92 | 5.7 | > 5 | 5.8 | 5.82 | 6.2 | 6.13 | 5.75 | 1
| 1 | 2 | 5.95 | 6.14 | 5.4775 | 5.8 | 5.61 | 5.18 | 5.93 | 6.24 | 6.22 | 6.0602 | 6.15 | 6.19 | 5.0358 | 5.93 | 6.0941 | 6.1 | 5.94 | 5.73 | > 5 | 5.78 | 5.81 | 6.25 | 6.33 | 5.77 | #N/A
| 1 | 3 | 6 | 6.16 | 5.4267 | 5.9 | 5.51 | 5.21 | 6.04 | 6.02 | 6.23 | 6.0188 | 6.19 | 6.09 | 5.0613 | 5.76 | 6.1674 | 5.94 | 6 | 5.61 | > 5 | 5.75 | 5.81 | 6.13 | 6.18 | 5.7 | 16
| 1 | 4 | 6.03 | 6.21 | 5.5175 | 5.8 | 5.59 | 5.14 | 5.94 | 5.98 | 6.2 | 6.1189 | 6.2 | 6.22 | 5.0151 | 5.79 | 5.9978 | 6.08 | 6.01 | 5.79 | > 5 | 5.78 | 5.78 | 6.16 | 6.25 | 5.67 | 5.951
| 1 | 5 | 5.91 | 6.13 | 5.5067 | 5.84 | 5.59 | 5.15 | 6 | 6.3 | 6.02 | 6.0612 | 6.04 | 5.93 | 5.1041 | 5.78 | 6.1011 | 6.03 | 6.1 | 5.77 | > 5 | 5.83 | 5.81 | 6.21 | 6.1 | 5.73 | 92
| 1 | 6 | 6 | 5.82 | 5.5046 | 5.85 | 5.57 | 5.22 | 5.78 | 6.15 | 6.05 | 6.0187 | 6.16 | 6 | 5.0023 | 5.89 | 6.0264 | 6.11 | 6.05 | 5.68 | > 5 | 5.84 | 5.79 | 6.18 | 6.38 | 5.77
| Mean |  | 5.9567 | 6.0683 | 5.4841 | 5.845 | 5.6083 | 5.1783 | 5.9033 | 6.17 | 6.145 | 6.0662 | 6.1433 | 6.0983 | 5.0466 | 5.8683 | 6.0556 | 6.0783 | 6.0033 | 5.7133 |  | 5.7967 | 5.8033 | 6.1883 | 6.2283 | 5.7317
| Median |  | 5.975 | 6.135 | 5.491 | 5.845 | 5.59 | 5.175 | 5.935 | 6.195 | 6.175 | 6.0607 | 6.155 | 6.125 | 5.0484 | 5.84 | 6.0602 | 6.09 | 6.005 | 5.715 |  | 5.79 | 5.81 | 6.19 | 6.215 | 5.74
| Std Dev. |  | 0.067429 | 0.15039 | 0.03329 | 0.040866 | 0.090866 | 0.031885 | 0.12275 | 0.14588 | 0.090056 | 0.04502 | 0.058195 | 0.11409 | 0.036873 | 0.11548 | 0.080166 | 0.089759 | 0.067132 | 0.06532 |  | 0.033862 | 0.015055 | 0.041673 | 0.11161 | 0.040208
| Rel.Std.Dev. |  | 1.132% | 2.478% | 0.607% | 0.6992% | 1.62% | 0.6157% | 2.079% | 2.364% | 1.466% | 0.7421% | 0.9473% | 1.871% | 0.7307% | 1.968% | 1.324% | 1.477% | 1.118% | 1.143% |  | 0.5842% | 0.2594% | 0.6734% | 1.792% | 0.7015%
| PDM3 |  | 0.09511% | 1.972% | -7.846% | -1.781% | -5.758% | -12.98% | -0.8011% | 3.68% | 3.26% | 1.935% | 3.232% | 2.476% | -15.2% | -1.389% | 1.757% | 2.14% | 0.8793% | -3.994% |  | -2.594% | -2.481% | 3.988% | 4.66% | -3.686%
| Z-Score (Absolute) |  | 0 | 0.47 | 1.99 | 0.47 | 1.47 | 3.28 | 0.22 | 0.9 | 0.79 | 0.46 | 0.79 | 0.6 | 3.84 | 0.37 | 0.42 | 0.51 | 0.2 | 1.03 | NA | 0.67 | 0.65 | 0.98 | 1.14 | 0.95
| Analytical results for Ga in OREAS 236 (Certified Value 11.6 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 10.7 | 10.8 | 12.6153 | 11.2 | 11.1 | 10 | 10.19 | 12.07 | 12.5 | 13.264 | 10.88 | 12.3 | 14.81 | 11.4 | 12.2 | 11.5 | 11.5 | 13 | 11 | 11.25 | 12.59 | 10.9 |  |  | 1
| 1 | 2 | 11.55 | 11 | 12.746 | 11.2 | 11.1 | 11 | 10.43 | 12.43 | 12.5 | 12.855 | 10.89 | 12.1 | 15.3 | 11.6 | 12.35 | 11.5 | 11.9 | 13 | 11 | 10.9 | 12.56 | 11.15 |  |  | 32
| 1 | 3 | 11.05 | 10.8 | 12.4347 | 11 | 11.4 | 11 | 10.62 | 12.5 | 12.5 | 11.784 | 10.8 | 11.9 | 15.04 | 11.8 | 11.7 | 11.3 | 12 | 13 | 11 | 11.1 | 12.2 | 11 |  |  | 16
| 1 | 4 | 12.25 | 11.4 | 12.3993 | 10.9 | 11.2 | 11 | 10.55 | 12.04 | 12 | 13.163 | 10.86 | 12.1 | 15.11 | 11.4 | 12 | 11.7 | 11.6 | 10 | 10 | 10.9 | 12.33 | 10.9 |  |  | 11.6266
| 1 | 5 | 12.35 | 11.2 | 12.6239 | 11.5 | 11.1 | 11 | 10.76 | 12.18 | 11.9 | 12.65 | 10.61 | 11.9 | 15.7 | 12 | 12 | 11.5 | 11.8 | 11 | 11 | 11.05 | 12.41 | 10.95 |  |  | 93
| 1 | 6 | 12.4 | 11 | 12.588 | 11.4 | 11.6 | 11 | 10.25 | 11.25 | 12.1 | 11.681 | 10.94 | 11.6 | 15.3 | 11.8 | 12.15 | 11.4 | 11 | 12 | 11 | 11.1 | 12.73 | 10.9
| Mean |  | 11.7167 | 11.0333 | 12.5679 | 11.2 | 11.25 | 10.8333 | 10.4667 | 12.0783 | 12.25 | 12.5662 | 10.83 | 11.9833 | 15.21 | 11.6667 | 12.0667 | 11.4833 | 11.6333 | 12 | 10.8333 | 11.05 | 12.47 | 10.9667
| Median |  | 11.9 | 11 | 12.6017 | 11.2 | 11.15 | 11 | 10.49 | 12.125 | 12.3 | 12.7525 | 10.87 | 12 | 15.205 | 11.7 | 12.075 | 11.5 | 11.7 | 12.5 | 11 | 11.075 | 12.485 | 10.925
| Std Dev. |  | 0.72915 | 0.23381 | 0.12936 | 0.22804 | 0.20736 | 0.40825 | 0.2197 | 0.4471 | 0.28107 | 0.68249 | 0.11696 | 0.24014 | 0.30173 | 0.24221 | 0.22286 | 0.13292 | 0.36148 | 1.2649 | 0.40825 | 0.13416 | 0.19298 | 0.098319
| Rel.Std.Dev. |  | 6.223% | 2.119% | 1.029% | 2.036% | 1.843% | 3.768% | 2.099% | 3.702% | 2.294% | 5.431% | 1.08% | 2.004% | 1.984% | 2.076% | 1.847% | 1.157% | 3.107% | 10.54% | 3.768% | 1.214% | 1.548% | 0.8965%
| PDM3 |  | 0.7745% | -5.103% | 8.096% | -3.669% | -3.239% | -6.823% | -9.977% | 3.885% | 5.362% | 8.081% | -6.852% | 3.068% | 30.82% | 0.3445% | 3.785% | -1.232% | 0.0578% | 3.211% | -6.823% | -4.959% | 7.254% | -5.676%
| Z-Score (Absolute) |  | 0.06 | 0.73 | 1.04 | 0.54 | 0.48 | NA | 1.39 | 0.48 | 0.67 | 1.04 | 0.97 | 0.37 | 4.1 | 0 | 0.46 | 0.21 | 0.04 | NA | NA | 0.71 | 0.93 | 0.81
Results from laboratories 7, 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 236 (Certified Value 2.24 ppm) | [25 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 3
| 1 | 1 | 1.98 | 2.138 | 2.47 | 2.3 | 2.3171 | [19 empty] | 1
| 1 | 2 | 2 | 2.123 | 2.39 | 2.4 | 2.3199 | [19 empty] | 33
| 1 | 3 | 2.04 | 2.272 | 2.51 | 2.2 | 2.3197 | [19 empty] | 16
| 1 | 4 | 2.09 | 2.197 | 2.45 | 2.3 | 2.2562 | [19 empty] | 2.2411
| 1 | 5 | 2.08 | 2.245 | 2.37 | 2.3 | 2.3073 | [19 empty] | 94
| 1 | 6 | 2.04 | 2.134 | 2.18 | 2.2 | 2.4505
| Mean |  | 2.0383 | 2.1848 | 2.395 | 2.2833 | 2.3284
| Median |  | 2.04 | 2.1675 | 2.42 | 2.3 | 2.3184
| Std Dev. |  | 0.043089 | 0.063193 | 0.11726 | 0.075277 | 0.064561
| Rel.Std.Dev. |  | 2.114% | 2.892% | 4.896% | 3.297% | 2.773%
| PDM3 |  | -9.048% | -2.511% | 6.867% | 1.884% | 3.897%
| Z-Score (Absolute) |  | 1.68 | 0.67 | 0.76 | 0 | 0.31
| Analytical results for Ge in OREAS 236 (Certified Value 0.15 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.04 | 1.07 | 1.13 | 1.14 | 1.17 | 1.18 | 1.23 | 1.25 | 1.26 | 1.27 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.5g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | 0.1 | 0.3 | < 0.1 | 0.1 | < 0.1 | < 3 | 0.13 | < 10 | 0.16 | 0.14 | 0.16 | [13 empty] | 1
| 1 | 2 | 0.12 | 0.3 | < 0.1 | 0.1 | < 0.1 | < 3 | 0.12 | < 10 | 0.16 | 0.14 | 0.16 | [13 empty] | 1
| 1 | 3 | 0.12 | 0.3 | < 0.1 | 0.1 | < 0.1 | < 3 | 0.15 | < 10 | 0.16 | 0.17 | 0.15 | [13 empty] | 16
| 1 | 4 | 0.15 | 0.3 | < 0.1 | 0.1 | < 0.1 | < 3 | 0.15 | < 10 | 0.17 | 0.15 | 0.16 | [13 empty] | 0.14633
| 1 | 5 | 0.15 | 0.3 | < 0.1 | < 0.1 | < 0.1 | < 3 | 0.14 | < 10 | 0.18 | 0.14 | 0.15 | [13 empty] | 95
| 1 | 6 | 0.15 | 0.3 | < 0.1 | < 0.1 | < 0.1 | < 3 | 0.13 | < 10 | 0.16 | 0.12 | 0.15
| Mean |  | 0.13167 | 0.3 |  | 0.1 |  |  | 0.13667 |  | 0.165 | 0.14333 | 0.155
| Median |  | 0.135 | 0.3 |  | 0.1 |  |  | 0.135 |  | 0.16 | 0.14 | 0.155
| Std Dev. |  | 0.02137 | 0 |  | 0 |  |  | 0.012111 |  | 0.0083666 | 0.01633 | 0.0054772
| Rel.Std.Dev. |  | 16.23% | 0% |  | 0% |  |  | 8.861% |  | 5.071% | 11.39% | 3.534%
| PDM3 |  | -10.02% | 105% |  | -31.66% |  |  | -6.606% |  | 12.76% | -2.05% | 5.923%
| Z-Score (Absolute) |  | 0.13 | 1.76 | 1.05 | 0.67 | 1.05 | 15.25 | 0.07 | 54.58 | 0.24 | 0 | 0.13
| Analytical results for Hf in OREAS 236 (Certified Value 0.54 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.18 | 1.19 | 1.25 | 1.26 | 1.27 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | 0.53 | < 0.2 | 0.47 | 0.54 | 0.4 | 0.44 | 0.47 | 0.65 | 0.71 | 0.3 | 0.70415 | 0.55 | 0.56 | 0.56 | 0.55 | 0.54 | [8 empty] | 1
| 1 | 2 | 0.54 | < 0.2 | 0.46 | 0.52 | 0.4 | 0.44 | 0.53 | 0.65 | 0.7 | 0.4 | 0.75809 | 0.55 | 0.57 | 0.54 | 0.55 | 0.55 | [8 empty] | 18
| 1 | 3 | 0.51 | < 0.2 | 0.45 | 0.54 | 0.5 | 0.45 | 0.54 | 0.66 | 0.71 | 0.3 | 0.70688 | 0.55 | 0.55 | 0.56 | 0.54 | 0.51 | [8 empty] | 16
| 1 | 4 | 0.52 | < 0.2 | 0.46 | 0.54 | 0.5 | 0.44 | 0.53 | 0.6 | 0.72 | 0.4 | 0.67601 | 0.57 | 0.57 | 0.53 | 0.51 | 0.54 | [8 empty] | 0.54117
| 1 | 5 | 0.53 | < 0.2 | 0.48 | 0.52 | 0.4 | 0.47 | 0.5 | 0.6 | 0.68 | 0.4 | 0.75496 | 0.57 | 0.57 | 0.54 | 0.54 | 0.51 | [8 empty] | 96
| 1 | 6 | 0.56 | < 0.2 | 0.47 | 0.49 | 0.4 | 0.41 | 0.38 | 0.66 | 0.71 | 0.4 | 0.62168 | 0.54 | 0.58 | 0.54 | 0.55 | 0.51
| Mean |  | 0.53167 |  | 0.465 | 0.525 | 0.43333 | 0.44167 | 0.49167 | 0.63667 | 0.705 | 0.36667 | 0.70363 | 0.555 | 0.56667 | 0.545 | 0.54 | 0.52667
| Median |  | 0.53 |  | 0.465 | 0.53 | 0.4 | 0.44 | 0.515 | 0.65 | 0.71 | 0.4 | 0.70551 | 0.55 | 0.57 | 0.54 | 0.545 | 0.525
| Std Dev. |  | 0.017224 |  | 0.010488 | 0.019748 | 0.05164 | 0.019408 | 0.06047 | 0.028752 | 0.013784 | 0.05164 | 0.051166 | 0.012247 | 0.010328 | 0.012247 | 0.015492 | 0.018619
| Rel.Std.Dev. |  | 3.24% |  | 2.256% | 3.762% | 11.92% | 4.394% | 12.3% | 4.516% | 1.955% | 14.08% | 7.272% | 2.207% | 1.823% | 2.247% | 2.869% | 3.535%
| PDM3 |  | -1.755% |  | -14.07% | -2.987% | -19.93% | -18.39% | -9.147% | 17.65% | 30.27% | -32.25% | 30.02% | 2.556% | 4.712% | 0.7083% | -0.2156% | -2.679%
| Z-Score (Absolute) |  | 0.07 | 7.84 | 1.27 | 0.19 | NA | 1.69 | 0.79 | 1.81 | 3.04 | NA | 3.02 | 0.34 | 0.55 | 0.16 | 0.07 | 0.16
Results from laboratories 7 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 236 (Certified Value 0.046 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.03 | 0.04 | < 5 | 0.04 | 0.23 | 0.338 | < 0.1 | 0.08 | < 0.1 | 0.1 | 0.19 | 0.04 | 0.031 | < 1 | < 2 | 0.05 | 0.047 | 0.05 | [6 empty] | 1
| 1 | 2 | 0.03 | 0.04 | < 5 | 0.04 | 0.2 | 0.356 | < 0.1 | 0.07 | < 0.1 | 0.14 | 0.19 | 0.04 | 0.032 | < 1 | < 2 | 0.05 | 0.052 | 0.04 | [6 empty] | 1
| 1 | 3 | 0.04 | 0.04 | < 5 | 0.04 | 0.2 | 0.336 | < 0.1 | 0.08 | < 0.1 | 0.13 | 0.14 | 0.04 | 0.027 | < 1 | < 2 | 0.05 | 0.059 | 0.04 | [6 empty] | 16
| 1 | 4 | 0.04 | 0.04 | < 5 | 0.04 | 0.22 | 0.27 | < 0.1 | 0.07 | < 0.1 | 0.1 | 0.2 | 0.04 | 0.033 | < 1 | < 2 | 0.06 | 0.052 | 0.05 | [6 empty] | 0.046422
| 1 | 5 | 0.04 | 0.04 | < 5 | 0.05 | 0.22 | 0.335 | < 0.1 | 0.07 | < 0.1 | 0.12 | 0.17 | 0.05 | 0.036 | < 1 | < 2 | 0.06 | 0.051 | 0.04 | [6 empty] | 97
| 1 | 6 | 0.05 | 0.04 | < 5 | 0.04 | 0.24 | 0.306 | < 0.1 | 0.07 | < 0.1 | 0.09 | 0.19 | 0.05 | 0.018 | < 1 | < 2 | 0.06 | 0.045 | 0.04
| Mean |  | 0.038333 | 0.04 |  | 0.041667 | 0.21833 | 0.3235 |  | 0.073333 |  | 0.11333 | 0.18 | 0.043333 | 0.0295 |  |  | 0.055 | 0.051 | 0.043333
| Median |  | 0.04 | 0.04 |  | 0.04 | 0.22 | 0.3355 |  | 0.07 |  | 0.11 | 0.19 | 0.04 | 0.0315 |  |  | 0.055 | 0.0515 | 0.04
| Std Dev. |  | 0.0075277 | 0 |  | 0.0040825 | 0.016021 | 0.030736 |  | 0.005164 |  | 0.019664 | 0.021909 | 0.005164 | 0.0063482 |  |  | 0.0054772 | 0.004858 | 0.005164
| Rel.Std.Dev. |  | 19.64% | 0% |  | 9.798% | 7.338% | 9.501% |  | 7.042% |  | 17.35% | 12.17% | 11.92% | 21.52% |  |  | 9.959% | 9.525% | 11.92%
| PDM3 |  | -17.42% | -13.83% |  | -10.24% | 370.3% | 596.9% |  | 57.97% |  | 144.1% | 287.7% | -6.654% | -36.45% |  |  | 18.48% | 9.861% | -6.654%
| Z-Score (Absolute) |  | 0.57 | 0.5 | 94.29 | 0.44 | 6.37 | 10.42 | 0.12 | 0.78 | 0.12 | 2.32 | 4.89 | 0.37 | 0.91 | 17.22 | 36.49 | 0.08 | 0.08 | 0.37
| Analytical results for Ho in OREAS 236 (Certified Value 0.53 ppm) | [25 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.415 | 0.5 | 0.58 | 0.6 | 0.52689 | [19 empty] | 1
| 1 | 2 | 0.43 | 0.5 | 0.56 | 0.6 | 0.55941 | [19 empty] | 19
| 1 | 3 | 0.43 | 0.5 | 0.6 | 0.6 | 0.53031 | [19 empty] | 16
| 1 | 4 | 0.455 | 0.5 | 0.58 | 0.6 | 0.51307 | [19 empty] | 0.52543
| 1 | 5 | 0.44 | 0.5 | 0.54 | 0.6 | 0.54734 | [19 empty] | 98
| 1 | 6 | 0.42 | 0.5 | 0.49 | 0.6 | 0.54601
| Mean |  | 0.43167 | 0.5 | 0.55833 | 0.6 | 0.53717
| Median |  | 0.43 | 0.5 | 0.57 | 0.6 | 0.53816
| Std Dev. |  | 0.014376 | 0 | 0.0392 | 0 | 0.016791
| Rel.Std.Dev. |  | 3.33% | 0% | 7.021% | 0% | 3.126%
| PDM3 |  | -17.85% | -4.841% | 6.261% | 14.19% | 2.234%
| Z-Score (Absolute) |  | 1.91 | 0.67 | 0.38 | 1.14 | 0
| Analytical results for In in OREAS 236 (Certified Value 0.039 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.18 | 1.19 | 1.23 | 1.25 | 1.26 | 1.27 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 01g | 0.15g | 0.5g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.037 | 0.03 | < 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.045 | < 0.05 | 0.04 | 0.05 | 0.039 | < 0.05 | 27 | 0.035 | 0.037 | 0.04 | [7 empty] | 1
| 1 | 2 | 0.038 | 0.03 | < 5 | 0.04 | 0.04 | 0.04 | 0.03 | 0.04 | < 0.05 | 0.04 | 0.04 | 0.038 | < 0.05 | 25 | 0.036 | 0.038 | 0.037 | [7 empty] | 20
| 1 | 3 | 0.037 | 0.035 | < 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.041 | < 0.05 | 0.04 | 0.04 | 0.037 | < 0.05 | 26 | 0.035 | 0.035 | 0.039 | [7 empty] | 16
| 1 | 4 | 0.034 | 0.035 | < 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.049 | < 0.05 | 0.04 | 0.04 | 0.04 | < 0.05 | 29 | 0.036 | 0.037 | 0.037 | [7 empty] | 0.038569
| 1 | 5 | 0.038 | 0.03 | < 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046 | < 0.05 | 0.03 | 0.04 | 0.037 | < 0.05 | 28 | 0.036 | 0.036 | 0.034 | [7 empty] | 99
| 1 | 6 | 0.039 | 0.03 | < 5 | 0.04 | 0.03 | 0.04 | 0.04 | 0.029 | < 0.05 | 0.04 | 0.04 | 0.043 | < 0.05 | 25 | 0.037 | 0.037 | 0.038
| Mean |  | 0.037167 | 0.031667 |  | 0.04 | 0.038333 | 0.04 | 0.038333 | 0.041667 |  | 0.038333 | 0.041667 | 0.039 |  | 26.6667 | 0.035833 | 0.036667 | 0.0375
| Median |  | 0.0375 | 0.03 |  | 0.04 | 0.04 | 0.04 | 0.04 | 0.043 |  | 0.04 | 0.04 | 0.0385 |  | 26.5 | 0.036 | 0.037 | 0.0375
| Std Dev. |  | 0.0017224 | 0.002582 |  | 0 | 0.0040825 | 0 | 0.0040825 | 0.0070333 |  | 0.0040825 | 0.0040825 | 0.0022804 |  | 1.633 | 0.00075277 | 0.0010328 | 0.0020736
| Rel.Std.Dev. |  | 4.634% | 8.154% |  | 0% | 10.65% | 0% | 10.65% | 16.88% |  | 10.65% | 9.798% | 5.847% |  | 6.124% | 2.101% | 2.817% | 5.53%
| PDM3 |  | -3.637% | -17.9% |  | 3.709% | -0.6122% | 3.709% | -0.6122% | 8.03% |  | -0.6122% | 8.03% | 1.116% |  | 69039% | -7.094% | -4.933% | -2.773%
| Z-Score (Absolute) |  | 0.47 | 2.7 | 995.95 | 0.67 | 0 | 0.67 | 0 | 1.35 | 5.39 | 0 | 1.35 | 0.27 | 5.39 | 10773.43 | 1.01 | 0.67 | 0.34
| Analytical results for K in OREAS 236 (Certified Value 0.134 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 0.13 | 0.13 | 0.11783 | 0.17 | 0.14 | 0.15 | 0.1293 | 0.14 | 0.138 | 0.16223 | 0.1 | 0.12 | 0.16535 | 0.15 | 0.22373 | 0.14 | 0.137 | 0.14 | 0.13477 | 0.14 | 0.13 | 0.14 | 0.12 | 0.13 | 1
| 1 | 2 | 0.14 | 0.14 | 0.12167 | 0.17 | 0.14 | 0.15 | 0.1317 | 0.14 | 0.139 | 0.1608 | 0.09 | 0.12 | 0.15977 | 0.14 | 0.20372 | 0.14 | 0.133 | 0.14 | 0.1372 | 0.14 | 0.13 | 0.14 | 0.13 | 0.13 | #N/A
| 1 | 3 | 0.13 | 0.14 | 0.11783 | 0.17 | 0.13 | 0.15 | 0.135 | 0.14 | 0.14 | 0.16428 | 0.09 | 0.12 | 0.16718 | 0.14 | 0.20256 | 0.13 | 0.132 | 0.14 | 0.13404 | 0.14 | 0.13 | 0.14 | 0.12 | 0.12 | 16
| 1 | 4 | 0.13 | 0.14 | 0.11783 | 0.17 | 0.14 | 0.15 | 0.1296 | 0.13 | 0.139 | 0.16253 | 0.09 | 0.12 | 0.15533 | 0.14 | 0.20646 | 0.14 | 0.139 | 0.15 | 0.13727 | 0.14 | 0.13 | 0.14 | 0.12 | 0.12 | 0.13424
| 1 | 5 | 0.13 | 0.14 | 0.11783 | 0.17 | 0.14 | 0.15 | 0.1306 | 0.14 | 0.135 | 0.1611 | 0.09 | 0.12 | 0.15462 | 0.14 | 0.20621 | 0.13 | 0.139 | 0.14 | 0.14068 | 0.14 | 0.13 | 0.14 | 0.12 | 0.12 | 100
| 1 | 6 | 0.13 | 0.13 | 0.11783 | 0.17 | 0.14 | 0.15 | 0.1281 | 0.13 | 0.134 | 0.16408 | 0.1 | 0.12 | 0.1676 | 0.14 | 0.20903 | 0.14 | 0.139 | 0.14 | 0.13367 | 0.14 | 0.13 | 0.14 | 0.13 | 0.13
| Mean |  | 0.13167 | 0.13667 | 0.11847 | 0.17 | 0.13833 | 0.15 | 0.13072 | 0.13667 | 0.1375 | 0.16251 | 0.093333 | 0.12 | 0.16164 | 0.14167 | 0.20862 | 0.13667 | 0.1365 | 0.14167 | 0.13627 | 0.14 | 0.13 | 0.14 | 0.12333 | 0.125
| Median |  | 0.13 | 0.14 | 0.11783 | 0.17 | 0.14 | 0.15 | 0.1301 | 0.14 | 0.1385 | 0.16238 | 0.09 | 0.12 | 0.16256 | 0.14 | 0.20633 | 0.14 | 0.138 | 0.14 | 0.13598 | 0.14 | 0.13 | 0.14 | 0.12 | 0.125
| Std Dev. |  | 0.0040825 | 0.005164 | 0.001565 | 0 | 0.0040825 | 0 | 0.002426 | 0.005164 | 0.002429 | 0.0014535 | 0.005164 | 0 | 0.0058749 | 0.0040825 | 0.0077419 | 0.005164 | 0.0032094 | 0.0040825 | 0.0026583 | 0 | 0 | 0 | 0.005164 | 0.0054772
| Rel.Std.Dev. |  | 3.101% | 3.779% | 1.321% | 0% | 2.951% | 0% | 1.856% | 3.779% | 1.767% | 0.8944% | 5.533% | 0% | 3.635% | 2.882% | 3.711% | 3.779% | 2.351% | 2.882% | 1.951% | 0% | 0% | 0% | 4.187% | 4.382%
| PDM3 |  | -1.914% | 1.811% | -11.74% | 26.64% | 3.052% | 11.74% | -2.622% | 1.811% | 2.431% | 21.06% | -30.47% | -10.61% | 20.42% | 5.535% | 55.41% | 1.811% | 1.686% | 5.535% | 1.516% | 4.294% | -3.156% | 4.294% | -8.122% | -6.881%
| Z-Score (Absolute) |  | 0.62 | 0 | 2.24 | 4.11 | 0.21 | 1.64 | 0.73 | 0 | 0.1 | 3.18 | 5.34 | 2.05 | 3.08 | 0.62 | 8.86 | 0 | 0.02 | 0.62 | 0.05 | 0.41 | 0.82 | 0.41 | 1.64 | 1.44
| Analytical results for La in OREAS 236 (Certified Value 4.72 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 4.4 | 4.41 | < 5 | 5 | 4.5 | 5 | 4.705 | 5.1 | 4.86 | < 5 | 3.5 | 4.7 | 4.8503 | 4.9 | 4.9 | 4.59 | 5.2 | 4 | 5 | 4.6 | 4.3 | 4.3 |  |  | 1
| 1 | 2 | 4.8 | 4.48 | < 5 | 4.8 | 4.3 | 5 | 4.704 | 4.9 | 4.9 | < 5 | 3.46 | 4.8 | 4.99 | 5 | 5 | 4.5 | 5.3 | 5 | 5 | 4.3 | 4.3 | 4.4 |  |  | 21
| 1 | 3 | 4.6 | 4.53 | < 5 | 4.7 | 4.3 | 5 | 4.8 | 5.2 | 4.93 | < 5 | 3.47 | 4.8 | 4.8718 | 4.9 | 4.6 | 4.46 | 5.3 | 4 | 5 | 4.6 | 4.2 | 4.7 |  |  | 16
| 1 | 4 | 4.6 | 4.58 | < 5 | 4.9 | 4.4 | 5 | 4.712 | 5.1 | 4.79 | < 5 | 3.52 | 4.9 | 4.785 | 5 | 4.9 | 4.6 | 5.2 | 4 | 5 | 4.4 | 4.2 | 4.5 |  |  | 4.7234
| 1 | 5 | 4.6 | 4.56 | < 5 | 4.9 | 4.4 | 5 | 4.724 | 5 | 4.8 | < 5 | 3.47 | 4.9 | 4.8281 | 5 | 4.9 | 4.59 | 5.1 | 5 | 5 | 4.5 | 4.2 | 4.7 |  |  | 101
| 1 | 6 | 4.6 | 4.66 | < 5 | 4.9 | 4.3 | 5 | 4.668 | 4.8 | 4.72 | < 5 | 3.52 | 4.8 | 4.8694 | 4.9 | 4.9 | 4.74 | 4.6 | 5 | 5 | 4.4 | 4.4 | 4.7
| Mean |  | 4.6 | 4.5367 |  | 4.8667 | 4.3667 | 5 | 4.7188 | 5.0167 | 4.8333 |  | 3.49 | 4.8167 | 4.8658 | 4.95 | 4.8667 | 4.58 | 5.1167 | 4.5 | 5 | 4.4667 | 4.2667 | 4.55
| Median |  | 4.6 | 4.545 |  | 4.9 | 4.35 | 5 | 4.7085 | 5.05 | 4.83 |  | 3.485 | 4.8 | 4.8599 | 4.95 | 4.9 | 4.59 | 5.2 | 4.5 | 5 | 4.45 | 4.25 | 4.6
| Std Dev. |  | 0.12649 | 0.085946 |  | 0.10328 | 0.08165 | 0 | 0.043947 | 0.1472 | 0.077889 |  | 0.026833 | 0.075277 | 0.068818 | 0.054772 | 0.13663 | 0.096954 | 0.26394 | 0.54772 | 0 | 0.12111 | 0.08165 | 0.17607
| Rel.Std.Dev. |  | 2.75% | 1.894% |  | 2.122% | 1.87% | 0% | 0.9313% | 2.934% | 1.611% |  | 0.7688% | 1.563% | 1.414% | 1.107% | 2.807% | 2.117% | 5.159% | 12.17% | 0% | 2.711% | 1.914% | 3.87%
| PDM3 |  | -2.613% | -3.954% |  | 3.033% | -7.553% | 5.856% | -0.09695% | 6.209% | 2.327% |  | -26.11% | 1.974% | 3.014% | 4.797% | 3.033% | -3.036% | 8.326% | -4.73% | 5.856% | -5.436% | -9.67% | -3.671%
| Z-Score (Absolute) |  | 0 | 0.16 | 5.33 | 0.68 | 0.59 | NA | 0.3 | 1.06 | 0.59 | 5.33 | 2.82 | 0.55 | 0.67 | 0.89 | 0.68 | 0.05 | 1.31 | NA | NA | 0.34 | 0.85 | 0.13
Results from laboratories 7, 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 236 (Certified Value 9.28 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 8.8 | 9.5 | 8.812 | 12 | 11.6 | 9 | 9.7 | 9.7 | 11.577 | 9 | 9.1 | 7.28 | 9 | 10 | 8.7 | < 10 | 9 | 9 | 9.7 | 9.7 | 8.9 |  |  |  | 1
| 1 | 2 | 9.3 | 9.8 | 8.976 | 11 | 10.7 | 9 | 9.7 | 9.2 | 11.552 | 8.8 | 9 | 7.3371 | 10 | 10.2 | 8.9 | < 10 | 9 | 9 | 9.6 | 9.7 | 9.2 |  |  |  | #N/A
| 1 | 3 | 8.8 | 9.7 | 8.7937 | 12 | 10.7 | 9.1 | 9.9 | 9.2 | 10.874 | 9.1 | 8.9 | 7.0229 | 9 | 9.5 | 8.8 | < 10 | 9 | 9 | 9.7 | 9.4 | 9.1 |  |  |  | 16
| 1 | 4 | 8.9 | 10.4 | 8.944 | 10 | 11.9 | 8.9 | 9.5 | 9.7 | 11.51 | 9.1 | 9 | 6.9385 | 9 | 9.8 | 9.1 | < 10 | 10 | 9 | 9.6 | 9.5 | 8.9 |  |  |  | 9.2802
| 1 | 5 | 8.9 | 10 | 8.984 | 12 | 10.3 | 8.9 | 9.7 | 9.6 | 11.55 | 8.7 | 8.8 | 7.2117 | 9 | 9.7 | 8.8 | < 10 | 10 | 9 | 9.6 | 9.4 | 8.9 |  |  |  | 102
| 1 | 6 | 8.9 | 10.3 | 9.047 | 12 | 10.1 | 9 | 9.7 | 9.5 | 10.86 | 8.8 | 8.7 | 7.1478 | 9 | 10.1 | 9.1 | < 10 | 10 | 9 | 9.7 | 9.5 | 8.9
| Mean |  | 8.9333 | 9.95 | 8.9261 | 11.5 | 10.8833 | 8.9833 | 9.7 | 9.4833 | 11.3205 | 8.9167 | 8.9167 | 7.1563 | 9.1667 | 9.8833 | 8.9 |  | 9.5 | 9 | 9.65 | 9.5333 | 8.9833
| Median |  | 8.9 | 9.9 | 8.96 | 12 | 10.7 | 9 | 9.7 | 9.55 | 11.53 | 8.9 | 8.95 | 7.1798 | 9 | 9.9 | 8.85 |  | 9.5 | 9 | 9.65 | 9.5 | 8.9
| Std Dev. |  | 0.18619 | 0.35071 | 0.10133 | 0.83666 | 0.71671 | 0.075277 | 0.12649 | 0.23166 | 0.35196 | 0.17224 | 0.1472 | 0.15254 | 0.40825 | 0.26394 | 0.16733 |  | 0.54772 | 0 | 0.054772 | 0.13663 | 0.13292
| Rel.Std.Dev. |  | 2.084% | 3.525% | 1.135% | 7.275% | 6.585% | 0.838% | 1.304% | 2.443% | 3.109% | 1.932% | 1.651% | 2.132% | 4.454% | 2.671% | 1.88% |  | 5.766% | 0% | 0.5676% | 1.433% | 1.48%
| PDM3 |  | -3.738% | 7.217% | -3.816% | 23.92% | 17.27% | -3.199% | 4.523% | 2.189% | 21.99% | -3.917% | -3.917% | -22.89% | -1.224% | 6.499% | -4.097% |  | 2.368% | -3.019% | 3.985% | 2.728% | -3.199%
| Z-Score (Absolute) |  | 0.06 | 1.19 | 0.07 | NA | 2.33 | 0 | 0.88 | 0.61 | 2.87 | 0.08 | 0.08 | 2.24 | NA | 1.1 | 0.1 | 4.88 | NA | NA | 0.82 | 0.67 | 0
Results from laboratories 4, 16, 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 236 (Certified Value 0.18 ppm) | [25 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 40g | 0.25g | 0.25g | 25g | 25g | 0.5g | 0.5g | [17 empty] | 3
| 1 | 1 | 0.145 | 0.19 | 0.17 | 0.175 | 0.2 | 0.2 | 0.20448 | [17 empty] | 1
| 1 | 2 | 0.155 | 0.18 | 0.17 | 0.172 | 0.19 | 0.2 | 0.20964 | [17 empty] | 22
| 1 | 3 | 0.16 | 0.18 | 0.19 | 0.173 | 0.21 | 0.2 | 0.19852 | [17 empty] | 16
| 1 | 4 | 0.165 | 0.18 | 0.17 | 0.179 | 0.2 | 0.2 | 0.20991 | [17 empty] | 0.18221
| 1 | 5 | 0.16 | 0.19 | 0.19 | 0.185 | 0.19 | 0.2 | 0.19892 | [17 empty] | 103
| 1 | 6 | 0.155 | 0.19 | 0.18 | 0.17 | 0.17 | 0.2 | 0.20408
| Mean |  | 0.15667 | 0.185 | 0.17833 | 0.17567 | 0.19333 | 0.2 | 0.20426
| Median |  | 0.1575 | 0.185 | 0.175 | 0.174 | 0.195 | 0.2 | 0.20428
| Std Dev. |  | 0.0068313 | 0.0054772 | 0.0098319 | 0.0055015 | 0.013663 | 3.0405e-17 | 0.0049504
| Rel.Std.Dev. |  | 4.36% | 2.961% | 5.513% | 3.132% | 7.067% | 1.52e-14% | 2.424%
| PDM3 |  | -14.02% | 1.531% | -2.127% | -3.591% | 6.105% | 9.764% | 12.1%
| Z-Score (Absolute) |  | 1.91 | 0.25 | 0.25 | 0.46 | 0.89 | NA | 1.72
Results from laboratory 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 236 (Certified Value 1.72 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 1.72 | 1.7 | 1.6 | 1.63 | 1.68 | 1.68 | 1.74 | 1.69 | 1.8055 | 1.73 | 1.78 | 1.6486 | 1.76 | 1.6886 | 1.77 | 1.67 | 1.71 | 1.6435 | 1.74 | 1.76 | 1.78 | 1.76 | 1.59 |  | 1
| 1 | 2 | 1.75 | 1.76 | 1.6108 | 1.63 | 1.7 | 1.74 | 1.72 | 1.71 | 1.7755 | 1.7 | 1.79 | 1.6773 | 1.75 | 1.7503 | 1.74 | 1.67 | 1.74 | 1.6355 | 1.75 | 1.76 | 1.8 | 1.82 | 1.6 |  | #N/A
| 1 | 3 | 1.76 | 1.76 | 1.5933 | 1.63 | 1.69 | 1.76 | 1.66 | 1.72 | 1.7777 | 1.71 | 1.78 | 1.6491 | 1.76 | 1.7632 | 1.69 | 1.7 | 1.72 | 1.6339 | 1.71 | 1.74 | 1.76 | 1.78 | 1.59 |  | 16
| 1 | 4 | 1.76 | 1.77 | 1.6067 | 1.64 | 1.67 | 1.73 | 1.65 | 1.71 | 1.8054 | 1.73 | 1.86 | 1.6194 | 1.73 | 1.7035 | 1.73 | 1.7 | 1.75 | 1.6422 | 1.77 | 1.75 | 1.77 | 1.8 | 1.58 |  | 1.7239
| 1 | 5 | 1.73 | 1.75 | 1.6258 | 1.63 | 1.69 | 1.72 | 1.73 | 1.7 | 1.7759 | 1.71 | 1.79 | 1.6552 | 1.76 | 1.7407 | 1.72 | 1.69 | 1.74 | 1.6364 | 1.77 | 1.74 | 1.79 | 1.76 | 1.59 |  | 104
| 1 | 6 | 1.76 | 1.67 | 1.6283 | 1.63 | 1.67 | 1.68 | 1.69 | 1.69 | 1.7778 | 1.73 | 1.77 | 1.6899 | 1.75 | 1.7053 | 1.74 | 1.68 | 1.73 | 1.674 | 1.75 | 1.76 | 1.78 | 1.83 | 1.61
| Mean |  | 1.7467 | 1.735 | 1.6108 | 1.6317 | 1.6833 | 1.7183 | 1.6983 | 1.7033 | 1.7863 | 1.7183 | 1.795 | 1.6566 | 1.7517 | 1.7253 | 1.7317 | 1.685 | 1.7317 | 1.6443 | 1.7483 | 1.7517 | 1.78 | 1.7917 | 1.5933
| Median |  | 1.755 | 1.755 | 1.6087 | 1.63 | 1.685 | 1.725 | 1.705 | 1.705 | 1.7777 | 1.72 | 1.785 | 1.6521 | 1.755 | 1.723 | 1.735 | 1.685 | 1.735 | 1.6393 | 1.75 | 1.755 | 1.78 | 1.79 | 1.59
| Std Dev. |  | 0.017512 | 0.040373 | 0.013944 | 0.0040825 | 0.012111 | 0.032506 | 0.037639 | 0.012111 | 0.014868 | 0.013292 | 0.032711 | 0.024679 | 0.01169 | 0.030071 | 0.026394 | 0.013784 | 0.01472 | 0.01508 | 0.022286 | 0.0098319 | 0.014142 | 0.029944 | 0.010328
| Rel.Std.Dev. |  | 1.003% | 2.327% | 0.8657% | 0.2502% | 0.7194% | 1.892% | 2.216% | 0.711% | 0.8323% | 0.7735% | 1.822% | 1.49% | 0.6674% | 1.743% | 1.524% | 0.818% | 0.85% | 0.9171% | 1.275% | 0.5613% | 0.7945% | 1.671% | 0.6482%
| PDM3 |  | 1.323% | 0.6461% | -6.557% | -5.348% | -2.351% | -0.3207% | -1.481% | -1.191% | 3.622% | -0.3207% | 4.127% | -3.903% | 1.613% | 0.08246% | 0.4527% | -2.254% | 0.4527% | -4.617% | 1.42% | 1.613% | 3.257% | 3.933% | -7.572%
| Z-Score (Absolute) |  | 0.53 | 0.24 | 2.86 | 2.34 | 1.05 | 0.17 | 0.67 | 0.55 | 1.53 | 0.17 | 1.74 | 1.72 | 0.66 | 0 | 0.16 | 1.01 | 0.16 | 2.03 | 0.58 | 0.66 | 1.37 | 1.66 | 3.3
| Analytical results for Mn in OREAS 236 (Certified Value 0.071 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 0.0677 | 0.0732 | 0.064191 | 0.078 | 0.0691 | 0.0743 | 0.0677 | 0.0754 | 0.0773 | 0.079022 | 0.0723 | 0.075 | 0.066414 | 0.0741 | < 0.08 | 0.0746 | 0.067 | 0.073 | 0.069723 | 0.0664 | 0.0721 | 0.0689 | 0.0682 | 0.0657 | 1
| 1 | 2 | 0.0674 | 0.0756 | 0.064711 | 0.0766 | 0.0678 | 0.0746 | 0.067 | 0.075 | 0.0769 | 0.07837 | 0.0715 | 0.0736 | 0.066139 | 0.0724 | 0.07876 | 0.0726 | 0.0663 | 0.0732 | 0.073288 | 0.0667 | 0.0721 | 0.0696 | 0.0705 | 0.0664 | 23
| 1 | 3 | 0.0675 | 0.0759 | 0.063981 | 0.0755 | 0.066 | 0.0747 | 0.0671 | 0.0725 | 0.078 | 0.077701 | 0.0712 | 0.0731 | 0.067473 | 0.0739 | 0.07915 | 0.0708 | 0.0665 | 0.0718 | 0.070441 | 0.0663 | 0.072 | 0.069 | 0.0687 | 0.066 | 16
| 1 | 4 | 0.0677 | 0.0769 | 0.064326 | 0.0753 | 0.0676 | 0.0734 | 0.067 | 0.072 | 0.0756 | 0.079012 | 0.073 | 0.0743 | 0.065359 | 0.0725 | < 0.08 | 0.0725 | 0.0689 | 0.0738 | 0.071966 | 0.0674 | 0.0721 | 0.0689 | 0.0694 | 0.0654 | 0.071399
| 1 | 5 | 0.0669 | 0.0755 | 0.064953 | 0.0797 | 0.0681 | 0.074 | 0.0675 | 0.0757 | 0.0742 | 0.078064 | 0.0705 | 0.0721 | 0.066662 | 0.0735 | 0.07814 | 0.0722 | 0.0688 | 0.0737 | 0.073142 | 0.0683 | 0.0719 | 0.0693 | 0.068 | 0.066 | 105
| 1 | 6 | 0.0696 | 0.0715 | 0.064949 | 0.0768 | 0.0672 | 0.0754 | 0.0676 | 0.0732 | 0.0744 | 0.077802 | 0.0722 | 0.0708 | 0.066051 | 0.0735 | < 0.08 | 0.0727 | 0.0684 | 0.0729 | 0.071159 | 0.068 | 0.072 | 0.0691 | 0.0709 | 0.0664
| Mean |  | 0.0678 | 0.074767 | 0.064518 | 0.076983 | 0.067633 | 0.0744 | 0.067317 | 0.073967 | 0.076067 | 0.078328 | 0.071783 | 0.07315 | 0.06635 | 0.073317 | 0.078683 | 0.072567 | 0.06765 | 0.073067 | 0.07162 | 0.067183 | 0.072033 | 0.069133 | 0.069283 | 0.065983
| Median |  | 0.0676 | 0.07555 | 0.064518 | 0.0767 | 0.0677 | 0.07445 | 0.0673 | 0.0741 | 0.07625 | 0.078217 | 0.07185 | 0.07335 | 0.066277 | 0.0735 | 0.07876 | 0.07255 | 0.0677 | 0.0731 | 0.071562 | 0.06705 | 0.07205 | 0.06905 | 0.06905 | 0.066
| Std Dev. |  | 0.00092952 | 0.0020096 | 0.00041079 | 0.001651 | 0.001025 | 0.00067823 | 0.00031885 | 0.0015958 | 0.0015769 | 0.00058141 | 0.00089312 | 0.0015215 | 0.00070411 | 0.0007111 | 0.00050935 | 0.0012176 | 0.0011845 | 0.00072019 | 0.0014436 | 0.00084715 | 8.165e-05 | 0.00027325 | 0.0012057 | 0.000392
| Rel.Std.Dev. |  | 1.371% | 2.688% | 0.6367% | 2.145% | 1.516% | 0.9116% | 0.4737% | 2.157% | 2.073% | 0.7423% | 1.244% | 2.08% | 1.061% | 0.9699% | 0.6473% | 1.678% | 1.751% | 0.9857% | 2.016% | 1.261% | 0.1133% | 0.3953% | 1.74% | 0.5941%
| PDM3 |  | -5.041% | 4.716% | -9.637% | 7.821% | -5.275% | 4.203% | -5.718% | 3.596% | 6.537% | 9.705% | 0.5378% | 2.452% | -7.072% | 2.685% | 10.2% | 1.635% | -5.251% | 2.335% | 0.3086% | -5.905% | 0.8879% | -3.174% | -2.964% | -7.586%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 1.39 | 1.02 | 0.79 | 0.52 | 0.85 | 0.44 | 0.85 | 1.28 | 0.02 | 0.28 | 1.04 | 0.31 | 2.52 | 0.17 | 0.78 | 0.26 | 0.02 | 0.87 | 0.06 | 0.5 | 0.47 | 1.11
| Analytical results for Mo in OREAS 236 (Certified Value 1.49 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 1.33 | 1.4 | < 2 | 1.52 | 1.56 | 1.5 | 1.4 | 1.61 | 1.3 | < 5 | 1.5 | 1.47 | 1.49 | < 1 | 1.52 | 1.4 | 1.6 | < 10 | < 2 | 3 | 1.34 | 1.51 | 1.43 |  | 1
| 1 | 2 | 1.53 | 1.4 | < 2 | 1.48 | 1.52 | 1.6 | 1.4 | 1.52 | 1.4 | < 5 | 1.5 | 1.53 | 1.53 | < 1 | 1.47 | 1.5 | 1.6 | < 10 | < 2 | 3 | 1.36 | 1.51 | 1.44 |  | 24
| 1 | 3 | 1.52 | 1.4 | < 2 | 1.5 | 1.58 | 1.5 | 1.5 | 1.5 | 1.4 | < 5 | 1.5 | 1.48 | 1.51 | < 1 | 1.42 | 1.4 | 1.63 | < 10 | < 2 | 2 | 1.33 | 1.62 | 1.45 |  | 16
| 1 | 4 | 1.56 | 1.5 | < 2 | 1.52 | 1.54 | 1.6 | 1.5 | 1.54 | 1.2 | < 5 | 1.5 | 1.46 | 1.44 | < 1 | 1.49 | 1.4 | 1.58 | < 10 | < 2 | 2 | 1.33 | 1.57 | 1.4 |  | 1.4887
| 1 | 5 | 1.43 | 1.5 | < 2 | 1.54 | 1.57 | 1.5 | 1.4 | 1.53 | 1.3 | < 5 | 1.4 | 1.48 | 1.55 | < 1 | 1.48 | 1.5 | 1.65 | < 10 | < 2 | 2 | 1.38 | 1.58 | 1.44 |  | 106
| 1 | 6 | 1.47 | 1.5 | < 2 | 1.51 | 1.51 | 1.5 | 1.4 | 1.54 | 1.2 | < 5 | 1.5 | 1.45 | 1.49 | < 1 | 1.5 | 1.5 | 1.47 | < 10 | < 2 | 2 | 1.34 | 1.59 | 1.42
| Mean |  | 1.4733 | 1.45 |  | 1.5117 | 1.5467 | 1.5333 | 1.4333 | 1.54 | 1.3 |  | 1.4833 | 1.4783 | 1.5017 |  | 1.48 | 1.45 | 1.5883 |  |  | 2.3333 | 1.3467 | 1.5633 | 1.43
| Median |  | 1.495 | 1.45 |  | 1.515 | 1.55 | 1.5 | 1.4 | 1.535 | 1.3 |  | 1.5 | 1.475 | 1.5 |  | 1.485 | 1.45 | 1.6 |  |  | 2 | 1.34 | 1.575 | 1.435
| Std Dev. |  | 0.084063 | 0.054772 |  | 0.020412 | 0.028048 | 0.05164 | 0.05164 | 0.037417 | 0.089443 |  | 0.040825 | 0.027869 | 0.038166 |  | 0.034059 | 0.054772 | 0.063061 |  |  | 0.5164 | 0.019664 | 0.044572 | 0.017889
| Rel.Std.Dev. |  | 5.706% | 3.777% |  | 1.35% | 1.813% | 3.368% | 3.603% | 2.43% | 6.88% |  | 2.752% | 1.885% | 2.542% |  | 2.301% | 3.777% | 3.97% |  |  | 22.13% | 1.46% | 2.851% | 1.251%
| PDM3 |  | -1.034% | -2.601% |  | 1.541% | 3.892% | 2.996% | -3.721% | 3.444% | -12.68% |  | -0.3624% | -0.6983% | 0.869% |  | -0.5864% | -2.601% | 6.691% |  |  | 56.73% | -9.543% | 5.011% | -3.945%
| Z-Score (Absolute) |  | 0.07 | 0.34 | 5.62 | 0.38 | 0.79 | 0.64 | 0.54 | 0.71 | 2.1 | 11.97 | 0.05 | 0.01 | 0.26 | 11.48 | 0.01 | 0.34 | 1.28 | 41.29 | 5.62 | NA | 1.55 | 0.99 | 0.58
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 236 (Certified Value 0.265 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 3
| 1 | 1 | 0.26 | 0.28 | 0.21533 | 0.25 | 0.27 | 0.41 | 0.26 | 0.26 | 0.25 | 0.26322 | 0.28 | 0.225 | 0.30204 | 0.28 | 0.37233 | 0.26 | 0.29 | 0.276 | 0.24911 | 0.27 | 0.25 | 0.27 | 0.28 | 0.27 | 1
| 1 | 2 | 0.26 | 0.28 | 0.21533 | 0.24 | 0.26 | 0.41 | 0.26 | 0.27 | 0.24 | 0.26138 | 0.28 | 0.23 | 0.29983 | 0.27 | 0.3322 | 0.26 | 0.28 | 0.273 | 0.24622 | 0.27 | 0.26 | 0.27 | 0.29 | 0.27 | #N/A
| 1 | 3 | 0.26 | 0.29 | 0.21533 | 0.25 | 0.26 | 0.41 | 0.26 | 0.26 | 0.24 | 0.26397 | 0.28 | 0.228 | 0.30333 | 0.27 | 0.29785 | 0.25 | 0.28 | 0.271 | 0.25098 | 0.27 | 0.25 | 0.27 | 0.28 | 0.26 | 16
| 1 | 4 | 0.26 | 0.29 | 0.21533 | 0.24 | 0.26 | 0.42 | 0.26 | 0.26 | 0.24 | 0.26352 | 0.28 | 0.229 | 0.29484 | 0.27 | 0.348 | 0.26 | 0.28 | 0.282 | 0.25721 | 0.27 | 0.25 | 0.27 | 0.29 | 0.26 | 0.26515
| 1 | 5 | 0.26 | 0.28 | 0.21533 | 0.25 | 0.26 | 0.42 | 0.26 | 0.27 | 0.25 | 0.26187 | 0.27 | 0.224 | 0.29881 | 0.27 | 0.37945 | 0.26 | 0.29 | 0.276 | 0.26183 | 0.28 | 0.24 | 0.27 | 0.28 | 0.26 | 107
| 1 | 6 | 0.26 | 0.26 | 0.21667 | 0.24 | 0.26 | 0.42 | 0.25 | 0.26 | 0.24 | 0.26365 | 0.28 | 0.228 | 0.30606 | 0.27 | 0.37018 | 0.26 | 0.28 | 0.275 | 0.25391 | 0.28 | 0.26 | 0.27 | 0.3 | 0.26
| Mean |  | 0.26 | 0.28 | 0.21556 | 0.245 | 0.26167 | 0.415 | 0.25833 | 0.26333 | 0.24333 | 0.26294 | 0.27833 | 0.22733 | 0.30082 | 0.27167 | 0.35 | 0.25833 | 0.28333 | 0.2755 | 0.25321 | 0.27333 | 0.25167 | 0.27 | 0.28667 | 0.26333
| Median |  | 0.26 | 0.28 | 0.21533 | 0.245 | 0.26 | 0.415 | 0.26 | 0.26 | 0.24 | 0.26337 | 0.28 | 0.228 | 0.30094 | 0.27 | 0.35909 | 0.26 | 0.28 | 0.2755 | 0.25245 | 0.27 | 0.25 | 0.27 | 0.285 | 0.26
| Std Dev. |  | 0 | 0.010954 | 0.00054433 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0040825 | 0.005164 | 0.005164 | 0.0010546 | 0.0040825 | 0.0023381 | 0.0038994 | 0.0040825 | 0.031028 | 0.0040825 | 0.005164 | 0.0037283 | 0.0056805 | 0.005164 | 0.0075277 | 0 | 0.008165 | 0.005164
| Rel.Std.Dev. |  | 0% | 3.912% | 0.2525% | 2.236% | 1.56% | 1.32% | 1.58% | 1.961% | 2.122% | 0.4011% | 1.467% | 1.028% | 1.296% | 1.503% | 8.865% | 1.58% | 1.823% | 1.353% | 2.243% | 1.889% | 2.991% | 0% | 2.848% | 1.961%
| PDM3 |  | -1.942% | 5.601% | -18.7% | -7.599% | -1.313% | 56.52% | -2.571% | -0.6848% | -8.228% | -0.834% | 4.972% | -14.26% | 13.45% | 2.458% | 32% | -2.571% | 6.858% | 3.904% | -4.503% | 3.087% | -5.085% | 1.829% | 8.115% | -0.6848%
| Z-Score (Absolute) |  | 0.19 | 0.94 | 2.7 | 1.04 | 0.09 | 8.58 | 0.28 | 0 | 1.13 | 0.02 | 0.85 | 2.04 | 2.12 | 0.47 | 4.9 | 0.28 | 1.13 | 0.69 | 0.57 | 0.57 | 0.66 | 0.38 | 1.32 | 0
| Analytical results for Nb in OREAS 236 (Certified Value 0.16 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.16 | < 0.2 | < 5 | 0.15 | 0.21 | 0.2 | 0.23 | < 0.05 | 0.3 | 0.13 | 0.2 | 35.6 | 0.21 | 0.3 | < 10 | < 1 | 0.15 | 0.09 | 0.11 | [5 empty] | 1
| 1 | 2 | 0.15 | < 0.2 | < 5 | 0.14 | 0.21 | 0.2 | 0.24 | < 0.05 | 0.3 | 0.16 | 0.2 | 36.86 | 0.2 | 0.3 | < 10 | < 1 | 0.14 | 0.1 | 0.13 | [5 empty] | 25
| 1 | 3 | 0.14 | < 0.2 | < 5 | 0.14 | 0.19 | 0.2 | 0.23 | < 0.05 | 0.3 | 0.15 | 0.2 | 35.12 | 0.17 | 0.3 | < 10 | < 1 | 0.14 | 0.09 | 0.1 | [5 empty] | 16
| 1 | 4 | 0.13 | < 0.2 | < 5 | 0.14 | 0.21 | 0.2 | 0.26 | < 0.05 | 0.3 | 0.15 | 0.1 | 38.67 | 0.18 | 0.3 | < 10 | < 1 | 0.14 | 0.08 | 0.11 | [5 empty] | 0.15752
| 1 | 5 | 0.15 | < 0.2 | < 5 | 0.15 | 0.21 | 0.2 | 0.25 | < 0.05 | 0.3 | 0.13 | 0.2 | 32.02 | 0.19 | 0.3 | < 10 | < 1 | 0.14 | 0.09 | 0.1 | [5 empty] | 108
| 1 | 6 | 0.16 | < 0.2 | < 5 | 0.15 | 0.18 | 0.2 | 0.24 | < 0.05 | 0.3 | 0.12 | 0.1 | 36.49 | 0.2 | 0.3 | < 10 | < 1 | 0.15 | 0.1 | 0.11
| Mean |  | 0.14833 |  |  | 0.145 | 0.20167 | 0.2 | 0.24167 |  | 0.3 | 0.14 | 0.16667 | 35.7933 | 0.19167 | 0.3 |  |  | 0.14333 | 0.091667 | 0.11
| Median |  | 0.15 |  |  | 0.145 | 0.21 | 0.2 | 0.24 |  | 0.3 | 0.14 | 0.2 | 36.045 | 0.195 | 0.3 |  |  | 0.14 | 0.09 | 0.11
| Std Dev. |  | 0.01169 |  |  | 0.0054772 | 0.013292 | 3.0405e-17 | 0.01169 |  | 0 | 0.015492 | 0.05164 | 2.2193 | 0.01472 | 0 |  |  | 0.005164 | 0.0075277 | 0.010954
| Rel.Std.Dev. |  | 7.881% |  |  | 3.777% | 6.591% | 1.52e-14% | 4.837% |  | 0% | 11.07% | 30.98% | 6.2% | 7.68% | 0% |  |  | 3.603% | 8.212% | 9.959%
| PDM3 |  | -5.831% |  |  | -7.947% | 28.03% | 26.97% | 53.42% |  | 90.45% | -11.12% | 5.808% | 22623% | 21.68% | 90.45% |  |  | -9.005% | -41.81% | -30.17%
| Z-Score (Absolute) |  | 0 | 0.61 | 29.73 | 0.04 | 0.67 | NA | 1.18 | 1.56 | NA | 0.11 | NA | 450.67 | 0.55 | NA | 61.34 | 4.45 | 0.06 | 0.72 | 0.48
| Analytical results for Nd in OREAS 236 (Certified Value 6.89 ppm) | [25 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 3
| 1 | 1 | 6.46 | 6.449 | 6.95 | 7.58 | 6.8392 | [19 empty] | 1
| 1 | 2 | 6.68 | 6.595 | 6.86 | 7.34 | 6.9451 | [19 empty] | 26
| 1 | 3 | 6.48 | 6.775 | 7.25 | 7.41 | 6.8641 | [19 empty] | 16
| 1 | 4 | 6.84 | 6.668 | 7.12 | 7.54 | 6.6413 | [19 empty] | 6.8859
| 1 | 5 | 6.65 | 6.579 | 6.93 | 7.31 | 6.8726 | [19 empty] | 109
| 1 | 6 | 6.73 | 6.483 | 6.53 | 7.04 | 6.918
| Mean |  | 6.64 | 6.5915 | 6.94 | 7.37 | 6.8467
| Median |  | 6.665 | 6.587 | 6.94 | 7.375 | 6.8684
| Std Dev. |  | 0.14683 | 0.11981 | 0.24625 | 0.19391 | 0.10768
| Rel.Std.Dev. |  | 2.211% | 1.818% | 3.548% | 2.631% | 1.573%
| PDM3 |  | -3.571% | -4.275% | 0.7863% | 7.031% | -0.5684%
| Z-Score (Absolute) |  | 0.67 | 0.83 | 0.3 | 1.71 | 0
| Analytical results for Ni in OREAS 236 (Certified Value 48.8 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 44.9 | 53 | 41.5307 | 50.8 | 51.6 | 58.6 | 45.7 | 47.4 | 49 | 44.921 | 48.8 | 48.6 | 46.838 | 49 | 438.7897 | 54.2 | 58 | 53.5 | 47.2949 | 45 | 48 | 49.9 | 55.2 | 48.8 | 1
| 1 | 2 | 47.8 | 55 | 41.6125 | 50.9 | 50.6 | 58.1 | 47.1 | 45.4 | 49 | 44.914 | 48.1 | 47.6 | 47.663 | 49 | 48.5441 | 54 | 57 | 52.4 | 47.8912 | 46 | 50 | 49.6 | 55.5 | 49.1 | 10
| 1 | 3 | 46.7 | 54 | 41.4953 | 50.5 | 51.9 | 58.5 | 47.7 | 46.5 | 50 | 44.5 | 47.9 | 47.3 | 46.937 | 48 | 49.6561 | 51.7 | 58 | 53.4 | 45.8905 | 46 | 49 | 49.2 | 55.1 | 49.6 | 16
| 1 | 4 | 51.2 | 55 | 41.2263 | 49.1 | 51.7 | 58.1 | 46.7 | 45.6 | 49 | 44.674 | 49 | 47.8 | 46.739 | 48 | 47.2557 | 52.8 | 57 | 54.8 | 46.0194 | 46 | 50 | 48.6 | 55.6 | 48.1 | 48.8399
| 1 | 5 | 50.7 | 55 | 41.5792 | 51.2 | 51.5 | 57.8 | 47.7 | 47.3 | 48 | 44.145 | 47.5 | 46.8 | 47.795 | 49 | 48.3417 | 52.6 | 58 | 52.4 | 45.7082 | 46 | 49 | 49.5 | 55.2 | 49.1 | 110
| 1 | 6 | 51.7 | 52 | 41.7695 | 52.2 | 50.7 | 59.6 | 46 | 45.5 | 47 | 44.65 | 48.6 | 46 | 48.752 | 50 | 48.1222 | 53.5 | 58 | 50.4 | 46.7188 | 46 | 48 | 49.6 | 56.3 | 49.3
| Mean |  | 48.8333 | 54 | 41.5356 | 50.7833 | 51.3333 | 58.45 | 46.8167 | 46.2833 | 48.6667 | 44.634 | 48.3167 | 47.35 | 47.454 | 48.8333 | 113.4516 | 53.1333 | 57.6667 | 52.8167 | 46.5872 | 45.8333 | 49 | 49.4 | 55.4833 | 49
| Median |  | 49.25 | 54.5 | 41.5549 | 50.85 | 51.55 | 58.3 | 46.9 | 46.05 | 49 | 44.662 | 48.35 | 47.45 | 47.3 | 49 | 48.4429 | 53.15 | 58 | 52.9 | 46.3691 | 46 | 49 | 49.55 | 55.35 | 49.1
| Std Dev. |  | 2.771 | 1.2649 | 0.17878 | 1.0108 | 0.5465 | 0.63482 | 0.84479 | 0.91524 | 1.0328 | 0.28972 | 0.57764 | 0.88938 | 0.77495 | 0.75277 | 159.3844 | 0.94587 | 0.5164 | 1.4784 | 0.87166 | 0.40825 | 0.89443 | 0.45166 | 0.4446 | 0.51381
| Rel.Std.Dev. |  | 5.674% | 2.342% | 0.4304% | 1.99% | 1.065% | 1.086% | 1.804% | 1.977% | 2.122% | 0.6491% | 1.196% | 1.878% | 1.633% | 1.542% | 140.5% | 1.78% | 0.8955% | 2.799% | 1.871% | 0.8907% | 1.825% | 0.9143% | 0.8013% | 1.049%
| PDM3 |  | -0.01347% | 10.57% | -14.96% | 3.979% | 5.105% | 19.68% | -4.143% | -5.235% | -0.3547% | -8.612% | -1.071% | -3.051% | -2.838% | -0.01347% | 132.3% | 8.791% | 18.07% | 8.142% | -4.613% | -6.156% | 0.3278% | 1.147% | 13.6% | 0.3278%
| Z-Score (Absolute) |  | 0.02 | 1.44 | 2.1 | 0.53 | 0.69 | 2.71 | 0.6 | 0.75 | 0.07 | 1.22 | 0.17 | 0.45 | 0.42 | 0.02 | 18.34 | 1.2 | 2.49 | 1.11 | 0.66 | 0.88 | 0.02 | 0.14 | 1.87 | 0.02
| Analytical results for P in OREAS 236 (Certified Value 0.045 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.045 | 0.046 | 0.044467 | 0.0466 | 0.057 | 0.0429 | 0.0476 | 0.043 | 0.04451 | 0.0452 | 0.043 | 0.044553 | 0.05 | 0.07235 | 0.047 | 0.0421 | 0.045 | 0.041429 | 0.0438 | 0.0445 | 0.045 | 0.0449 | 0.044 |  | 1
| 1 | 2 | 0.045 | 0.048 | 0.045146 | 0.047 | 0.056 | 0.0453 | 0.0476 | 0.044 | 0.044087 | 0.0464 | 0.043 | 0.045458 | 0.05 | 0.07309 | 0.046 | 0.0421 | 0.051 | 0.041318 | 0.0443 | 0.0448 | 0.046 | 0.0465 | 0.044 |  | #N/A
| 1 | 3 | 0.045 | 0.048 | 0.044451 | 0.0457 | 0.057 | 0.0458 | 0.0461 | 0.045 | 0.044121 | 0.0444 | 0.043 | 0.044507 | 0.05 | 0.07423 | 0.045 | 0.0426 | 0.045 | 0.040627 | 0.044 | 0.0448 | 0.045 | 0.0458 | 0.045 |  | 16
| 1 | 4 | 0.045 | 0.048 | 0.044241 | 0.0447 | 0.056 | 0.0426 | 0.0454 | 0.042 | 0.044504 | 0.045 | 0.042 | 0.044941 | 0.05 | 0.07139 | 0.046 | 0.0423 | 0.045 | 0.040793 | 0.0448 | 0.0447 | 0.046 | 0.0458 | 0.043 |  | 0.04488
| 1 | 5 | 0.045 | 0.048 | 0.045068 | 0.0467 | 0.057 | 0.0441 | 0.0482 | 0.044 | 0.044286 | 0.0434 | 0.043 | 0.045404 | 0.05 | 0.07331 | 0.046 | 0.0421 | 0.044 | 0.040053 | 0.0441 | 0.0446 | 0.046 | 0.0446 | 0.044 |  | 111
| 1 | 6 | 0.046 | 0.046 | 0.044969 | 0.0464 | 0.058 | 0.0441 | 0.0474 | 0.042 | 0.044333 | 0.0463 | 0.042 | 0.046076 | 0.05 | 0.07266 | 0.046 | 0.0426 | 0.045 | 0.041037 | 0.0444 | 0.0449 | 0.045 | 0.0466 | 0.044
| Mean |  | 0.045167 | 0.047333 | 0.044724 | 0.046183 | 0.056833 | 0.044133 | 0.04705 | 0.043333 | 0.044307 | 0.045117 | 0.042667 | 0.045156 | 0.05 | 0.072838 | 0.046 | 0.0423 | 0.045833 | 0.040876 | 0.044233 | 0.044717 | 0.0455 | 0.0457 | 0.044
| Median |  | 0.045 | 0.048 | 0.044718 | 0.0465 | 0.057 | 0.0441 | 0.0475 | 0.0435 | 0.04431 | 0.0451 | 0.043 | 0.045172 | 0.05 | 0.072875 | 0.046 | 0.0422 | 0.045 | 0.040915 | 0.0442 | 0.04475 | 0.0455 | 0.0458 | 0.044
| Std Dev. |  | 0.00040825 | 0.0010328 | 0.00038226 | 0.00084715 | 0.00075277 | 0.001266 | 0.0010654 | 0.0012111 | 0.00018121 | 0.0011427 | 0.0005164 | 0.00060478 | 7.6012e-18 | 0.00095817 | 0.00063246 | 0.00024495 | 0.0025626 | 0.00050466 | 0.00035024 | 0.0001472 | 0.00054772 | 0.00081486 | 0.00063246
| Rel.Std.Dev. |  | 0.9039% | 2.182% | 0.8547% | 1.834% | 1.325% | 2.869% | 2.264% | 2.795% | 0.409% | 2.533% | 1.21% | 1.339% | 1.52e-14% | 1.315% | 1.375% | 0.5791% | 5.591% | 1.235% | 0.7918% | 0.3292% | 1.204% | 1.783% | 1.437%
| PDM3 |  | 0.6391% | 5.467% | -0.348% | 2.904% | 26.63% | -1.663% | 4.836% | -3.446% | -1.277% | 0.5277% | -4.931% | 0.616% | 11.41% | 62.3% | 2.496% | -5.748% | 2.125% | -8.921% | -1.44% | -0.3635% | 1.382% | 1.827% | -1.96%
| Z-Score (Absolute) |  | 0.01 | 1.44 | 0.29 | 0.68 | 7.7 | 0.67 | 1.25 | 1.2 | 0.56 | 0.03 | 1.64 | 0 | 3.19 | 18.25 | 0.56 | 1.88 | 0.45 | 2.82 | 0.61 | 0.29 | 0.23 | 0.36 | 0.76
| Analytical results for Pb in OREAS 236 (Certified Value 30.5 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 27.5 | 31 | 25.4736 | 32.8 | 29.7 | 31.4 | 30 | 35 | 26.9 | 32.791 | 30 | 33 | 30.5 | 72.7586 | 29.8 | 27.8 | 31.62 | 36.0382 | 28 | 29 | 31.2 | 29.7 | 29.9 |  | 1
| 1 | 2 | 30.6 | 31 | 26.0283 | 32.9 | 28.8 | 30.4 | 30.2 | 33.3 | 27.7 | 31.529 | 29.3 | 32.6 | 30.2 | 65.3627 | 29.2 | 28.3 | 32.79 | 34.8546 | 32 | 28 | 31.1 | 30.2 | 30 |  | 12
| 1 | 3 | 28.9 | 31 | 25.6373 | 31.5 | 29 | 30.7 | 31 | 34.2 | 28.4 | 32.74 | 30.1 | 32.3 | 30.2 | 54.1942 | 28.7 | 28.1 | 31.9 | 31.4574 | 31 | 30 | 32.7 | 29.6 | 30.1 |  | 16
| 1 | 4 | 29.9 | 32 | 25.7167 | 33.7 | 29.3 | 30.2 | 29.9 | 34.1 | 27.6 | 32.56 | 30 | 32.7 | 30.3 | 42.4379 | 29.2 | 28.6 | 32.78 | 33.4272 | 30 | 30 | 32.2 | 29.9 | 29.8 |  | 30.4578
| 1 | 5 | 29.3 | 31 | 25.5606 | 32.6 | 28.6 | 30.9 | 30.4 | 34.2 | 28.5 | 31.78 | 29 | 32.6 | 30.2 | 112.2623 | 29.4 | 27.8 | 32.32 | 32.1608 | 30 | 29 | 32.2 | 29.4 | 30.1 |  | 112
| 1 | 6 | 29.7 | 32 | 25.7343 | 32.7 | 28.7 | 30.8 | 29.7 | 34 | 27.1 | 32.312 | 29.5 | 31.8 | 30.7 | 63.0875 | 29.8 | 29 | 31.89 | 32.7297 | 31 | 29 | 32.4 | 30.2 | 29.4
| Mean |  | 29.3167 | 31.3333 | 25.6918 | 32.7 | 29.0167 | 30.7333 | 30.2 | 34.1333 | 27.7 | 32.2853 | 29.65 | 32.5 | 30.35 | 68.3505 | 29.35 | 28.2667 | 32.2167 | 33.4446 | 30.3333 | 29.1667 | 31.9667 | 29.8333 | 29.8833
| Median |  | 29.5 | 31 | 25.677 | 32.75 | 28.9 | 30.75 | 30.1 | 34.15 | 27.65 | 32.436 | 29.75 | 32.6 | 30.25 | 64.2251 | 29.3 | 28.2 | 32.11 | 33.0784 | 30.5 | 29 | 32.2 | 29.8 | 29.95
| Std Dev. |  | 1.0591 | 0.5164 | 0.19156 | 0.70711 | 0.41673 | 0.41793 | 0.46043 | 0.54283 | 0.65422 | 0.52268 | 0.45056 | 0.40988 | 0.20736 | 23.9005 | 0.41833 | 0.47188 | 0.49391 | 1.7218 | 1.3663 | 0.75277 | 0.65929 | 0.3266 | 0.26394
| Rel.Std.Dev. |  | 3.613% | 1.648% | 0.7456% | 2.162% | 1.436% | 1.36% | 1.525% | 1.59% | 2.362% | 1.619% | 1.52% | 1.261% | 0.6832% | 34.97% | 1.425% | 1.669% | 1.533% | 5.148% | 4.504% | 2.581% | 2.062% | 1.095% | 0.8832%
| PDM3 |  | -3.747% | 2.875% | -15.65% | 7.362% | -4.732% | 0.9046% | -0.8464% | 12.07% | -9.055% | 6% | -2.652% | 6.705% | -0.354% | 124.4% | -3.637% | -7.194% | 5.775% | 9.806% | -0.4087% | -4.239% | 4.954% | -2.05% | -1.886%
| Z-Score (Absolute) |  | 0.52 | 0.51 | 2.38 | 1.21 | 0.67 | 0.2 | 0.07 | 1.95 | 1.35 | 1 | 0.35 | 1.11 | 0.01 | 19.47 | 0.5 | 1.06 | 0.96 | 1.59 | 0 | 0.6 | 0.84 | 0.26 | 0.23
| Analytical results for Pd in OREAS 236 (Indicative Value < 10 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.11 | 1.13 | [19 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 1
|  |  | 40g | 25g | 25g | 01g | 0.5g | [19 empty] | 1
| 1 | 1 | < 10 | < 10 | < 5 | < 10 | 14 | [19 empty] | 1
| 1 | 2 | < 10 | < 10 | < 5 | < 10 | 14 | [19 empty] | 3
| 1 | 3 | < 10 | < 10 | < 5 | < 10 | 21 | [19 empty] | 16
| 1 | 4 | 10 | < 10 | < 5 | < 10 | 14 | [19 empty] | < 10
| 1 | 5 | < 10 | 10 | < 5 | < 10 | < 10 | [19 empty] | 9
| 1 | 6 | < 10 | < 10 | < 5 | < 10 | 19
| Mean |  | 10 | 10 |  |  | 16.4
| Median |  | 10 | 10 |  |  | 14
| Std Dev. | [5 empty] | 3.3615
| Rel.Std.Dev. | [5 empty] | 20.5%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 2.7 | 0.67 | 7.01
| Analytical results for Pr in OREAS 236 (Certified Value 1.43 ppm) | [25 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 3
| 1 | 1 | 1.32 | 1.427 | 1.71 | 1.3 | 1.3406 | [19 empty] | 1
| 1 | 2 | 1.34 | 1.425 | 1.64 | 1.3 | 1.3881 | [19 empty] | 29
| 1 | 3 | 1.36 | 1.471 | 1.72 | 1.3 | 1.3587 | [19 empty] | 16
| 1 | 4 | 1.4 | 1.432 | 1.69 | 1.4 | 1.3163 | [19 empty] | 1.4299
| 1 | 5 | 1.39 | 1.478 | 1.63 | 1.3 | 1.3327 | [19 empty] | 113
| 1 | 6 | 1.4 | 1.404 | 1.59 | 1.4 | 1.3338
| Mean |  | 1.3683 | 1.4395 | 1.6633 | 1.3333 | 1.345
| Median |  | 1.375 | 1.4295 | 1.665 | 1.3 | 1.3372
| Std Dev. |  | 0.033714 | 0.028836 | 0.051251 | 0.05164 | 0.025149
| Rel.Std.Dev. |  | 2.464% | 2.003% | 3.081% | 3.873% | 1.87%
| PDM3 |  | -4.306% | 0.671% | 16.32% | -6.754% | -5.936%
| Z-Score (Absolute) |  | 0 | 1.37 | 5.68 | 0.67 | 0.45
| Analytical results for Pt in OREAS 236 (Indicative Value 9.8 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.08 | 1.10 | 1.11 | 1.13 | [20 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [20 empty] | 2
|  |  | 25g | 25g | 01g | 0.5g | [20 empty] | 2
| 1 | 1 | 10 | 10 | 9 | 18 | [20 empty] | 1
| 1 | 2 | 7 | 9 | 10 | 14 | [20 empty] | 3
| 1 | 3 | 8 | 9 | 10 | 10 | [20 empty] | 16
| 1 | 4 | 7 | 9 | 10 | 15 | [20 empty] | 9.8
| 1 | 5 | 10 | 10 | 9 | 11 | [20 empty] | 9
| 1 | 6 | 7 | 9 | 9 | 11
| Mean |  | 8.1667 | 9.3333 | 9.5 | 13.1667
| Median |  | 7.5 | 9 | 9.5 | 12.5
| Std Dev. |  | 1.472 | 0.5164 | 0.54772 | 3.0605
| Rel.Std.Dev. |  | 18.02% | 5.533% | 5.766% | 23.24%
| PDM3 |  | -16.67% | -4.762% | -3.061% | 34.35%
| Z-Score (Absolute) |  | 1.26 | 0.08 | 0.08 | 3.79
| Analytical results for Rb in OREAS 236 (Certified Value 5.87 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.25 | 1.26 | 1.27 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 5.7 | 5.6 | 6.8 | 5.9 | 5.8 | 5.49 | 6.33 | 5.89 | 13.537 | 4.29 | 4.8 | 5.4691 | 4.7 | 6.3 | 5.96 | 5.7 | 6.2 | 5.5 | [6 empty] | 1
| 1 | 2 | 6.3 | 5.6 | 6.9 | 5.8 | 5.9 | 5.65 | 6.48 | 5.99 | 14.405 | 4.13 | 4.9 | 5.6017 | 4.7 | 6.3 | 5.91 | 5.7 | 6.3 | 5.7 | [6 empty] | 30
| 1 | 3 | 5.9 | 5.8 | 6.5 | 6 | 5.9 | 5.91 | 6.36 | 5.93 | 13.68 | 4.21 | 4.8 | 5.3877 | 4.8 | 6 | 5.79 | 5.7 | 6.1 | 5.7 | [6 empty] | 16
| 1 | 4 | 6.2 | 6 | 7.3 | 5.9 | 5.8 | 5.71 | 6.13 | 5.89 | 13.634 | 4.33 | 4.9 | 5.3766 | 4.8 | 6.1 | 5.93 | 5.6 | 6.2 | 5.6 | [6 empty] | 5.8727
| 1 | 5 | 6.1 | 5.8 | 7 | 6 | 5.7 | 5.73 | 6.31 | 5.72 | 14.326 | 4.05 | 4.8 | 5.5895 | 5 | 6.1 | 5.85 | 5.7 | 6.2 | 5.7 | [6 empty] | 114
| 1 | 6 | 6.2 | 6 | 6.6 | 5.8 | 5.8 | 5.51 | 5.99 | 5.7 | 13.06 | 4.26 | 4.7 | 5.6617 | 4.8 | 6.3 | 6.14 | 5.6 | 6.3 | 5.6
| Mean |  | 6.0667 | 5.8 | 6.85 | 5.9 | 5.8167 | 5.6667 | 6.2667 | 5.8533 | 13.7737 | 4.2117 | 4.8167 | 5.5144 | 4.8 | 6.1833 | 5.93 | 5.6667 | 6.2167 | 5.6333
| Median |  | 6.15 | 5.8 | 6.85 | 5.9 | 5.8 | 5.68 | 6.32 | 5.89 | 13.657 | 4.235 | 4.8 | 5.5293 | 4.8 | 6.2 | 5.92 | 5.7 | 6.2 | 5.65
| Std Dev. |  | 0.22509 | 0.17889 | 0.2881 | 0.089443 | 0.075277 | 0.15565 | 0.17626 | 0.11708 | 0.50937 | 0.10515 | 0.075277 | 0.12005 | 0.10954 | 0.13292 | 0.1195 | 0.05164 | 0.075277 | 0.08165
| Rel.Std.Dev. |  | 3.71% | 3.084% | 4.206% | 1.516% | 1.294% | 2.747% | 2.813% | 2% | 3.698% | 2.497% | 1.563% | 2.177% | 2.282% | 2.15% | 2.015% | 0.9113% | 1.211% | 1.449%
| PDM3 |  | 3.302% | -1.239% | 16.64% | 0.4643% | -0.9547% | -3.509% | 6.708% | -0.3304% | 134.5% | -28.28% | -17.98% | -6.102% | -18.27% | 5.289% | 0.9751% | -3.509% | 5.856% | -4.077%
| Z-Score (Absolute) |  | 0.57 | 0.09 | 2.48 | 0.16 | 0.04 | 0.41 | 1.05 | 0.04 | 19.38 | 3.96 | 2.49 | 0.78 | 2.53 | 0.85 | 0.23 | 0.41 | 0.93 | 0.49
| Analytical results for Re in OREAS 236 (Certified Value 0.002 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.18 | 1.19 | 1.23 | 1.25 | 1.26 | 1.27 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 01g | 0.15g | 0.5g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.002 | 0.003 | 0.007 | 0.002 | 0.002 | < 0.05 | 0.003 | 0.002 | 0.002 | < 0.05 | < 5 | 0.001 | 0.002 | 0.002 | [10 empty] | 1
| 1 | 2 | 0.002 | 0.002 | 0.006 | 0.002 | 0.003 | < 0.05 | 0.002 | 0.002 | 0.002 | < 0.05 | < 5 | 0.001 | 0.002 | 0.002 | [10 empty] | 31
| 1 | 3 | 0.002 | 0.002 | 0.006 | 0.003 | 0.002 | < 0.05 | 0.003 | 0.002 | 0.001 | < 0.05 | < 5 | 0.002 | 0.002 | 0.003 | [10 empty] | 16
| 1 | 4 | 0.002 | 0.002 | 0.007 | 0.002 | 0.002 | < 0.05 | 0.003 | 0.002 | 0.002 | < 0.05 | < 5 | 0.002 | 0.002 | 0.002 | [10 empty] | 0.0021333
| 1 | 5 | 0.002 | 0.002 | 0.008 | 0.002 | < 0.001 | < 0.05 | 0.002 | 0.002 | 0.002 | < 0.05 | < 5 | 0.002 | 0.002 | 0.002 | [10 empty] | 115
| 1 | 6 | 0.002 | 0.002 | 0.006 | 0.002 | 0.002 | < 0.05 | 0.003 | 0.002 | 0.002 | < 0.05 | < 5 | 0.001 | 0.002 | 0.002
| Mean |  | 0.002 | 0.0021667 | 0.0066667 | 0.0021667 | 0.0022 |  | 0.0026667 | 0.002 | 0.0018333 |  |  | 0.0015 | 0.002 | 0.0021667
| Median |  | 0.002 | 0.002 | 0.0065 | 0.002 | 0.002 |  | 0.003 | 0.002 | 0.002 |  |  | 0.0015 | 0.002 | 0.002
| Std Dev. |  | 0 | 0.00040825 | 0.0008165 | 0.00040825 | 0.00044721 |  | 0.0005164 | 0 | 0.00040825 |  |  | 0.00054772 | 0 | 0.00040825
| Rel.Std.Dev. |  | 0% | 18.84% | 12.25% | 18.84% | 20.33% |  | 19.36% | 0% | 22.27% |  |  | 36.51% | 0% | 18.84%
| PDM3 |  | -6.25% | 1.562% | 212.5% | 1.562% | 3.125% |  | 25% | -6.25% | -14.06% |  |  | -29.69% | -6.25% | 1.562%
| Z-Score (Absolute) |  | 0.39 | 0 | 10.4 | 0 | 0.58 | 52.79 | 1.16 | 0.39 | 0.77 | 52.79 | 5774.78 | 1.54 | 0.39 | 0
| Analytical results for S in OREAS 236 (Certified Value 0.445 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.46 | 0.465 | 0.4095 | 0.48 | 0.48 | 0.55 | 0.44 | 0.43 | 0.46 | 0.418 | 0.43479 | 0.48 | 0.17063 | 0.48 | 0.44 | 0.42 | 0.44724 | 0.44 | 0.41 | 0.46 | 0.45 | 0.42 |  |  | 1
| 1 | 2 | 0.47 | 0.475 | 0.41333 | 0.48 | 0.47 | 0.55 | 0.44 | 0.44 | 0.46 | 0.414 | 0.44825 | 0.47 | 0.17944 | 0.46 | 0.43 | 0.42 | 0.43949 | 0.45 | 0.4 | 0.47 | 0.47 | 0.42 |  |  | 16
| 1 | 3 | 0.47 | 0.475 | 0.40617 | 0.47 | 0.46 | 0.56 | 0.47 | 0.45 | 0.47 | 0.412 | 0.43413 | 0.46 | 0.18056 | 0.46 | 0.43 | 0.41 | 0.42444 | 0.44 | 0.4 | 0.46 | 0.46 | 0.42 |  |  | 16
| 1 | 4 | 0.47 | 0.48 | 0.4095 | 0.47 | 0.48 | 0.54 | 0.48 | 0.44 | 0.45 | 0.399 | 0.43536 | 0.47 | 0.17279 | 0.47 | 0.43 | 0.42 | 0.42961 | 0.45 | 0.4 | 0.46 | 0.46 | 0.41 |  |  | 0.44514
| 1 | 5 | 0.46 | 0.475 | 0.4095 | 0.49 | 0.48 | 0.55 | 0.44 | 0.43 | 0.45 | 0.408 | 0.44415 | 0.48 | 0.17868 | 0.46 | 0.44 | 0.42 | 0.42363 | 0.45 | 0.38 | 0.47 | 0.45 | 0.42 |  |  | 116
| 1 | 6 | 0.47 | 0.455 | 0.4095 | 0.48 | 0.47 | 0.56 | 0.44 | 0.43 | 0.45 | 0.404 | 0.44158 | 0.48 | 0.17199 | 0.47 | 0.44 | 0.42 | 0.43205 | 0.45 | 0.4 | 0.46 | 0.47 | 0.42
| Mean |  | 0.46667 | 0.47083 | 0.40958 | 0.47833 | 0.47333 | 0.55167 | 0.45167 | 0.43667 | 0.45667 | 0.40917 | 0.43971 | 0.47333 | 0.17568 | 0.46667 | 0.435 | 0.41833 | 0.43274 | 0.44667 | 0.39833 | 0.46333 | 0.46 | 0.41833
| Median |  | 0.47 | 0.475 | 0.4095 | 0.48 | 0.475 | 0.55 | 0.44 | 0.435 | 0.455 | 0.41 | 0.43847 | 0.475 | 0.17574 | 0.465 | 0.435 | 0.42 | 0.43083 | 0.45 | 0.4 | 0.46 | 0.46 | 0.42
| Std Dev. |  | 0.005164 | 0.0091742 | 0.00227 | 0.0075277 | 0.008165 | 0.0075277 | 0.018348 | 0.008165 | 0.008165 | 0.0069402 | 0.0058352 | 0.008165 | 0.0043456 | 0.008165 | 0.0054772 | 0.0040825 | 0.0091437 | 0.005164 | 0.0098319 | 0.005164 | 0.0089443 | 0.0040825
| Rel.Std.Dev. |  | 1.107% | 1.949% | 0.5542% | 1.574% | 1.725% | 1.365% | 4.062% | 1.87% | 1.788% | 1.696% | 1.327% | 1.725% | 2.474% | 1.75% | 1.259% | 0.9759% | 2.113% | 1.156% | 2.468% | 1.115% | 1.944% | 0.9759%
| PDM3 |  | 4.835% | 5.771% | -7.988% | 7.456% | 6.333% | 23.93% | 1.465% | -1.904% | 2.589% | -8.082% | -1.22% | 6.333% | -60.53% | 4.835% | -2.279% | -6.023% | -2.786% | 0.3422% | -10.52% | 4.086% | 3.337% | -6.023%
| Z-Score (Absolute) |  | 0.6 | 0.75 | 1.36 | 1 | 0.83 | 3.53 | 0.09 | 0.43 | 0.26 | 1.38 | 0.33 | 0.83 | 9.42 | 0.6 | 0.49 | 1.06 | 0.57 | 0.09 | 1.75 | 0.49 | 0.37 | 1.06
| Analytical results for Sb in OREAS 236 (Certified Value 0.9 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 1.25 | 0.7 | < 5 | 0.54 | 0.69 | 0.99 | 1.36 | 0.96 | 1.19 | < 10 | 0.52 | 0.49 | 0.60888 | 0.61 | < 10 | 0.96 | 0.94 | 1.27 | 6.4271 | < 5 | < 5 | 1.26 | 0.57 | 1.21 | 1
| 1 | 2 | 1.37 | 0.66 | < 5 | 0.52 | 0.7 | 0.97 | 1.41 | 0.97 | 1.2 | < 10 | 0.55 | 0.5 | 0.64818 | 0.63 | < 10 | 0.95 | 0.95 | 1.32 | 8.6524 | < 5 | < 5 | 1.31 | 0.61 | 1.21 | 32
| 1 | 3 | 1.38 | 0.6 | < 5 | 0.51 | 0.69 | 0.96 | 1.41 | 0.98 | 1.2 | < 10 | 0.51 | 0.49 | 0.58941 | 0.63 | < 10 | 0.95 | 0.93 | 1.33 | 7.3115 | < 5 | < 5 | 1.28 | 0.62 | 1.11 | 16
| 1 | 4 | 1.18 | 0.6 | < 5 | 0.56 | 0.7 | 0.99 | 1.42 | 1.01 | 1.14 | < 10 | 0.49 | 0.49 | 0.56347 | 0.62 | < 10 | 1.04 | 0.98 | 1.32 | 8.4801 | < 5 | < 5 | 1.27 | 0.63 | 1.15 | 0.89803
| 1 | 5 | 1.18 | 0.58 | < 5 | 0.54 | 0.68 | 1 | 1.4 | 0.97 | 1.16 | < 10 | 0.47 | 0.48 | 0.5454 | 0.66 | < 10 | 0.99 | 0.91 | 1.27 | 7.7154 | < 5 | < 5 | 1.37 | 0.61 | 1.1 | 117
| 1 | 6 | 1.2 | 0.72 | < 5 | 0.54 | 0.65 | 0.96 | 1.44 | 0.92 | 1.14 | < 10 | 0.51 | 0.51 | 0.51947 | 0.62 | < 10 | 1.05 | 0.97 | 1.22 | 6.3666 | < 5 | < 5 | 1.38 | 0.59 | 1.17
| Mean |  | 1.26 | 0.64333 |  | 0.535 | 0.685 | 0.97833 | 1.4067 | 0.96833 | 1.1717 |  | 0.50833 | 0.49333 | 0.57914 | 0.62833 |  | 0.99 | 0.94667 | 1.2883 | 7.4922 |  |  | 1.3117 | 0.605 | 1.1583
| Median |  | 1.225 | 0.63 |  | 0.54 | 0.69 | 0.98 | 1.41 | 0.97 | 1.175 |  | 0.51 | 0.49 | 0.57644 | 0.625 |  | 0.975 | 0.945 | 1.295 | 7.5134 |  |  | 1.295 | 0.61 | 1.16
| Std Dev. |  | 0.092736 | 0.058538 |  | 0.017607 | 0.018708 | 0.017224 | 0.026583 | 0.029269 | 0.028577 |  | 0.027142 | 0.010328 | 0.046263 | 0.017224 |  | 0.045166 | 0.02582 | 0.042622 | 0.98042 |  |  | 0.051929 | 0.021679 | 0.047504
| Rel.Std.Dev. |  | 7.36% | 9.099% |  | 3.291% | 2.731% | 1.761% | 1.89% | 3.023% | 2.439% |  | 5.339% | 2.094% | 7.988% | 2.741% |  | 4.562% | 2.727% | 3.308% | 13.09% |  |  | 3.959% | 3.583% | 4.101%
| PDM3 |  | 40.31% | -28.36% |  | -40.42% | -23.72% | 8.943% | 56.64% | 7.829% | 30.47% |  | -43.39% | -45.06% | -35.51% | -30.03% |  | 10.24% | 5.416% | 43.46% | 734.3% |  |  | 46.06% | -32.63% | 28.99%
| Z-Score (Absolute) |  | 0.29 | 0.66 | 2.19 | 0.83 | 0.6 | 0.15 | 0.51 | 0.16 | 0.15 | 6.04 | 0.87 | 0.89 | 0.76 | 0.69 | 6.04 | 0.13 | 0.2 | 0.33 | 9.87 | 2.19 | 2.19 | 0.37 | 0.72 | 0.13
| Analytical results for Sc in OREAS 236 (Certified Value 6.47 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 5.7 | 5 | 5.985 | 5.6 | 5.1 | 6.7 | 5.7 | 6.6 | 7 | 6.8 | 7.9 | 7.3577 | 6.5 | 7.6 | 7 | 6.3 | 6 | 6 | 7.2 | 6.6 | 6.3 |  |  |  | 1
| 1 | 2 | 6.1 | 5 | 6.051 | 5.5 | 5.1 | 6.8 | 6 | 6.9 | 7 | 6.5 | 8.4 | 7.4539 | 6.7 | 7.9 | 7 | 6.4 | 6 | 6 | 7.1 | 6.6 | 6.5 |  |  |  | 33
| 1 | 3 | 5.9 | 5 | 5.98 | 5.6 | 5 | 6.8 | 5.9 | 7.2 | 7 | 6.5 | 8.2 | 7.2346 | 6.8 | 7.2 | 7 | 6.6 | 6 | 6 | 7.1 | 6.3 | 6.2 |  |  |  | 16
| 1 | 4 | 6.8 | 5 | 6.084 | 5.4 | 5.2 | 6.9 | 6 | 7.2 | 7 | 6.7 | 8.6 | 6.8945 | 6.3 | 7.6 | 7 | 6.6 | 6 | 6 | 6.9 | 6.3 | 6.3 |  |  |  | 6.472
| 1 | 5 | 6.8 | 5 | 6.1115 | 5.5 | 5 | 6.7 | 6 | 6.7 | 7 | 6.7 | 8.2 | 7.1349 | 7 | 7.5 | 7 | 6.7 | 6 | 6 | 6.9 | 6.4 | 6.2 |  |  |  | 118
| 1 | 6 | 7 | 4 | 6.0767 | 5.6 | 5 | 6.9 | 5.7 | 5.5 | 7 | 6.6 | 8 | 7.2986 | 6.7 | 7.7 | 7 | 6.4 | 6 | 6 | 7.1 | 6.7 | 6.3
| Mean |  | 6.3833 | 4.8333 | 6.048 | 5.5333 | 5.0667 | 6.8 | 5.8833 | 6.6833 | 7 | 6.6333 | 8.2167 | 7.2291 | 6.6667 | 7.5833 | 7 | 6.5 | 6 | 6 | 7.05 | 6.4833 | 6.3
| Median |  | 6.45 | 5 | 6.0639 | 5.55 | 5.05 | 6.8 | 5.95 | 6.8 | 7 | 6.65 | 8.2 | 7.2666 | 6.7 | 7.6 | 7 | 6.5 | 6 | 6 | 7.1 | 6.5 | 6.3
| Std Dev. |  | 0.54924 | 0.40825 | 0.054326 | 0.08165 | 0.08165 | 0.089443 | 0.1472 | 0.63061 | 0 | 0.12111 | 0.25626 | 0.19638 | 0.24221 | 0.23166 | 0 | 0.15492 | 0 | 0 | 0.12247 | 0.17224 | 0.10954
| Rel.Std.Dev. |  | 8.604% | 8.447% | 0.8982% | 1.476% | 1.612% | 1.315% | 2.502% | 9.436% | 0% | 1.826% | 3.119% | 2.717% | 3.633% | 3.055% | 0% | 2.383% | 0% | 0% | 1.737% | 2.657% | 1.739%
| PDM3 |  | -1.37% | -25.32% | -6.551% | -14.5% | -21.71% | 5.068% | -9.096% | 3.265% | 8.158% | 2.492% | 26.96% | 11.7% | 3.007% | 17.17% | 8.158% | 0.4322% | -7.293% | -7.293% | 8.93% | 0.1747% | -2.658%
| Z-Score (Absolute) |  | 0.33 | NA | 0.93 | 1.86 | 2.7 | 0.42 | 1.23 | 0.21 | NA | 0.12 | 2.97 | 1.19 | 0.18 | 1.83 | NA | 0.12 | NA | NA | 0.87 | 0.15 | 0.48
Results from laboratories 2, 11, 19, 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 236 (Certified Value < 1 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 0.7 | 1 | < 5 | < 1 | 1.3 | < 1 | 0.8 | < 1 | < 10 | < 1 | 0.3 | < 10 | 0.5 | < 1 | 0.3 | < 10 | < 5 | < 10 | 0.7 | 0.5 | 0.7 |  |  |  | 1
| 1 | 2 | 0.7 | 1 | < 5 | < 1 | 1.2 | < 1 | 0.6 | < 1 | < 10 | < 1 | 0.2 | < 10 | 0.6 | < 1 | 0.3 | < 10 | < 5 | < 10 | 0.6 | 0.5 | 0.9 |  |  |  | 34
| 1 | 3 | 0.8 | < 1 | < 5 | < 1 | 1.1 | < 1 | 0.6 | < 1 | < 10 | < 1 | 0.2 | < 10 | 0.6 | < 1 | 0.4 | < 10 | < 5 | < 10 | 0.8 | 0.5 | 0.5 |  |  |  | 16
| 1 | 4 | 0.8 | 1 | < 5 | < 1 | 1.1 | < 1 | < 0.5 | < 1 | < 10 | < 1 | 0.2 | < 10 | 0.5 | < 1 | 0.3 | < 10 | < 5 | < 10 | 0.9 | 0.5 | 0.5 |  |  |  | < 1
| 1 | 5 | 0.8 | 1 | < 5 | < 1 | 1 | < 1 | < 0.5 | < 1 | < 10 | < 1 | 0.1 | < 10 | 0.8 | < 1 | 0.2 | < 10 | < 5 | < 10 | 0.8 | 0.5 | 0.7 |  |  |  | 119
| 1 | 6 | 0.8 | < 1 | < 5 | < 1 | 1.2 | < 1 | 1.2 | < 1 | < 10 | < 1 | 0.3 | < 10 | 0.5 | < 1 | 0.3 | < 10 | < 5 | < 10 | 0.7 | 0.5 | 0.8
| Mean |  | 0.76667 | 1 |  |  | 1.15 |  | 0.8 |  |  |  | 0.21667 |  | 0.58333 |  | 0.3 |  |  |  | 0.75 | 0.5 | 0.68333
| Median |  | 0.8 | 1 |  |  | 1.15 |  | 0.7 |  |  |  | 0.2 |  | 0.55 |  | 0.3 |  |  |  | 0.75 | 0.5 | 0.7
| Std Dev. |  | 0.05164 | 0 |  |  | 0.10488 |  | 0.28284 |  |  |  | 0.075277 |  | 0.1169 |  | 0.063246 |  |  |  | 0.10488 | 0 | 0.16021
| Rel.Std.Dev. |  | 6.736% | 0% |  |  | 9.12% |  | 35.36% |  |  |  | 34.74% |  | 20.04% |  | 21.08% |  |  |  | 13.98% | 0% | 23.45%
PDM3
| Z-Score (Absolute) |  | 0.31 | 0.55 | 6.68 | 0.67 | 1.72 | 0.67 | 0.25 | 0.67 | 15.88 | 0.67 | 1.72 | 15.88 | 0.37 | 0.67 | 1.41 | 15.88 | 6.68 | 15.88 | 0.25 | 0.67 | 0
| Analytical results for Sm in OREAS 236 (Certified Value 1.89 ppm) | [25 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 3
| 1 | 1 | 1.58 | 1.73 | 2.18 | 1.7 | 1.9297 | [19 empty] | 1
| 1 | 2 | 1.66 | 1.815 | 2.1 | 2 | 1.9185 | [19 empty] | 18
| 1 | 3 | 1.65 | 1.815 | 2.22 | 2.1 | 1.9418 | [19 empty] | 16
| 1 | 4 | 1.7 | 1.772 | 2.19 | 2.1 | 1.8586 | [19 empty] | 1.8878
| 1 | 5 | 1.66 | 1.83 | 2.07 | 1.8 | 1.922 | [19 empty] | 120
| 1 | 6 | 1.65 | 1.771 | 1.92 | 2.1 | 1.9499
| Mean |  | 1.65 | 1.7888 | 2.1133 | 1.9667 | 1.9201
| Median |  | 1.655 | 1.7935 | 2.14 | 2.05 | 1.9259
| Std Dev. |  | 0.038987 | 0.037754 | 0.11057 | 0.17512 | 0.032364
| Rel.Std.Dev. |  | 2.363% | 2.111% | 5.232% | 8.904% | 1.686%
| PDM3 |  | -12.6% | -5.242% | 11.95% | 4.179% | 1.711%
| Z-Score (Absolute) |  | 1.39 | 0.67 | 0.99 | 0.24 | 0
| Analytical results for Sn in OREAS 236 (Certified Value 0.68 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.6 | 0.4 | < 10 | 0.8 | 0.4 | 0.7 | 0.59 | 0.7 | 0.7 | < 10 | 0.65 | 0.74 | 0.7 | < 5 | 0.7 | 0.9 | < 5 | < 10 | < 20 | 0.7 | 0.7 | 0.6 |  |  | 1
| 1 | 2 | 0.7 | 0.4 | < 10 | 0.7 | 0.4 | 0.7 | 0.61 | 0.7 | 0.7 | < 10 | 0.62 | 0.75 | 0.7 | < 5 | 0.8 | 0.8 | < 5 | < 10 | < 20 | 0.7 | 0.7 | 0.7 |  |  | 7
| 1 | 3 | 0.7 | 0.4 | < 10 | 0.7 | 0.4 | 0.7 | 0.64 | 0.7 | 0.7 | < 10 | 0.64 | 0.73 | 0.7 | < 5 | 0.7 | 0.8 | < 5 | < 10 | < 20 | 0.7 | 0.7 | 0.6 |  |  | 16
| 1 | 4 | 0.6 | 0.5 | < 10 | 0.7 | 0.4 | 0.7 | 0.65 | 0.7 | 0.7 | < 10 | 0.65 | 0.77 | 0.7 | < 5 | 0.7 | 0.9 | < 5 | < 10 | < 20 | 0.7 | 0.7 | 0.6 |  |  | 0.68423
| 1 | 5 | 0.6 | 0.4 | < 10 | 0.7 | 0.4 | 0.7 | 0.64 | 0.7 | 0.7 | < 10 | 0.64 | 0.74 | 0.7 | < 5 | 0.7 | 0.8 | < 5 | < 10 | < 20 | 0.7 | 0.7 | 0.6 |  |  | 121
| 1 | 6 | 0.7 | 0.4 | < 10 | 0.7 | 0.4 | 0.7 | 0.56 | 0.6 | 0.7 | < 10 | 0.62 | 0.73 | 0.7 | < 5 | 0.7 | 0.8 | < 5 | < 10 | < 20 | 0.7 | 0.8 | 0.6
| Mean |  | 0.65 | 0.41667 |  | 0.71667 | 0.4 | 0.7 | 0.615 | 0.68333 | 0.7 |  | 0.63667 | 0.74333 | 0.7 |  | 0.71667 | 0.83333 |  |  |  | 0.7 | 0.71667 | 0.61667
| Median |  | 0.65 | 0.4 |  | 0.7 | 0.4 | 0.7 | 0.625 | 0.7 | 0.7 |  | 0.64 | 0.74 | 0.7 |  | 0.7 | 0.8 |  |  |  | 0.7 | 0.7 | 0.6
| Std Dev. |  | 0.054772 | 0.040825 |  | 0.040825 | 6.0809e-17 | 1.2162e-16 | 0.035071 | 0.040825 | 1.2162e-16 |  | 0.013663 | 0.015055 | 1.2162e-16 |  | 0.040825 | 0.05164 |  |  |  | 1.2162e-16 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 8.427% | 9.798% |  | 5.696% | 1.52e-14% | 1.737e-14% | 5.703% | 5.974% | 1.737e-14% |  | 2.146% | 2.025% | 1.737e-14% |  | 5.696% | 6.197% |  |  |  | 1.737e-14% | 5.696% | 6.62%
| PDM3 |  | -5.003% | -39.1% |  | 4.74% | -41.54% | 2.305% | -10.12% | -0.1312% | 2.305% |  | -6.951% | 8.638% | 2.305% |  | 4.74% | 21.79% |  |  |  | 2.305% | 4.74% | -9.874%
| Z-Score (Absolute) |  | 0.48 | 2.41 | 35.44 | 0.07 | 2.55 | 0.07 | 0.77 | 0.21 | 0.07 | 35.44 | 0.59 | 0.29 | 0.07 | 14.79 | 0.07 | 1.03 | 14.79 | 35.44 | 76.72 | 0.07 | 0.07 | 0.76
| Analytical results for Sr in OREAS 236 (Certified Value 29.7 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | 2
| 1 | 1 | 27.9 | 28 | 35.1563 | 29.3 | 26.3 | 26 | 26.61 | 32.8 | 30.6 | 25.15 | 33.8 | 34.9 | 28.4155 | 29.6 | 26.489 | 32.4 | 38.3 | 31.5 | 34.3716 | 27 | 30 | 29.7 | 30.7 | 28.4 | 1
| 1 | 2 | 30.1 | 28.4 | 35.0413 | 29 | 25.9 | 27 | 27.28 | 33.6 | 31.5 | 25.9 | 32.55 | 34.7 | 29.0547 | 29.4 | 28.0638 | 33 | 39 | 32.4 | 32.6162 | 27 | 29 | 29.1 | 30.5 | 28.7 | 8
| 1 | 3 | 28.3 | 28.6 | 34.7023 | 28.6 | 25.5 | 27 | 27.67 | 34.3 | 31 | 25.63 | 32.22 | 35.3 | 28.8604 | 29.3 | 28.3865 | 30 | 39.5 | 31.2 | 31.0478 | 26 | 29 | 29.5 | 29.9 | 27.8 | 16
| 1 | 4 | 29.2 | 29.2 | 35.0283 | 30.1 | 26 | 26 | 28.05 | 33.5 | 31.4 | 25.26 | 33.52 | 34.9 | 28.2895 | 28.7 | 26.7367 | 31.2 | 38.3 | 31.2 | 31.6356 | 27 | 29 | 29.1 | 30 | 28.6 | 29.7497
| 1 | 5 | 29.1 | 29 | 35.6463 | 29.7 | 25.7 | 26 | 27.73 | 32.4 | 30.3 | 25.74 | 32.5 | 36.1 | 29.4398 | 29.4 | 27.3456 | 31.2 | 38.5 | 30.9 | 31.9234 | 28 | 28 | 29.1 | 30.1 | 28.2 | 122
| 1 | 6 | 30.1 | 28.8 | 35.2678 | 29.6 | 25.6 | 27 | 26.52 | 28.2 | 30.3 | 25.49 | 33.14 | 32.6 | 29.6868 | 29.1 | 27.098 | 32.3 | 39.4 | 29.8 | 31.8541 | 28 | 29 | 29.2 | 31.2 | 28.6
| Mean |  | 29.1167 | 28.6667 | 35.1404 | 29.3833 | 25.8333 | 26.5 | 27.31 | 32.4667 | 30.85 | 25.5283 | 32.955 | 34.75 | 28.9578 | 29.25 | 27.3533 | 31.6833 | 38.8333 | 31.1667 | 32.2414 | 27.1667 | 29 | 29.2833 | 30.4 | 28.3833
| Median |  | 29.15 | 28.7 | 35.0988 | 29.45 | 25.8 | 26.5 | 27.475 | 33.15 | 30.8 | 25.56 | 32.845 | 34.9 | 28.9576 | 29.35 | 27.2218 | 31.75 | 38.75 | 31.2 | 31.8887 | 27 | 29 | 29.15 | 30.3 | 28.5
| Std Dev. |  | 0.90425 | 0.43205 | 0.312 | 0.53448 | 0.29439 | 0.54772 | 0.62747 | 2.1924 | 0.53198 | 0.28632 | 0.62909 | 1.1658 | 0.55199 | 0.31464 | 0.74361 | 1.0889 | 0.54283 | 0.84538 | 1.1594 | 0.75277 | 0.63246 | 0.25626 | 0.498 | 0.33714
| Rel.Std.Dev. |  | 3.106% | 1.507% | 0.8879% | 1.819% | 1.14% | 2.067% | 2.298% | 6.753% | 1.724% | 1.122% | 1.909% | 3.355% | 1.906% | 1.076% | 2.719% | 3.437% | 1.398% | 2.712% | 3.596% | 2.771% | 2.181% | 0.8751% | 1.638% | 1.188%
| PDM3 |  | -2.128% | -3.641% | 18.12% | -1.232% | -13.16% | -10.92% | -8.201% | 9.133% | 3.698% | -14.19% | 10.77% | 16.81% | -2.662% | -1.68% | -8.055% | 6.5% | 30.53% | 4.763% | 8.376% | -8.683% | -2.52% | -1.568% | 2.186% | -4.593%
| Z-Score (Absolute) |  | 0.05 | 0.21 | 2.05 | 0.04 | 1.2 | 0.96 | 0.68 | 1.12 | 0.55 | 1.3 | 1.29 | 1.91 | 0.11 | 0.01 | 0.67 | 0.84 | 3.33 | 0.66 | 1.04 | 0.73 | 0.09 | 0.01 | 0.39 | 0.31
| Analytical results for Ta in OREAS 236 (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.18 | 1.19 | 1.24 | 1.25 | 1.26 | 1.27 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.005 | 5.88 | 0.05 | 0.16 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 10 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.005 | 5.9167 | < 0.05 | 0.16 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | < 0.05 | < 10 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 21
| 1 | 3 | < 0.01 | < 0.005 | 5.72 | 0.08 | 0.15 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.005 | 6.005 | < 0.05 | 0.16 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 | [7 empty] | < 0.01
| 1 | 5 | 0.01 | < 0.005 | 5.845 | 0.08 | 0.15 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 10 | < 0.01 | < 0.01 | < 0.01 | [7 empty] | 123
| 1 | 6 | 0.01 | < 0.005 | 5.97 | 0.09 | 0.15 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 10 | < 0.01 | < 0.01 | < 0.01
| Mean |  | 0.01 |  | 5.8894 | 0.075 | 0.155 | [6 empty] | 0.01 |  |  | 0.01
| Median |  | 0.01 |  | 5.8983 | 0.08 | 0.155 | [6 empty] | 0.01 |  |  | 0.01
| Std Dev. |  | 0 |  | 0.10135 | 0.017321 | 0.0054772 | [6 empty] | 0 |  |  | 0
| Rel.Std.Dev. |  | 0% |  | 1.721% | 23.09% | 3.534% | [6 empty] | 0% |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 952.03 | 8.36 | 24.01 | 7.01 | 0.27 | 0.27 | 2.97 | 0.27 | 2.97 | 0 | 2.97 | 808.09 | 0 | 0.27 | 0.27
| Analytical results for Tb in OREAS 236 (Certified Value 0.39 ppm) | [25 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 40g | 0.25g | 0.25g | 25g | 25g | 0.5g | 0.5g | [17 empty] | 3
| 1 | 1 | 0.355 | 0.39 | 0.38 | 0.358 | 0.41 | 0.4 | 0.40441 | [17 empty] | 1
| 1 | 2 | 0.37 | 0.39 | 0.38 | 0.367 | 0.41 | 0.4 | 0.41788 | [17 empty] | 22
| 1 | 3 | 0.365 | 0.38 | 0.38 | 0.371 | 0.43 | 0.4 | 0.41535 | [17 empty] | 16
| 1 | 4 | 0.38 | 0.38 | 0.39 | 0.369 | 0.42 | 0.4 | 0.38884 | [17 empty] | 0.38707
| 1 | 5 | 0.375 | 0.39 | 0.4 | 0.372 | 0.39 | 0.4 | 0.41247 | [17 empty] | 124
| 1 | 6 | 0.365 | 0.4 | 0.37 | 0.356 | 0.36 | 0.4 | 0.41784
| Mean |  | 0.36833 | 0.38833 | 0.38333 | 0.3655 | 0.40333 | 0.4 | 0.40946
| Median |  | 0.3675 | 0.39 | 0.38 | 0.368 | 0.41 | 0.4 | 0.41391
| Std Dev. |  | 0.008756 | 0.0075277 | 0.010328 | 0.0068337 | 0.025033 | 6.0809e-17 | 0.011275
| Rel.Std.Dev. |  | 2.377% | 1.938% | 2.694% | 1.87% | 6.207% | 1.52e-14% | 2.754%
| PDM3 |  | -4.841% | 0.3263% | -0.9655% | -5.573% | 4.202% | 3.34% | 5.785%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 0.1 | 0.78 | 0.67 | NA | 0.91
Results from laboratory 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 236 (Certified Value 0.11 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.1 | 0.08 | 7.9833 | 0.1 | < 0.1 | 0.1 | 0.06 | 0.12 | < 20 | 0.1 | 0.07 | 0.1 | 0.11 | 0.1 | < 5 | < 10 | 0.13 | 0.1 | 0.1 | [5 empty] | 1
| 1 | 2 | 0.11 | 0.12 | 7.8489 | 0.09 | 0.1 | 0.1 | 0.29 | 0.1 | < 20 | 0.1 | 0.07 | 0.11 | 0.11 | 0.1 | < 5 | < 10 | 0.1 | 0.09 | 0.11 | [5 empty] | 23
| 1 | 3 | 0.09 | 0.12 | 7.99 | 0.1 | 0.1 | 0.1 | 0.12 | 0.11 | < 20 | 0.1 | 0.07 | 0.09 | 0.11 | 0.11 | < 5 | < 10 | 0.12 | 0.09 | 0.1 | [5 empty] | 16
| 1 | 4 | 0.09 | 0.14 | 7.9233 | 0.12 | 0.1 | 0.1 | 0.13 | 0.11 | < 20 | 0.1 | 0.07 | 0.1 | 0.1 | 0.1 | < 5 | < 10 | 0.12 | 0.09 | 0.1 | [5 empty] | 0.10545
| 1 | 5 | 0.1 | 0.08 | 7.6533 | 0.12 | 0.1 | 0.1 | 0.12 | 0.09 | < 20 | < 0.1 | 0.07 | 0.1 | 0.11 | 0.11 | < 5 | < 10 | 0.12 | 0.09 | 0.1 | [5 empty] | 125
| 1 | 6 | 0.11 | 0.14 | 7.9767 | 0.11 | < 0.1 | 0.1 | 0.09 | 0.1 | < 20 | 0.1 | 0.07 | 0.11 | 0.11 | 0.08 | < 5 | < 10 | 0.11 | 0.09 | 0.1
| Mean |  | 0.1 | 0.11333 | 7.8959 | 0.10667 | 0.1 | 0.1 | 0.135 | 0.105 |  | 0.1 | 0.07 | 0.10167 | 0.10833 | 0.1 |  |  | 0.11667 | 0.091667 | 0.10167
| Median |  | 0.1 | 0.12 | 7.95 | 0.105 | 0.1 | 0.1 | 0.12 | 0.105 |  | 0.1 | 0.07 | 0.1 | 0.11 | 0.1 |  |  | 0.12 | 0.09 | 0.1
| Std Dev. |  | 0.0089443 | 0.027325 | 0.13025 | 0.012111 | 0 | 1.5202e-17 | 0.080187 | 0.010488 |  | 0 | 0 | 0.0075277 | 0.0040825 | 0.010954 |  |  | 0.010328 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 8.944% | 24.11% | 1.65% | 11.35% | 0% | 1.52e-14% | 59.4% | 9.989% |  | 0% | 0% | 7.404% | 3.768% | 10.95% |  |  | 8.853% | 4.454% | 4.016%
| PDM3 |  | -5.172% | 7.471% | 7388% | 1.149% | -5.172% | -5.172% | 28.02% | -0.431% |  | -5.172% | -33.62% | -3.592% | 2.73% | -5.172% |  |  | 10.63% | -13.07% | -3.592%
| Z-Score (Absolute) |  | 0.61 | 0.47 | 630.22 | 0.07 | NA | NA | 2.23 | 0.2 | 800.47 | NA | 3.03 | 0.47 | 0.07 | 0.61 | 193.59 | 395.89 | 0.74 | 1.28 | 0.47
Results from laboratories 7, 8 and 13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 236 (Certified Value 0.88 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.8 | 0.86 | 1 | 1 | 0.9 | 0.94 | 0.9 | 0.9 | 0.83 | 0.9 | 0.7517 | 1 | 0.9 | 0.91 | 1 | < 10 | 0.8 | 0.9 | 0.8 | [5 empty] | 1
| 1 | 2 | 0.9 | 0.88 | 0.9 | 0.9 | 0.9 | 0.93 | 0.8 | 0.92 | 0.8 | 0.9 | 0.81751 | 0.9 | 0.9 | 0.93 | 1.1 | < 10 | 0.8 | 0.9 | 0.8 | [5 empty] | 24
| 1 | 3 | 0.9 | 0.88 | 0.9 | 0.9 | 0.9 | 0.95 | 0.9 | 0.91 | 0.82 | 0.9 | 0.78806 | 0.9 | 0.8 | 0.93 | 1 | < 10 | 0.8 | 0.9 | 0.8 | [5 empty] | 16
| 1 | 4 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.91 | 0.9 | 0.91 | 0.83 | 0.9 | 0.76085 | 1 | 0.8 | 0.95 | 1.2 | < 10 | 0.8 | 0.9 | 0.8 | [5 empty] | 0.87875
| 1 | 5 | 0.9 | 0.9 | 0.9 | 1 | 0.9 | 0.94 | 0.8 | 0.89 | 0.8 | 0.9 | 0.81829 | 0.9 | 0.9 | 0.92 | 1.1 | < 10 | 0.8 | 0.9 | 0.8 | [5 empty] | 126
| 1 | 6 | 0.9 | 0.92 | 0.9 | 0.9 | 0.9 | 0.91 | 0.8 | 0.88 | 0.84 | 0.9 | 0.7564 | 0.9 | 0.9 | 0.95 | 1.4 | < 10 | 0.8 | 0.9 | 0.8
| Mean |  | 0.88333 | 0.89 | 0.91667 | 0.93333 | 0.9 | 0.93 | 0.85 | 0.90167 | 0.82 | 0.9 | 0.78214 | 0.93333 | 0.86667 | 0.93167 | 1.1333 |  | 0.8 | 0.9 | 0.8
| Median |  | 0.9 | 0.89 | 0.9 | 0.9 | 0.9 | 0.935 | 0.85 | 0.905 | 0.825 | 0.9 | 0.77446 | 0.9 | 0.9 | 0.93 | 1.1 |  | 0.8 | 0.9 | 0.8
| Std Dev. |  | 0.040825 | 0.020976 | 0.040825 | 0.05164 | 0 | 0.016733 | 0.054772 | 0.01472 | 0.016733 | 0 | 0.030448 | 0.05164 | 0.05164 | 0.016021 | 0.15055 |  | 1.2162e-16 | 0 | 1.2162e-16
| Rel.Std.Dev. |  | 4.622% | 2.357% | 4.454% | 5.533% | 0% | 1.799% | 6.444% | 1.632% | 2.041% | 0% | 3.893% | 5.533% | 5.958% | 1.72% | 13.28% |  | 1.52e-14% | 0% | 1.52e-14%
| PDM3 |  | 0.5212% | 1.28% | 4.314% | 6.211% | 2.418% | 5.832% | -3.272% | 2.608% | -6.686% | 2.418% | -10.99% | 6.211% | -1.375% | 6.021% | 28.97% |  | -8.962% | 2.418% | -8.962%
| Z-Score (Absolute) |  | 0.34 | 0.2 | 0.34 | 0.67 | 0 | 0.61 | 1.01 | 0.03 | 1.62 | 0 | 2.38 | 0.67 | 0.67 | 0.64 | 4.72 | 82.94 | 2.02 | 0 | 2.02
| Analytical results for Ti in OREAS 236 (Certified Value 0.387 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.359 | 0.15 | 0.332 | 0.28 | 0.39 | 0.3083 | 0.36 | 0.4624 | 0.4984 | 0.45 | 0.46805 | 0.44 | 0.39686 | 0.419 | 0.441 | 0.345 | 0.46057 | 0.37 | 0.3 | 0.398 | 0.384 | 0.303 |  |  | 1
| 1 | 2 | 0.348 | 0.16 | 0.33667 | 0.28 | 0.396 | 0.3188 | 0.39 | 0.4599 | 0.4878 | 0.45 | 0.45635 | 0.41 | 0.43397 | 0.417 | 0.442 | 0.363 | 0.47016 | 0.35 | 0.28 | 0.397 | 0.389 | 0.331 |  |  | 25
| 1 | 3 | 0.35 | 0.17 | 0.33506 | 0.28 | 0.394 | 0.3149 | 0.39 | 0.4509 | 0.4868 | 0.44 | 0.45822 | 0.4 | 0.44026 | 0.394 | 0.445 | 0.346 | 0.45899 | 0.34 | 0.29 | 0.391 | 0.36 | 0.293 |  |  | 16
| 1 | 4 | 0.352 | 0.19 | 0.33067 | 0.28 | 0.387 | 0.3341 | 0.39 | 0.4607 | 0.4975 | 0.44 | 0.45567 | 0.41 | 0.42354 | 0.407 | 0.442 | 0.376 | 0.46442 | 0.35 | 0.28 | 0.378 | 0.355 | 0.307 |  |  | 0.38698
| 1 | 5 | 0.352 | 0.185 | 0.337 | 0.28 | 0.372 | 0.332 | 0.37 | 0.4427 | 0.4904 | 0.44 | 0.45694 | 0.41 | 0.42983 | 0.404 | 0.446 | 0.377 | 0.4521 | 0.38 | 0.28 | 0.388 | 0.368 | 0.291 |  |  | 127
| 1 | 6 | 0.369 | 0.13 | 0.33668 | 0.28 | 0.399 | 0.3042 | 0.26 | 0.4412 | 0.4893 | 0.43 | 0.46465 | 0.41 | 0.41406 | 0.411 | 0.449 | 0.343 | 0.47775 | 0.38 | 0.27 | 0.39 | 0.404 | 0.315
| Mean |  | 0.355 | 0.16417 | 0.33468 | 0.28 | 0.38967 | 0.31872 | 0.36 | 0.45297 | 0.4917 | 0.44167 | 0.45998 | 0.41333 | 0.42309 | 0.40867 | 0.44417 | 0.35833 | 0.464 | 0.36167 | 0.28333 | 0.39033 | 0.37667 | 0.30667
| Median |  | 0.352 | 0.165 | 0.33586 | 0.28 | 0.392 | 0.31685 | 0.38 | 0.4554 | 0.48985 | 0.44 | 0.45758 | 0.41 | 0.42668 | 0.409 | 0.4435 | 0.3545 | 0.4625 | 0.36 | 0.28 | 0.3905 | 0.376 | 0.305
| Std Dev. |  | 0.0077974 | 0.022454 | 0.0027108 | 0 | 0.0096471 | 0.012222 | 0.050596 | 0.0094339 | 0.005004 | 0.0075277 | 0.0051191 | 0.013663 | 0.015674 | 0.0091797 | 0.0030605 | 0.015795 | 0.0090066 | 0.017224 | 0.010328 | 0.0072296 | 0.018843 | 0.014882
| Rel.Std.Dev. |  | 2.196% | 13.68% | 0.81% | 0% | 2.476% | 3.835% | 14.05% | 2.083% | 1.018% | 1.704% | 1.113% | 3.305% | 3.705% | 2.246% | 0.689% | 4.408% | 1.941% | 4.762% | 3.645% | 1.852% | 5.003% | 4.853%
| PDM3 |  | -8.263% | -57.58% | -13.51% | -27.64% | 0.6953% | -17.64% | -6.971% | 17.05% | 27.06% | 14.13% | 18.87% | 6.811% | 9.332% | 5.605% | 14.78% | -7.402% | 19.9% | -6.54% | -26.78% | 0.8676% | -2.664% | -20.75%
| Z-Score (Absolute) |  | 0.36 | 2.76 | 0.61 | 1.3 | 0.08 | 0.81 | 0.29 | 0.88 | 1.37 | 0.74 | 0.97 | 0.38 | 0.5 | 0.32 | 0.77 | 0.31 | 1.02 | 0.27 | 1.26 | 0.09 | 0.08 | 0.96
| Analytical results for Tl in OREAS 236 (Certified Value 0.13 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.13 | 0.13 | < 10 | 0.13 | 0.11 | 0.14 | 0.12 | 0.12 | 0.13 | 0.11 | 0.13 | 0.13 | 0.13 | 0.1 | 0.13 | < 5 | < 5 | 0.13 | 0.13 | 0.12 |  |  |  |  | 1
| 1 | 2 | 0.13 | 0.13 | < 10 | 0.13 | 0.12 | 0.14 | 0.12 | 0.12 | 0.13 | 0.11 | 0.13 | 0.14 | 0.12 | 0.1 | 0.14 | < 5 | < 5 | 0.12 | 0.13 | 0.12 |  |  |  |  | 26
| 1 | 3 | 0.13 | 0.13 | < 10 | 0.12 | 0.12 | 0.14 | 0.13 | 0.12 | 0.13 | 0.11 | 0.13 | 0.13 | 0.12 | 0.11 | 0.14 | < 5 | < 5 | 0.12 | 0.13 | 0.12 |  |  |  |  | 16
| 1 | 4 | 0.14 | 0.13 | < 10 | 0.13 | 0.12 | 0.14 | 0.13 | 0.12 | 0.13 | 0.1 | 0.13 | 0.14 | 0.12 | 0.1 | 0.16 | < 5 | < 5 | 0.12 | 0.13 | 0.12 |  |  |  |  | 0.12677
| 1 | 5 | 0.12 | 0.14 | < 10 | 0.13 | 0.13 | 0.14 | 0.13 | 0.12 | 0.13 | 0.1 | 0.13 | 0.13 | 0.12 | 0.1 | 0.17 | < 5 | < 5 | 0.12 | 0.13 | 0.12 |  |  |  |  | 128
| 1 | 6 | 0.13 | 0.15 | < 10 | 0.13 | 0.12 | 0.14 | 0.13 | 0.12 | 0.13 | 0.11 | 0.13 | 0.14 | 0.12 | 0.11 | 0.13 | < 5 | < 5 | 0.12 | 0.13 | 0.11
| Mean |  | 0.13 | 0.135 |  | 0.12833 | 0.12 | 0.14 | 0.12667 | 0.12 | 0.13 | 0.10667 | 0.13 | 0.135 | 0.12167 | 0.10333 | 0.145 |  |  | 0.12167 | 0.13 | 0.11833
| Median |  | 0.13 | 0.13 |  | 0.13 | 0.12 | 0.14 | 0.13 | 0.12 | 0.13 | 0.11 | 0.13 | 0.135 | 0.12 | 0.1 | 0.14 |  |  | 0.12 | 0.13 | 0.12
| Std Dev. |  | 0.0063246 | 0.0083666 |  | 0.0040825 | 0.0063246 | 0 | 0.005164 | 0 | 0 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.005164 | 0.016432 |  |  | 0.0040825 | 0 | 0.0040825
| Rel.Std.Dev. |  | 4.865% | 6.197% |  | 3.181% | 5.27% | 0% | 4.077% | 0% | 0% | 4.841% | 0% | 4.057% | 3.355% | 4.997% | 11.33% |  |  | 3.355% | 0% | 3.45%
| PDM3 |  | 2.547% | 6.491% |  | 1.233% | -5.341% | 10.44% | -0.08217% | -5.341% | 2.547% | -15.86% | 2.547% | 6.491% | -4.026% | -18.49% | 14.38% |  |  | -4.026% | 2.547% | -6.656%
| Z-Score (Absolute) |  | 0 | 0.37 | 358.24 | 0.12 | 0.74 | 0.74 | 0.25 | 0.74 | 0 | 1.72 | 0 | 0.37 | 0.61 | 1.96 | 1.1 | 174.34 | 174.34 | 0.61 | 0 | 0.86
| Analytical results for Tm in OREAS 236 (Certified Value 0.2 ppm) | [25 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.02 | 1.08 | 1.10 | 1.14 | 1.15 | [19 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [19 empty] | 2
|  |  | 40g | 25g | 25g | 0.5g | 0.5g | [19 empty] | 2
| 1 | 1 | 0.16 | 0.176 | 0.21 | 0.2 | 0.22103 | [19 empty] | 1
| 1 | 2 | 0.16 | 0.185 | 0.22 | 0.2 | 0.22644 | [19 empty] | 27
| 1 | 3 | 0.16 | 0.185 | 0.22 | 0.2 | 0.20877 | [19 empty] | 16
| 1 | 4 | 0.18 | 0.187 | 0.22 | 0.2 | 0.21431 | [19 empty] | 0.19545
| 1 | 5 | 0.17 | 0.19 | 0.21 | 0.2 | 0.2183 | [19 empty] | 129
| 1 | 6 | 0.16 | 0.18 | 0.18 | 0.2 | 0.22152
| Mean |  | 0.165 | 0.18383 | 0.21 | 0.2 | 0.2184
| Median |  | 0.16 | 0.185 | 0.215 | 0.2 | 0.21967
| Std Dev. |  | 0.0083666 | 0.0050365 | 0.015492 | 3.0405e-17 | 0.006175
| Rel.Std.Dev. |  | 5.071% | 2.74% | 7.377% | 1.52e-14% | 2.827%
| PDM3 |  | -15.58% | -5.942% | 7.447% | 2.33% | 11.74%
| Z-Score (Absolute) |  | 1.46 | 0.67 | 0.42 | 0 | 0.77
| Analytical results for U in OREAS 236 (Certified Value 0.24 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 40g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.23 | 0.2 | 0.24 | 0.24 | 0.2 | 0.24 | 0.27 | 0.26 | 0.24 | 0.3 | 0.28345 | 0.26 | 0.24 | 0.25 | 0.2 | < 10 | < 10 | 0.23 | 0.24 | 0.23 |  |  |  |  | 1
| 1 | 2 | 0.25 | 0.24 | 0.23 | 0.22 | 0.2 | 0.24 | 0.26 | 0.26 | 0.24 | 0.3 | 0.2812 | 0.26 | 0.25 | 0.25 | 0.3 | < 10 | < 10 | 0.22 | 0.23 | 0.23 |  |  |  |  | 28
| 1 | 3 | 0.23 | 0.24 | 0.22 | 0.23 | 0.2 | 0.25 | 0.28 | 0.26 | 0.24 | 0.3 | 0.29011 | 0.26 | 0.23 | 0.25 | 0.3 | < 10 | < 10 | 0.24 | 0.23 | 0.22 |  |  |  |  | 16
| 1 | 4 | 0.24 | 0.24 | 0.24 | 0.23 | 0.2 | 0.24 | 0.28 | 0.26 | 0.24 | 0.3 | 0.28007 | 0.26 | 0.24 | 0.26 | 0.3 | < 10 | < 10 | 0.23 | 0.23 | 0.23 |  |  |  |  | 0.24262
| 1 | 5 | 0.22 | 0.24 | 0.23 | 0.24 | 0.2 | 0.25 | 0.26 | 0.26 | 0.24 | 0.3 | 0.28187 | 0.27 | 0.24 | 0.25 | 0.2 | < 10 | < 10 | 0.23 | 0.23 | 0.23 |  |  |  |  | 130
| 1 | 6 | 0.23 | 0.24 | 0.24 | 0.23 | 0.2 | 0.23 | 0.24 | 0.26 | 0.26 | 0.3 | 0.28526 | 0.26 | 0.25 | 0.26 | 0.2 | < 10 | < 10 | 0.23 | 0.25 | 0.22
| Mean |  | 0.23333 | 0.23333 | 0.23333 | 0.23167 | 0.2 | 0.24167 | 0.265 | 0.26 | 0.24333 | 0.3 | 0.28366 | 0.26167 | 0.24167 | 0.25333 | 0.25 |  |  | 0.23 | 0.235 | 0.22667
| Median |  | 0.23 | 0.24 | 0.235 | 0.23 | 0.2 | 0.24 | 0.265 | 0.26 | 0.24 | 0.3 | 0.28266 | 0.26 | 0.24 | 0.25 | 0.25 |  |  | 0.23 | 0.23 | 0.23
| Std Dev. |  | 0.010328 | 0.01633 | 0.008165 | 0.0075277 | 3.0405e-17 | 0.0075277 | 0.015166 | 0 | 0.008165 | 0 | 0.0036428 | 0.0040825 | 0.0075277 | 0.005164 | 0.054772 |  |  | 0.0063246 | 0.0083666 | 0.005164
| Rel.Std.Dev. |  | 4.426% | 6.999% | 3.499% | 3.249% | 1.52e-14% | 3.115% | 5.723% | 0% | 3.355% | 0% | 1.284% | 1.56% | 3.115% | 2.038% | 21.91% |  |  | 2.75% | 3.56% | 2.278%
| PDM3 |  | -3.827% | -3.827% | -3.827% | -4.514% | -17.57% | -0.3925% | 9.225% | 7.164% | 0.2944% | 23.65% | 16.92% | 7.851% | -0.3925% | 4.416% | 3.042% |  |  | -5.201% | -3.14% | -6.575%
| Z-Score (Absolute) |  | 0.48 | 0.48 | 0.48 | 0.58 | NA | 0 | 1.35 | 1.06 | 0.1 | NA | 2.43 | 1.16 | 0 | 0.67 | NA | 275.02 | 275.02 | 0.67 | 0.39 | 0.87
Results from laboratories 7, 14 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 236 (Certified Value 158 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 0
| 1 | 1 | 153 | 126 | 148.688 | 136 | 144 | 173 | 139 | 154 | 176 | 152.673 | 161 | 180 | 171 | 157.2548 | 167 | 174 | 156 | 233.9682 | 153 | 153 | 165 | 161 | 141 |  | 1
| 1 | 2 | 153 | 130 | 150.3382 | 134 | 142 | 173 | 147 | 161 | 176 | 156.024 | 158 | 184 | 169 | 165.5916 | 164 | 172 | 156 | 236.5907 | 150 | 152 | 164 | 165 | 147 |  | 29
| 1 | 3 | 154 | 134 | 150.8315 | 131 | 140 | 173 | 144 | 160 | 174 | 154.364 | 158 | 181 | 167 | 170.6132 | 157 | 177 | 152 | 232.6156 | 147 | 150 | 161 | 156 | 138 |  | 16
| 1 | 4 | 155 | 142 | 150.4475 | 133 | 144 | 172 | 148 | 156 | 177 | 152.104 | 160 | 187 | 170 | 158.4021 | 162 | 179 | 162 | 233.8078 | 150 | 151 | 161 | 157 | 142 |  | 158.2314
| 1 | 5 | 152 | 138 | 153.9932 | 133 | 140 | 169 | 150 | 153 | 174 | 156.447 | 158 | 182 | 173 | 161.3321 | 161 | 179 | 159 | 231.6483 | 155 | 151 | 163 | 157 | 139 |  | 131
| 1 | 6 | 158 | 122 | 153.7633 | 133 | 140 | 172 | 139 | 131 | 176 | 155.648 | 158 | 180 | 173 | 159.0592 | 163 | 177 | 157 | 240.5553 | 155 | 153 | 163 | 168 | 145
| Mean |  | 154.1667 | 132 | 151.3436 | 133.3333 | 141.6667 | 172 | 144.5 | 152.5 | 175.5 | 154.5433 | 158.8333 | 182.3333 | 170.5 | 162.0422 | 162.3333 | 176.3333 | 157 | 234.8643 | 151.6667 | 151.6667 | 162.8333 | 160.6667 | 142
| Median |  | 153.5 | 132 | 150.6395 | 133 | 141 | 172.5 | 145.5 | 155 | 176 | 155.006 | 158 | 181.5 | 170.5 | 160.1957 | 162.5 | 177 | 156.5 | 233.888 | 151.5 | 151.5 | 163 | 159 | 141.5
| Std Dev. |  | 2.137 | 7.4833 | 2.0978 | 1.633 | 1.9664 | 1.5492 | 4.6797 | 11.0045 | 1.2247 | 1.8175 | 1.3292 | 2.7325 | 2.3452 | 5.1309 | 3.3267 | 2.8048 | 3.3466 | 3.2461 | 3.2042 | 1.2111 | 1.6021 | 4.9261 | 3.4641
| Rel.Std.Dev. |  | 1.386% | 5.669% | 1.386% | 1.225% | 1.388% | 0.9007% | 3.239% | 7.216% | 0.6979% | 1.176% | 0.8368% | 1.499% | 1.375% | 3.166% | 2.049% | 1.591% | 2.132% | 1.382% | 2.113% | 0.7985% | 0.9839% | 3.066% | 2.44%
| PDM3 |  | -2.569% | -16.58% | -4.353% | -15.74% | -10.47% | 8.702% | -8.678% | -3.622% | 10.91% | -2.331% | 0.3804% | 15.23% | 7.754% | 2.408% | 2.592% | 11.44% | -0.7782% | 48.43% | -4.149% | -4.149% | 2.908% | 1.539% | -10.26%
| Z-Score (Absolute) |  | 0.33 | 2.89 | 0.65 | 2.74 | 1.77 | 1.73 | 1.44 | 0.52 | 2.14 | 0.28 | 0.21 | 2.93 | 1.56 | 0.58 | 0.62 | 2.23 | 0 | 9 | 0.62 | 0.62 | 0.67 | 0.42 | 1.73
| Analytical results for W in OREAS 236 (Certified Value 22.6 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 15g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 21.8 | 11.8 | 33.6133 | 25.8 | 27.8 | 21.24 | 20.41 | 24.1 | 17.53 | 17.4 | 26 | 31.9592 | 23.8 | 11.3 | 18.8 | 19 | 24 | 26.1 | 19.99 | 24.2 |  |  |  |  | 1
| 1 | 2 | 23.8 | 11.7 | 32.54 | 25.5 | 27.6 | 21.13 | 21.3 | 24.9 | 18.02 | 18.4 | 26 | 23.8043 | 23.5 | 11.2 | 21.3 | 16 | 24 | 25.6 | 20.09 | 24.5 |  |  |  |  | 30
| 1 | 3 | 22.2 | 11.6 | 33.7317 | 24.7 | 27.5 | 21.25 | 23.42 | 24.8 | 18.35 | 17.2 | 22 | 23.0114 | 22.7 | 11.5 | 20.7 | 18 | 23 | 26.6 | 19.89 | 23.5 |  |  |  |  | 16
| 1 | 4 | 22.8 | 11.6 | 34.8667 | 26.1 | 27.5 | 21.63 | 22.63 | 24.1 | 17.72 | 19 | 23 | 25.1351 | 23.6 | 11.5 | 20.8 | 22 | 24 | 25.9 | 20.44 | 24.3 |  |  |  |  | 22.6049
| 1 | 5 | 23.6 | 11.5 | 34.27 | 25.3 | 28.7 | 21.61 | 22.92 | 23.7 | 16.16 | 18.7 | 24 | 24.457 | 24.5 | 11.2 | 21.9 | 17 | 23 | 26.2 | 20.37 | 23 |  |  |  |  | 132
| 1 | 6 | 24.2 | 12.3 | 34.0514 | 25.6 | 27.4 | 21.21 | 23.01 | 23.8 | 15.71 | 17 | 24 | 25.6076 | 23.4 | 11.3 | 19.1 | 20 | 24 | 26.3 | 20.49 | 23.6
| Mean |  | 23.0667 | 11.75 | 33.8455 | 25.5 | 27.75 | 21.345 | 22.2817 | 24.2333 | 17.2483 | 17.95 | 24.1667 | 25.6624 | 23.5833 | 11.3333 | 20.4333 | 18.6667 | 23.6667 | 26.1167 | 20.2117 | 23.85
| Median |  | 23.2 | 11.65 | 33.8915 | 25.55 | 27.55 | 21.245 | 22.775 | 24.1 | 17.625 | 17.9 | 24 | 24.7961 | 23.55 | 11.3 | 20.75 | 18.5 | 24 | 26.15 | 20.23 | 23.9
| Std Dev. |  | 0.95219 | 0.2881 | 0.7793 | 0.47749 | 0.48477 | 0.21723 | 1.1681 | 0.50465 | 1.0641 | 0.85264 | 1.6021 | 3.2207 | 0.58452 | 0.13663 | 1.2291 | 2.1602 | 0.5164 | 0.34303 | 0.2538 | 0.57533
| Rel.Std.Dev. |  | 4.128% | 2.452% | 2.303% | 1.873% | 1.747% | 1.018% | 5.242% | 2.082% | 6.169% | 4.75% | 6.629% | 12.55% | 2.479% | 1.206% | 6.015% | 11.57% | 2.182% | 1.313% | 1.256% | 2.412%
| PDM3 |  | 2.043% | -48.02% | 49.73% | 12.81% | 22.76% | -5.574% | -1.43% | 7.204% | -23.7% | -20.59% | 6.909% | 13.53% | 4.328% | -49.86% | -9.607% | -17.42% | 4.697% | 15.54% | -10.59% | 5.508%
| Z-Score (Absolute) |  | 0.07 | 3.04 | 2.77 | 0.57 | 1.16 | 0.52 | 0.27 | 0.24 | 1.6 | 1.41 | 0.22 | 0.61 | 0.07 | 3.15 | 0.76 | 1.22 | 0.09 | 0.73 | 0.82 | 0.14
| Analytical results for Y in OREAS 236 (Certified Value 12.9 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 12.55 | 11.6 | 11.8062 | 13.03 | 11.9 | 12.8 | 11.92 | 14.68 | 13.43 | 12.124 | 12.47 | 14.2 | 12.8745 | 12.58 | 9.5107 | 14.3 | 13.1 | 12 | 13 | 13.8 | 12.02 | 12.3 |  |  | 1
| 1 | 2 | 13.4 | 11.9 | 11.8888 | 12.71 | 11.7 | 12.7 | 12.11 | 14.86 | 13.6 | 12.396 | 12.21 | 14.5 | 13.1759 | 12.78 | 10.0881 | 14.5 | 13.1 | 12 | 13 | 13.45 | 12.18 | 13.2 |  |  | 31
| 1 | 3 | 13.05 | 12 | 11.7078 | 12.54 | 12.1 | 12.9 | 12.41 | 15.18 | 13.47 | 12.051 | 12.19 | 14.5 | 12.9491 | 12.91 | 10.1685 | 13.6 | 12.8 | 12 | 13 | 13.4 | 11.73 | 12.45 |  |  | 16
| 1 | 4 | 13.5 | 12.7 | 11.8563 | 12.6 | 12.1 | 12.8 | 12.28 | 14.82 | 13.4 | 12.04 | 12.58 | 14.3 | 12.5477 | 12.75 | 9.7956 | 14 | 13.7 | 12 | 13 | 13.4 | 11.63 | 12.85 |  |  | 12.9081
| 1 | 5 | 13.3 | 12.5 | 12.122 | 12.9 | 12 | 12.6 | 12.42 | 14.36 | 13.16 | 12.23 | 12.11 | 14.1 | 13.0985 | 13.3 | 10.0326 | 13.75 | 13.3 | 13 | 12 | 13.65 | 11.89 | 12.6 |  |  | 133
| 1 | 6 | 13.5 | 11.9 | 12.1337 | 12.8 | 11.6 | 12.7 | 11.87 | 12.43 | 13.13 | 12.14 | 12.33 | 14.2 | 13.3455 | 12.88 | 9.693 | 14.45 | 13 | 13 | 13 | 13.7 | 12.34 | 12.9
| Mean |  | 13.2167 | 12.1 | 11.9191 | 12.7633 | 11.9 | 12.75 | 12.1683 | 14.3883 | 13.365 | 12.1635 | 12.315 | 14.3 | 12.9985 | 12.8667 | 9.8814 | 14.1 | 13.1667 | 12.3333 | 12.8333 | 13.5667 | 11.965 | 12.7167
| Median |  | 13.35 | 11.95 | 11.8726 | 12.755 | 11.95 | 12.75 | 12.195 | 14.75 | 13.415 | 12.132 | 12.27 | 14.25 | 13.0238 | 12.83 | 9.9141 | 14.15 | 13.1 | 12 | 13 | 13.55 | 11.955 | 12.725
| Std Dev. |  | 0.36697 | 0.41473 | 0.17289 | 0.18468 | 0.20976 | 0.10488 | 0.24012 | 0.99564 | 0.18382 | 0.13304 | 0.18064 | 0.16733 | 0.27675 | 0.24213 | 0.25624 | 0.3755 | 0.30768 | 0.5164 | 0.40825 | 0.17224 | 0.26943 | 0.32965
| Rel.Std.Dev. |  | 2.777% | 3.428% | 1.451% | 1.447% | 1.763% | 0.8226% | 1.973% | 6.92% | 1.375% | 1.094% | 1.467% | 1.17% | 2.129% | 1.882% | 2.593% | 2.663% | 2.337% | 4.187% | 3.181% | 1.27% | 2.252% | 2.592%
| PDM3 |  | 2.39% | -6.261% | -7.662% | -1.122% | -7.81% | -1.225% | -5.731% | 11.47% | 3.539% | -5.769% | -4.595% | 10.78% | 0.7003% | -0.3212% | -23.45% | 9.233% | 2.003% | -4.453% | -0.5795% | 5.102% | -7.306% | -1.483%
| Z-Score (Absolute) |  | 0.52 | 0.74 | 0.94 | 0.01 | 0.96 | 0.01 | 0.66 | 1.83 | 0.68 | 0.67 | 0.5 | 1.73 | 0.27 | 0.12 | 3.23 | 1.51 | 0.46 | NA | NA | 0.91 | 0.89 | 0.04
Results from laboratories 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 236 (Certified Value 1.29 ppm) | [25 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.08 | 1.10 | 1.14 | 1.15 | [17 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 40g | 0.25g | 0.25g | 25g | 25g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 1.02 | 1.3 | 1.2 | 1.254 | 1.41 | 1.5 | 1.3731 | [17 empty] | 1
| 1 | 2 | 1.08 | 1.3 | 1.2 | 1.235 | 1.38 | 1.5 | 1.4605 | [17 empty] | 32
| 1 | 3 | 1.1 | 1.2 | 1.2 | 1.241 | 1.47 | 1.4 | 1.4023 | [17 empty] | 16
| 1 | 4 | 1.16 | 1.2 | 1.2 | 1.273 | 1.43 | 1.4 | 1.3888 | [17 empty] | 1.2904
| 1 | 5 | 1.14 | 1.3 | 1.2 | 1.304 | 1.34 | 1.4 | 1.4378 | [17 empty] | 134
| 1 | 6 | 1.03 | 1.2 | 1.1 | 1.221 | 1.17 | 1.4 | 1.4401
| Mean |  | 1.0883 | 1.25 | 1.1833 | 1.2547 | 1.3667 | 1.4333 | 1.4171
| Median |  | 1.09 | 1.25 | 1.2 | 1.2475 | 1.395 | 1.4 | 1.42
| Std Dev. |  | 0.056716 | 0.054772 | 0.040825 | 0.029924 | 0.10596 | 0.05164 | 0.034042
| Rel.Std.Dev. |  | 5.211% | 4.382% | 3.45% | 2.385% | 7.753% | 3.603% | 2.402%
| PDM3 |  | -15.66% | -3.13% | -8.297% | -2.769% | 5.911% | 11.08% | 9.818%
| Z-Score (Absolute) |  | 1 | 0.03 | 0.43 | 0 | 0.67 | 1.08 | 0.98
| Analytical results for Zn in OREAS 236 (Certified Value 133 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 15g | 0.25g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  | 0
| 1 | 1 | 132 | 138 | 135.1207 | 146 | 125 | 147 | 130 | 138 | 132 | 119.571 | 138.6 | 139 | 122.837 | 130 | 142 | 130 | 123.8 | 132.524 | 134 | 129 | 141 | 121 | 129 |  | 1
| 1 | 2 | 136 | 140 | 137.368 | 142 | 124 | 145 | 132 | 133 | 134 | 122.279 | 131.6 | 139 | 127.699 | 124 | 139 | 130 | 129.4 | 132.0592 | 136 | 129 | 143 | 124 | 129 |  | 11
| 1 | 3 | 136 | 142 | 135.7407 | 143 | 120 | 146 | 134 | 129 | 135 | 123.513 | 134.8 | 140 | 124.168 | 119 | 136 | 130 | 137.8 | 126.878 | 134 | 128 | 141 | 123 | 128 |  | 16
| 1 | 4 | 136 | 144 | 133.6213 | 144 | 123 | 146 | 132 | 128 | 133 | 119.14 | 130.2 | 141 | 126.599 | 121 | 140 | 132 | 136.7 | 128.1727 | 138 | 128 | 141 | 124 | 128 |  | 132.8469
| 1 | 5 | 132 | 141 | 136.578 | 151 | 126 | 147 | 133 | 137 | 130 | 122.81 | 125.3 | 138 | 125.697 | 121 | 138 | 130 | 136.5 | 126.8622 | 139 | 128 | 143 | 122 | 128 |  | 135
| 1 | 6 | 135 | 136 | 135.4933 | 148 | 123 | 148 | 130 | 139 | 131 | 123.03 | 130.1 | 139 | 126.049 | 123 | 139 | 133 | 128.6 | 128.6598 | 138 | 128 | 142 | 126 | 130
| Mean |  | 134.5 | 140.1667 | 135.6537 | 145.6667 | 123.5 | 146.5 | 131.8333 | 134 | 132.5 | 121.7238 | 131.7667 | 139.3333 | 125.5082 | 123 | 139 | 130.8333 | 132.1333 | 129.1927 | 136.5 | 128.3333 | 141.8333 | 123.3333 | 128.6667
| Median |  | 135.5 | 140.5 | 135.617 | 145 | 123.5 | 146.5 | 132 | 135 | 132.5 | 122.5445 | 130.9 | 139 | 125.873 | 122 | 139 | 130 | 132.95 | 128.4163 | 137 | 128 | 141.5 | 123.5 | 128.5
| Std Dev. |  | 1.9748 | 2.8577 | 1.2838 | 3.3862 | 2.0736 | 1.0488 | 1.6021 | 4.7329 | 1.8708 | 1.8818 | 4.5381 | 1.0328 | 1.7464 | 3.8471 | 2 | 1.3292 | 5.6821 | 2.5071 | 2.1679 | 0.5164 | 0.98319 | 1.7512 | 0.8165
| Rel.Std.Dev. |  | 1.468% | 2.039% | 0.9464% | 2.325% | 1.679% | 0.7159% | 1.215% | 3.532% | 1.412% | 1.546% | 3.444% | 0.7412% | 1.391% | 3.128% | 1.439% | 1.016% | 4.3% | 1.941% | 1.588% | 0.4024% | 0.6932% | 1.42% | 0.6346%
| PDM3 |  | 1.244% | 5.51% | 2.113% | 9.65% | -7.036% | 10.28% | -0.7629% | 0.868% | -0.2611% | -8.373% | -0.8131% | 4.883% | -5.524% | -7.412% | 4.632% | -1.516% | -0.5371% | -2.751% | 2.75% | -3.398% | 6.765% | -7.161% | -3.147%
| Z-Score (Absolute) |  | 0.37 | 1.24 | 0.54 | 2.09 | 1.33 | 2.22 | 0.05 | 0.29 | 0.06 | 1.61 | 0.06 | 1.11 | 1.02 | 1.41 | 1.06 | 0.2 | 0 | 0.45 | 0.67 | 0.59 | 1.5 | 1.36 | 0.54
| Analytical results for Zr in OREAS 236 (Certified Value 17.8 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 40g | 0.5g | 0.25g | 0.25g | 25g | 25g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 0.15g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 22.4 | 6 | 16.8273 | 15.9 | 16.7 | 14.5 | 15.7 | 20.9 | 17.065 | 20.6 | 17 | 25.8221 | 17.4 | 32.7726 | 21.2 | 14.8 | 17 | 12 | 20.9 | 19.8 | 17.2 |  |  |  | 1
| 1 | 2 | 19.1 | 6 | 17.062 | 15.6 | 16.9 | 15.2 | 18 | 20.6 | 18.56 | 20.2 | 17.1 | 25.8167 | 16.4 | 35.78 | 21.9 | 14.8 | 16 | 12 | 20.4 | 19.8 | 18.5 |  |  |  | 12
| 1 | 3 | 18.5 | 6 | 16.616 | 15.4 | 16.9 | 14.8 | 18.8 | 20.6 | 18.232 | 20 | 16.1 | 25.8578 | 16.6 | 29.031 | 20.8 | 15 | 16 | 12 | 20.4 | 19.4 | 16.8 |  |  |  | 16
| 1 | 4 | 18.5 | 7 | 16.8273 | 15.3 | 17.4 | 15.7 | 17.9 | 18.7 | 17.26 | 19.7 | 18.6 | 26.1322 | 17.2 | 36.9197 | 21.7 | 14.6 | 17 | 11 | 20.4 | 18.4 | 17.5 |  |  |  | 17.8391
| 1 | 5 | 19.1 | 7 | 17.6767 | 15 | 15.6 | 15.7 | 16.7 | 18.8 | 18.24 | 19.2 | 18.1 | 25.8288 | 17.4 | 36.6606 | 21.4 | 14.7 | 17 | 11 | 20.7 | 19.3 | 16.7 |  |  |  | 136
| 1 | 6 | 19 | 6 | 17.6833 | 15.6 | 16.8 | 14.2 | 12.5 | 19.5 | 18.06 | 19.6 | 16.6 | 25.5668 | 17.7 | 32.5767 | 21.8 | 15.1 | 17 | 11 | 20.6 | 20.4 | 17.1
| Mean |  | 19.4333 | 6.3333 | 17.1154 | 15.4667 | 16.7167 | 15.0167 | 16.6 | 19.85 | 17.9028 | 19.8833 | 17.25 | 25.8374 | 17.1167 | 33.9568 | 21.4667 | 14.8333 | 16.6667 | 11.5 | 20.5667 | 19.5167 | 17.3
| Median |  | 19.05 | 6 | 16.9447 | 15.5 | 16.85 | 15 | 17.3 | 20.05 | 18.146 | 19.85 | 17.05 | 25.8255 | 17.3 | 34.2763 | 21.55 | 14.8 | 17 | 11.5 | 20.5 | 19.6 | 17.15
| Std Dev. |  | 1.4801 | 0.5164 | 0.45952 | 0.30768 | 0.59805 | 0.62423 | 2.2856 | 0.97724 | 0.59897 | 0.4916 | 0.93541 | 0.17959 | 0.50761 | 3.0654 | 0.41793 | 0.18619 | 0.5164 | 0.54772 | 0.20656 | 0.67057 | 0.65422
| Rel.Std.Dev. |  | 7.616% | 8.154% | 2.685% | 1.989% | 3.578% | 4.157% | 13.77% | 4.923% | 3.346% | 2.472% | 5.423% | 0.6951% | 2.966% | 9.027% | 1.947% | 1.255% | 3.098% | 4.763% | 1.004% | 3.436% | 3.782%
| PDM3 |  | 8.937% | -64.5% | -4.056% | -13.3% | -6.292% | -15.82% | -6.946% | 11.27% | 0.3575% | 11.46% | -3.302% | 44.84% | -4.049% | 90.35% | 20.34% | -16.85% | -6.572% | -35.53% | 15.29% | 9.404% | -3.022%
| Z-Score (Absolute) |  | 0.61 | NA | 0.16 | 0.71 | 0.29 | 0.86 | 0.33 | 0.75 | 0.1 | 0.76 | 0.12 | 2.74 | 0.16 | 5.45 | 1.29 | 0.92 | NA | NA | 0.99 | 0.64 | 0.1
Results from laboratories 2, 23 and 24 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C204: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C222: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C240: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C259: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C277: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C295: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C423: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C533: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C552: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 236 (Indicative Value 12.67 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.67 | 1
| 1 | 2 | 12.66 | 5
| Mean |  | 12.665 | 16
| Median |  | 12.665 | 12.665
| Std Dev. |  | 0.0070711 | 11
| Rel.Std.Dev. |  | 0.05583%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 236 (Indicative Value 70 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 6
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0 | 12
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 236 (Indicative Value 285 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 290 | 1
| 1 | 2 | 280 | 7
| Mean |  | 285 | 16
| Median |  | 285 | 285
| Std Dev. |  | 7.0711 | 13
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 236 (Indicative Value 8.56 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 8.55 | 1
| 1 | 2 | 8.56 | 8
| Mean |  | 8.555 | 16
| Median |  | 8.555 | 8.555
| Std Dev. |  | 0.0070711 | 14
| Rel.Std.Dev. |  | 0.08265%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 236 (Indicative Value 2705 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 2720 | 1
| 1 | 2 | 2690 | 9
| Mean |  | 2705 | 16
| Median |  | 2705 | 2705
| Std Dev. |  | 21.2132 | 15
| Rel.Std.Dev. |  | 0.7842%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 236 (Indicative Value 50 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 10
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 236 (Indicative Value 145 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 140 | 1
| 1 | 2 | 150 | 11
| Mean |  | 145 | 16
| Median |  | 145 | 145
| Std Dev. |  | 7.0711 | 17
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 236 (Indicative Value 185 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 190 | 12
| Mean |  | 185 | 16
| Median |  | 185 | 185
| Std Dev. |  | 7.0711 | 18
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 236 (Indicative Value 11.81 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.81 | 1
| 1 | 2 | 11.81 | 5
| Mean |  | 11.81 | 16
| Median |  | 11.81 | 11.81
| Std Dev. |  | 0 | 11
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 236 (Indicative Value 0.669 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.668 | 1
| 1 | 2 | 0.669 | 6
| Mean |  | 0.6685 | 16
| Median |  | 0.6685 | 0.6685
| Std Dev. |  | 0.00070711 | 12
| Rel.Std.Dev. |  | 0.1058%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 236 (Indicative Value 5.85 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.85 | 1
| 1 | 2 | 5.85 | 7
| Mean |  | 5.85 | 16
| Median |  | 5.85 | 5.85
| Std Dev. |  | 0 | 13
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 236 (Indicative Value 0.177 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.177 | 1
| 1 | 2 | 0.176 | 8
| Mean |  | 0.1765 | 16
| Median |  | 0.1765 | 0.1765
| Std Dev. |  | 0.00070711 | 14
| Rel.Std.Dev. |  | 0.4006%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 236 (Indicative Value 2.9 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 2.89 | 9
| Mean |  | 2.895 | 16
| Median |  | 2.895 | 2.895
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.2443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 236 (Indicative Value 75 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 70 | 10
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 7.0711 | 16
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 236 (Indicative Value 0.11 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.109 | 11
| Mean |  | 0.1095 | 16
| Median |  | 0.1095 | 0.1095
| Std Dev. |  | 0.00070711 | 17
| Rel.Std.Dev. |  | 0.6458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 236 (Indicative Value 35 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 40 | 12
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 18
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 236 (Indicative Value 0.442 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.4405 | 1
| 1 | 2 | 0.4445 | 16
| Mean |  | 0.4425 | 16
| Median |  | 0.4425 | 0.4425
| Std Dev. |  | 0.0028284 | 11
| Rel.Std.Dev. |  | 0.6392%
| PDM3 |  | 0.0006215%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 236 (Indicative Value 52.79 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 52.77 | 1
| 1 | 2 | 52.8 | 6
| Mean |  | 52.785 | 16
| Median |  | 52.785 | 52.785
| Std Dev. |  | 0.021213 | 12
| Rel.Std.Dev. |  | 0.04019%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 236 (Indicative Value 7.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 2
2
| 1 | 1 | 10 | 1
| 1 | 2 | < 10 | 7
| Mean |  | 10 | 16
| Median |  | 10 | 7.5
| Std Dev. |  |  | 13
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 236 (Indicative Value 127 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 170 | 1
| 1 | 2 | 80 | 8
| Mean |  | 125 | 16
| Median |  | 125 | 126.8392
| Std Dev. |  | 63.6396 | 14
| Rel.Std.Dev. |  | 50.91%
| PDM3 |  | -1.45%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 236 (Indicative Value 1.1 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.101 | 1
| 1 | 2 | 1.103 | 9
| Mean |  | 1.102 | 16
| Median |  | 1.102 | 1.102
| Std Dev. |  | 0.0014142 | 15
| Rel.Std.Dev. |  | 0.1283%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 236 (Indicative Value 570 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 580 | 1
| 1 | 2 | 560 | 10
| Mean |  | 570 | 16
| Median |  | 570 | 570
| Std Dev. |  | 14.1421 | 16
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 236 (Indicative Value 155 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 150 | 11
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 7.0711 | 17
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 236 (Indicative Value 78 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 81 | 1
| 1 | 2 | 74 | 12
| Mean |  | 77.5 | 16
| Median |  | 77.5 | 77.7338
| Std Dev. |  | 4.9497 | 18
| Rel.Std.Dev. |  | 6.387%
| PDM3 |  | -0.3008%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 236 (Indicative Value 3.27 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.3 | 1
| 1 | 2 | 3.24 | 14
| Mean |  | 3.27 | 16
| Median |  | 3.27 | 3.27
| Std Dev. |  | 0.042426 | 20
| Rel.Std.Dev. |  | 1.297%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 236 (Indicative Value 0.195 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.21 | 16
| Mean |  | 0.195 | 16
| Median |  | 0.195 | 0.195
| Std Dev. |  | 0.021213 | 22
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 236 (Indicative Value 0.38 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.39 | 16
| Mean |  | 0.38 | 16
| Median |  | 0.38 | 0.38
| Std Dev. |  | 0.014142 | 22
| Rel.Std.Dev. |  | 3.722%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 236 (Indicative Value 0.5 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 18
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 236 (Indicative Value 67 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 67 | 1
| 1 | 2 | 66 | 6
| Mean |  | 66.5 | 16
| Median |  | 66.5 | 66.5
| Std Dev. |  | 0.70711 | 25
| Rel.Std.Dev. |  | 1.063%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 236 (Indicative Value 258 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 259 | 1
| 1 | 2 | 257 | 20
| Mean |  | 258 | 16
| Median |  | 258 | 258
| Std Dev. |  | 1.4142 | 26
| Rel.Std.Dev. |  | 0.5481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 236 (Indicative Value 0.35 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | 0.6 | 21
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.35
| Std Dev. |  |  | 27
Rel.Std.Dev.
| PDM3 |  | 71.43%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 236 (Indicative Value 0.06 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.06 | 22
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 236 (Indicative Value 0.7 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.7 | 23
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 236 (Indicative Value 13.7 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.7 | 1
| 1 | 2 | 13.6 | 24
| Mean |  | 13.65 | 16
| Median |  | 13.65 | 13.65
| Std Dev. |  | 0.070711 | 30
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 236 (Indicative Value 44.4 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45 | 1
| 1 | 2 | 43.8 | 10
| Mean |  | 44.4 | 16
| Median |  | 44.4 | 44.4
| Std Dev. |  | 0.84853 | 31
| Rel.Std.Dev. |  | 1.911%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 236 (Indicative Value 103 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 103 | 1
| 1 | 2 | 102 | 26
| Mean |  | 102.5 | 16
| Median |  | 102.5 | 102.5
| Std Dev. |  | 0.70711 | 32
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 236 (Indicative Value 1 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 0.98 | 27
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0.028284 | 33
| Rel.Std.Dev. |  | 2.828%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 236 (Indicative Value 172 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 172 | 1
| 1 | 2 | 172 | 12
| Mean |  | 172 | 16
| Median |  | 172 | 172
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 236 (Indicative Value 4.08 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.03 | 1
| 1 | 2 | 4.12 | 29
| Mean |  | 4.075 | 16
| Median |  | 4.075 | 4.075
| Std Dev. |  | 0.06364 | 35
| Rel.Std.Dev. |  | 1.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 236 (Indicative Value 2.6 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.58 | 30
| Mean |  | 2.595 | 16
| Median |  | 2.595 | 2.595
| Std Dev. |  | 0.021213 | 36
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 236 (Indicative Value 0.99 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.99 | 1
| 1 | 2 | 0.99 | 31
| Mean |  | 0.99 | 16
| Median |  | 0.99 | 0.99
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 236 (Indicative Value 15.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.1 | 1
| 1 | 2 | 15.5 | 32
| Mean |  | 15.8 | 16
| Median |  | 15.8 | 15.8
| Std Dev. |  | 0.42426 | 38
| Rel.Std.Dev. |  | 2.685%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 236 (Indicative Value 3.52 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.63 | 1
| 1 | 2 | 3.41 | 33
| Mean |  | 3.52 | 16
| Median |  | 3.52 | 3.52
| Std Dev. |  | 0.15556 | 39
| Rel.Std.Dev. |  | 4.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 236 (Indicative Value 1.43 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.45 | 1
| Mean |  | 1.425 | 16
| Median |  | 1.425 | 1.425
| Std Dev. |  | 0.035355 | 40
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 236 (Indicative Value 2.11 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.14 | 1
| 1 | 2 | 2.07 | 18
| Mean |  | 2.105 | 16
| Median |  | 2.105 | 2.105
| Std Dev. |  | 0.049497 | 24
| Rel.Std.Dev. |  | 2.351%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 236 (Indicative Value 0.92 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.92 | 1
| 1 | 2 | 0.92 | 19
| Mean |  | 0.92 | 16
| Median |  | 0.92 | 0.92
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 236 (Indicative Value 0.075 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.1 | 20
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.035355 | 26
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 236 (Indicative Value 6.12 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.17 | 1
| 1 | 2 | 6.07 | 21
| Mean |  | 6.12 | 16
| Median |  | 6.12 | 6.12
| Std Dev. |  | 0.070711 | 27
| Rel.Std.Dev. |  | 1.155%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 236 (Indicative Value 0.37 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.37 | 1
| 1 | 2 | 0.37 | 22
| Mean |  | 0.37 | 16
| Median |  | 0.37 | 0.37
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 236 (Indicative Value 0.14 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.141 | 1
| 1 | 2 | 0.139 | 23
| Mean |  | 0.14 | 16
| Median |  | 0.14 | 0.14
| Std Dev. |  | 0.0014142 | 29
| Rel.Std.Dev. |  | 1.01%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 236 (Indicative Value 1.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.6 | 24
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 236 (Indicative Value 3.8 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.84 | 1
| 1 | 2 | 3.75 | 25
| Mean |  | 3.795 | 16
| Median |  | 3.795 | 3.795
| Std Dev. |  | 0.06364 | 31
| Rel.Std.Dev. |  | 1.677%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 236 (Indicative Value 9.54 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.46 | 1
| 1 | 2 | 9.61 | 26
| Mean |  | 9.535 | 16
| Median |  | 9.535 | 9.535
| Std Dev. |  | 0.10607 | 32
| Rel.Std.Dev. |  | 1.112%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 236 (Indicative Value 75 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 76 | 1
| 1 | 2 | 74 | 10
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 236 (Indicative Value 33 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33 | 1
| 1 | 2 | 33 | 12
| Mean |  | 33 | 16
| Median |  | 33 | 33
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 236 (Indicative Value 2.07 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2.04 | 29
| Mean |  | 2.07 | 16
| Median |  | 2.07 | 2.07
| Std Dev. |  | 0.042426 | 35
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 236 (Indicative Value 14.1 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.3 | 1
| 1 | 2 | 13.9 | 30
| Mean |  | 14.1 | 16
| Median |  | 14.1 | 14.1
| Std Dev. |  | 0.28284 | 36
| Rel.Std.Dev. |  | 2.006%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 236 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 31
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 37
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 236 (Indicative Value 1.85 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.9 | 1
| 1 | 2 | 1.8 | 32
| Mean |  | 1.85 | 16
| Median |  | 1.85 | 1.85
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 236 (Indicative Value 39.5 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.7 | 1
| 1 | 2 | 39.3 | 33
| Mean |  | 39.5 | 16
| Median |  | 39.5 | 39.5
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 0.7161%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 236 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 34
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 236 (Indicative Value 2.79 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.84 | 1
| 1 | 2 | 2.73 | 18
| Mean |  | 2.785 | 16
| Median |  | 2.785 | 2.785
| Std Dev. |  | 0.077782 | 24
| Rel.Std.Dev. |  | 2.793%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 236 (Indicative Value 1.1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1 | 7
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.14142 | 25
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 236 (Indicative Value 95 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 96 | 1
| 1 | 2 | 94.4 | 8
| Mean |  | 95.2 | 16
| Median |  | 95.2 | 95.2
| Std Dev. |  | 1.1314 | 26
| Rel.Std.Dev. |  | 1.188%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 236 (Indicative Value 0.27 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.28 | 1
| 1 | 2 | 0.26 | 21
| Mean |  | 0.27 | 16
| Median |  | 0.27 | 0.27
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 236 (Indicative Value 0.63 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.63 | 1
| 1 | 2 | 0.63 | 22
| Mean |  | 0.63 | 16
| Median |  | 0.63 | 0.63
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 236 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 236 (Indicative Value 1.12 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.13 | 1
| 1 | 2 | 1.11 | 24
| Mean |  | 1.12 | 16
| Median |  | 1.12 | 1.12
| Std Dev. |  | 0.014142 | 30
| Rel.Std.Dev. |  | 1.263%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 236 (Indicative Value 0.659 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.661 | 1
| 1 | 2 | 0.657 | 25
| Mean |  | 0.659 | 16
| Median |  | 0.659 | 0.659
| Std Dev. |  | 0.0028284 | 31
| Rel.Std.Dev. |  | 0.4292%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 236 (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 26
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 32
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 236 (Indicative Value 0.38 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.39 | 1
| 1 | 2 | 0.36 | 27
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.021213 | 33
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 236 (Indicative Value 0.38 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.37 | 28
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.0070711 | 34
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 236 (Indicative Value 300 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 302 | 1
| 1 | 2 | 297 | 29
| Mean |  | 299.5 | 16
| Median |  | 299.5 | 299.5
| Std Dev. |  | 3.5355 | 35
| Rel.Std.Dev. |  | 1.18%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 236 (Indicative Value 32.8 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33 | 1
| 1 | 2 | 32.5 | 30
| Mean |  | 32.75 | 16
| Median |  | 32.75 | 32.75
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 1.08%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 236 (Indicative Value 23 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 22.9 | 31
| Mean |  | 22.95 | 16
| Median |  | 22.95 | 22.95
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.3081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 236 (Indicative Value 2.53 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.54 | 1
| 1 | 2 | 2.52 | 32
| Mean |  | 2.53 | 16
| Median |  | 2.53 | 2.53
| Std Dev. |  | 0.014142 | 38
| Rel.Std.Dev. |  | 0.559%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 236 (Indicative Value 153 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 155 | 1
| 1 | 2 | 150 | 11
| Mean |  | 152.5 | 16
| Median |  | 152.5 | 152.5
| Std Dev. |  | 3.5355 | 39
| Rel.Std.Dev. |  | 2.318%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 236 (Indicative Value 72 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.5 | 1
| 1 | 2 | 71.5 | 12
| Mean |  | 72 | 16
| Median |  | 72 | 72
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 0.9821%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.