## Abbreviations
Table 1. Abbreviations used for OREAS 922
Abbreviation | Explanation
Std.Dev. | one sigma standard deviation
Rel.Std.Dev. | one sigma relative standard deviation
PDM3 | percent deviation of lab mean from corrected mean of means
NR | not reported
< | ‘less than’
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
IRC | infrared combustion furnace
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*AAS | sodium peroxide fusion with AAS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 922
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 0.88812 | 0.10863 | 0.84224 | 0.934 | 0.75493 | 1.0213
Al, Aluminium (wt.%) | 7.3285 | 0.42619 | 7.1306 | 7.5264 | 7.1141 | 7.5429
As, Arsenic (ppm) | 6.9113 | 0.8228 | 6.6897 | 7.1328 | 6.1545 | 7.668
Ba, Barium (ppm) | 467.9714 | 19.187 | 459.1846 | 476.7581 | 453.5415 | 482.4012
Be, Beryllium (ppm) | 2.3983 | 0.23692 | 2.2899 | 2.5066 | 2.2149 | 2.5816
Bi, Bismuth (ppm) | 10.1327 | 1.2686 | 9.6474 | 10.618 | 9.1247 | 11.1406
Ca, Calcium (wt.%) | 0.47464 | 0.024218 | 0.46317 | 0.48612 | 0.46186 | 0.48743
Cd, Cadmium (ppm) | 0.29074 | 0.033959 | 0.27132 | 0.31015 | 0.26703 | 0.31444
Ce, Cerium (ppm) | 85.7246 | 3.3763 | 83.7378 | 87.7115 | 83.7657 | 87.6836
Co, Cobalt (ppm) | 20.3889 | 1.2138 | 19.7875 | 20.9904 | 19.8973 | 20.8805
Cr, Chromium (ppm) | 75.0614 | 9.7439 | 70.3304 | 79.7924 | 72.0048 | 78.118
Cs, Cesium (ppm) | 7.271 | 0.43835 | 7.0186 | 7.5234 | 7.0937 | 7.4483
Cu, Copper (ppm) | 2121.9946 | 86.4026 | 2089.2294 | 2154.7599 | 2054.7104 | 2189.2789
Dy, Dysprosium (ppm) | 5.3851 | 0.16598 | 5.2569 | 5.5134 | 5.2542 | 5.5161
Er, Erbium (ppm) | 3.1056 | 0.15086 | 2.9829 | 3.2283 | 3.02 | 3.1911
Eu, Europium (ppm) | 1.3893 | 0.08788 | 1.3146 | 1.4641 | 1.358 | 1.4207
Fe, Iron (wt.%) | 5.5342 | 0.33164 | 5.3755 | 5.6929 | 5.4054 | 5.663
Ga, Gallium (ppm) | 20.3923 | 1.0181 | 19.8572 | 20.9274 | 19.9139 | 20.8707
Gd, Gadolinium (ppm) | 6.1756 | 0.23248 | 5.9757 | 6.3754 | 6.0245 | 6.3266
Ge, Germanium (ppm) | < 2 | IND | IND | IND | IND | IND
Hf, Hafnium (ppm) | 3.7765 | 0.28186 | 3.6246 | 3.9285 | 3.6282 | 3.9249
Ho, Holmium (ppm) | 1.0433 | 0.038787 | 1.023 | 1.0637 | 1.0179 | 1.0688
In, Indium (ppm) | 0.30964 | 0.021148 | 0.29773 | 0.32154 | 0.29596 | 0.32331
K, Potassium (wt.%) | 2.5406 | 0.063497 | 2.5169 | 2.5643 | 2.4703 | 2.6109
La, Lanthanum (ppm) | 43.3738 | 1.9898 | 42.3148 | 44.4328 | 42.2022 | 44.5453
Li, Lithium (ppm) | 30.8012 | 1.7579 | 29.8724 | 31.7299 | 29.6743 | 31.928
Lu, Lutetium (ppm) | 0.41337 | 0.019729 | 0.39808 | 0.42866 | 0.40066 | 0.42608
Mg, Magnesium (wt.%) | 1.5805 | 0.082678 | 1.5417 | 1.6194 | 1.54 | 1.6211
Mn, Manganese (wt.%) | 0.083049 | 0.0040952 | 0.081209 | 0.084888 | 0.080796 | 0.085301
Mo, Molybdenum (ppm) | 0.74475 | 0.11976 | 0.70856 | 0.78093 | 0.64405 | 0.84544
Na, Sodium (wt.%) | 0.44101 | 0.023725 | 0.42932 | 0.45269 | 0.42871 | 0.4533
Nb, Niobium (ppm) | 14.9453 | 0.73962 | 14.5084 | 15.3822 | 14.3626 | 15.5279
Nd, Neodymium (ppm) | 36.666 | 1.4966 | 35.4843 | 37.8476 | 35.8229 | 37.509
Ni, Nickel (ppm) | 37.8739 | 2.6572 | 36.6805 | 39.0672 | 36.5495 | 39.1982
P, Phosphorus (wt.%) | 0.065825 | 0.0037599 | 0.064145 | 0.067505 | 0.062784 | 0.068866
Pb, Lead (ppm) | 59.2438 | 3.1791 | 58.0133 | 60.4743 | 56.6524 | 61.8352
Pr, Praseodymium (ppm) | 10.0315 | 0.486 | 9.6287 | 10.4343 | 9.8037 | 10.2593
Rb, Rubidium (ppm) | 164.3114 | 9.3579 | 158.859 | 169.7638 | 160.6941 | 167.9287
Re, Rhenium (ppb) | < 2 | IND | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.38568 | 0.019089 | 0.37729 | 0.39407 | 0.37314 | 0.39821
Sb, Antimony (ppm) | 1.3523 | 0.089181 | 1.3037 | 1.4009 | 1.2736 | 1.4311
Sc, Scandium (ppm) | 13.1346 | 0.63723 | 12.7859 | 13.4832 | 12.7482 | 13.5209
Se, Selenium (ppm) | 3.7634 | 0.82245 | 3.2796 | 4.2473 | 3.4584 | 4.0684
Sm, Samarium (ppm) | 6.954 | 0.3935 | 6.6255 | 7.2826 | 6.7877 | 7.1203
Sn, Tin (ppm) | 9.9538 | 0.69023 | 9.6363 | 10.2713 | 9.6826 | 10.225
Sr, Strontium (ppm) | 57.5915 | 3.0484 | 56.1099 | 59.0732 | 55.8867 | 59.2964
Ta, Tantalum (ppm) | 1.2412 | 0.14057 | 1.1431 | 1.3394 | 1.2014 | 1.2811
Tb, Terbium (ppm) | 0.9335 | 0.038708 | 0.90172 | 0.96528 | 0.91197 | 0.95504
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND | IND
Th, Thorium (ppm) | 17.1647 | 1.0749 | 16.606 | 17.7234 | 16.7997 | 17.5297
Ti, Titanium (wt.%) | 0.42731 | 0.025651 | 0.41734 | 0.43728 | 0.40878 | 0.44584
Tl, Thallium (ppm) | 0.84713 | 0.050808 | 0.8203 | 0.87395 | 0.81845 | 0.87581
Tm, Thulium (ppm) | 0.44313 | 0.045908 | 0.40653 | 0.47973 | 0.41482 | 0.47143
U, Uranium (ppm) | 3.3055 | 0.17804 | 3.2136 | 3.3974 | 3.2187 | 3.3923
V, Vanadium (ppm) | 90.8656 | 4.02 | 89.1989 | 92.5323 | 86.9145 | 94.8167
W, Tungsten (ppm) | 3.8074 | 0.30946 | 3.623 | 3.9918 | 3.6165 | 3.9983
Y, Yttrium (ppm) | 28.9273 | 1.9393 | 27.8501 | 30.0045 | 28.0745 | 29.7801
Yb, Ytterbium (ppm) | 2.7989 | 0.14178 | 2.7047 | 2.8931 | 2.6954 | 2.9023
Zn, Zinc (ppm) | 267.2858 | 8.2922 | 263.925 | 270.6466 | 257.3398 | 277.2317
Zr, Zirconium (ppm) | 127.3306 | 9.2053 | 123.3349 | 131.3263 | 121.7302 | 132.931
Aqua Regia Digestion
Ag, Silver (ppm) | 0.85127 | 0.077871 | 0.82039 | 0.88215 | 0.77131 | 0.93123
Al, Aluminium (wt.%) | 2.715 | 0.15217 | 2.6347 | 2.7953 | 2.6652 | 2.7648
As, Arsenic (ppm) | 6.1175 | 0.47662 | 5.9326 | 6.3024 | 5.6597 | 6.5753
Au, Gold (ppb) | < 5 | IND | IND | IND | IND | IND
Ba, Barium (ppm) | 69.9699 | 12.6767 | 63.1872 | 76.7526 | 67.8181 | 72.1217
Be, Beryllium (ppm) | 0.65166 | 0.075856 | 0.60503 | 0.69828 | 0.6237 | 0.67961
Bi, Bismuth (ppm) | 10.3066 | 1.2888 | 9.8126 | 10.8005 | 9.4119 | 11.2012
Ca, Calcium (wt.%) | 0.3238 | 0.01822 | 0.31331 | 0.33428 | 0.31106 | 0.33653
Cd, Cadmium (ppm) | 0.27605 | 0.022135 | 0.26453 | 0.28757 | 0.25526 | 0.29685
Ce, Cerium (ppm) | 63.2671 | 7.2722 | 58.6992 | 67.835 | 61.6326 | 64.9015
Co, Cobalt (ppm) | 19.3507 | 1.1433 | 18.8376 | 19.8638 | 18.8664 | 19.835
Cr, Chromium (ppm) | 40.7392 | 1.4796 | 40.0206 | 41.4578 | 39.2427 | 42.2357
Cs, Cesium (ppm) | 1.763 | 0.23436 | 1.612 | 1.914 | 1.709 | 1.817
Cu, Copper (ppm) | 2175.5245 | 82.5636 | 2142.1572 | 2208.8918 | 2114.2147 | 2236.8343
Dy, Dysprosium (ppm) | < 5 | IND | IND | IND | IND | IND
Er, Erbium (ppm) | < 3 | IND | IND | IND | IND | IND
Eu, Europium (ppm) | < 1.5 | IND | IND | IND | IND | IND
Fe, Iron (wt.%) | 5.0477 | 0.1892 | 4.9577 | 5.1377 | 4.9502 | 5.1452
Ga, Gallium (ppm) | 7.623 | 0.42643 | 7.3629 | 7.883 | 7.3698 | 7.8762
Gd, Gadolinium (ppm) | 4.4363 | 0.80028 | 3.6626 | 5.21 | 4.2123 | 4.6602
Ge, Germanium (ppm) | 0.102 | 0.011324 | 0.090643 | 0.11336 | IND | IND
Hf, Hafnium (ppm) | 0.6109 | 0.05505 | 0.56619 | 0.6556 | 0.58201 | 0.63978
Hg, Mercury (ppm) | < 0.05 | IND | IND | IND | IND | IND
Ho, Holmium (ppm) | < 0.7 | IND | IND | IND | IND | IND
In, Indium (ppm) | 0.237 | 0.023457 | 0.22327 | 0.25072 | 0.2262 | 0.2478
K, Potassium (wt.%) | 0.37562 | 0.057671 | 0.34489 | 0.40635 | 0.36077 | 0.39047
La, Lanthanum (ppm) | 32.4515 | 3.9575 | 30.3192 | 34.5838 | 31.4281 | 33.4749
Li, Lithium (ppm) | 22.7993 | 1.3168 | 22.0849 | 23.5136 | 21.9173 | 23.6813
Lu, Lutetium (ppm) | < 0.3 | IND | IND | IND | IND | IND
Mg, Magnesium (wt.%) | 1.3348 | 0.055882 | 1.3065 | 1.363 | 1.3065 | 1.363
Mn, Manganese (wt.%) | 0.073439 | 0.0023838 | 0.072278 | 0.074601 | 0.072178 | 0.0747
Mo, Molybdenum (ppm) | 0.69174 | 0.083194 | 0.64309 | 0.7404 | 0.63216 | 0.75132
Na, Sodium (wt.%) | 0.020567 | 0.0035509 | 0.018791 | 0.022342 | 0.018754 | 0.02238
Nb, Niobium (ppm) | 0.35274 | 0.032893 | 0.32912 | 0.37636 | 0.32327 | 0.38221
Nd, Neodymium (ppm) | 27.4845 | 4.9119 | 23.0964 | 31.8726 | 26.315 | 28.6541
Ni, Nickel (ppm) | 34.2683 | 1.6413 | 33.4462 | 35.0905 | 33.223 | 35.3137
P, Phosphorus (wt.%) | 0.063434 | 0.0030407 | 0.061528 | 0.065341 | 0.061471 | 0.065398
Pb, Lead (ppm) | 59.8774 | 3.7707 | 58.1652 | 61.5896 | 57.9596 | 61.7952
Pr, Praseodymium (ppm) | 7.3255 | 0.67381 | 6.4257 | 8.2253 | 7.1085 | 7.5425
Rb, Rubidium (ppm) | 22.7156 | 3.5791 | 20.3667 | 25.0644 | 21.8774 | 23.5538
Re, Rhenium (ppb) | < 1 | IND | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.3863 | 0.024788 | 0.37315 | 0.39944 | 0.37484 | 0.39775
Sb, Antimony (ppm) | 0.56872 | 0.081306 | 0.51731 | 0.62014 | 0.5217 | 0.61574
Sc, Scandium (ppm) | 3.1472 | 0.46415 | 2.8447 | 3.4497 | 2.9778 | 3.3167
Se, Selenium (ppm) | 3.4419 | 0.47235 | 3.1967 | 3.687 | 3.0588 | 3.8249
Sm, Samarium (ppm) | 4.9759 | 0.96679 | 4.1229 | 5.8288 | 4.7701 | 5.1816
Sn, Tin (ppm) | 3.8314 | 0.38839 | 3.5837 | 4.0791 | 3.701 | 3.9618
Sr, Strontium (ppm) | 14.9972 | 0.68748 | 14.6526 | 15.3418 | 14.5633 | 15.4312
Ta, Tantalum (ppm) | < 0.05 | IND | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.61837 | 0.12003 | 0.52909 | 0.70764 | 0.59609 | 0.64064
Te, Tellurium (ppm) | < 0.03 | IND | IND | IND | IND | IND
Th, Thorium (ppm) | 14.5424 | 1.0835 | 13.9214 | 15.1634 | 14.1611 | 14.9237
Ti, Titanium (wt.%) | < 0.15 | IND | IND | IND | IND | IND
Tl, Thallium (ppm) | 0.14227 | 0.020798 | 0.12747 | 0.15706 | IND | IND
Tm, Thulium (ppm) | < 0.4 | IND | IND | IND | IND | IND
U, Uranium (ppm) | 1.9777 | 0.25781 | 1.8263 | 2.1291 | 1.9077 | 2.0478
V, Vanadium (ppm) | 29.3908 | 2.2665 | 28.1085 | 30.6732 | 28.1811 | 30.6006
W, Tungsten (ppm) | 1.1178 | 0.18546 | 1.0117 | 1.2239 | 0.93627 | 1.2993
Y, Yttrium (ppm) | 16.0264 | 2.7958 | 14.4538 | 17.5989 | 15.4426 | 16.6101
Yb, Ytterbium (ppm) | < 2 | IND | IND | IND | IND | IND
Zn, Zinc (ppm) | 256.0471 | 12.766 | 250.2667 | 261.8275 | 248.3078 | 263.7864
Zr, Zirconium (ppm) | 22.304 | 3.0194 | 20.6411 | 23.9669 | 20.6794 | 23.9287
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 7.5925 | 0.23038 | 7.4091 | 7.7759 | 7.4212 | 7.7638
As, Arsenic (ppm) | < 8 | IND | IND | IND | IND | IND
Ba, Barium (ppm) | 480.8062 | 12.2083 | 474.0023 | 487.6101 | 467.791 | 493.8214
Be, Beryllium (ppm) | < 3 | IND | IND | IND | IND | IND
Bi, Bismuth (ppm) | 10.7943 | 1.0438 | 10.133 | 11.4555 | 9.7763 | 11.8122
Ca, Calcium (wt.%) | 0.48635 | 0.039755 | 0.4696 | 0.50311 | 0.4475 | 0.52521
Cd, Cadmium (ppm) | < 0.5 | IND | IND | IND | IND | IND
Ce, Cerium (ppm) | 87.5198 | 5.5026 | 82.8358 | 92.2037 | 85.4718 | 89.5677
Co, Cobalt (ppm) | 20.9039 | 1.369 | 20.2948 | 21.5131 | 19.503 | 22.3048
Cr, Chromium (ppm) | 90.2367 | 12.1442 | 82.7564 | 97.7169 | 84.1905 | 96.2829
Cs, Cesium (ppm) | 7.495 | 0.31471 | 7.2399 | 7.7501 | 7.3146 | 7.6754
Cu, Copper (ppm) | 2214.5991 | 89.4098 | 2165.6701 | 2263.5281 | 2153.1202 | 2276.078
Dy, Dysprosium (ppm) | 5.7482 | 0.21901 | 5.5539 | 5.9424 | 5.6196 | 5.8768
Er, Erbium (ppm) | 3.3803 | 0.27383 | 3.1445 | 3.6161 | 3.2601 | 3.5006
Eu, Europium (ppm) | 1.5229 | 0.10118 | 1.4411 | 1.6047 | IND | IND
Fe, Iron (wt.%) | 5.7119 | 0.20938 | 5.5938 | 5.83 | 5.6084 | 5.8154
Ga, Gallium (ppm) | 21.2177 | 1.3895 | 20.248 | 22.1874 | 20.6815 | 21.7539
Gd, Gadolinium (ppm) | 6.9411 | 0.66847 | 6.3611 | 7.5211 | 6.7563 | 7.1259
Hf, Hafnium (ppm) | 5.929 | 0.68942 | 5.1884 | 6.6696 | IND | IND
Ho, Holmium (ppm) | 1.2003 | 0.085043 | 1.1346 | 1.266 | IND | IND
In, Indium (ppm) | 0.33525 | 0.046286 | 0.30009 | 0.37041 | IND | IND
K, Potassium (wt.%) | 2.5969 | 0.09981 | 2.5327 | 2.6612 | 2.5079 | 2.686
La, Lanthanum (ppm) | 45.6006 | 3.5029 | 44.126 | 47.0751 | 43.6763 | 47.5248
Li, Lithium (ppm) | 28.8037 | 3.3145 | 26.4842 | 31.1232 | 27.2263 | 30.3811
Lu, Lutetium (ppm) | 0.48553 | 0.068233 | 0.39782 | 0.57325 | IND | IND
Mg, Magnesium (wt.%) | 1.6127 | 0.058687 | 1.5745 | 1.6508 | 1.5759 | 1.6494
Mn, Manganese (wt.%) | 0.087587 | 0.0027903 | 0.08557 | 0.089604 | 0.085587 | 0.089587
Mo, Molybdenum (ppm) | < 1 | IND | IND | IND | IND | IND
Nb, Niobium (ppm) | 15.227 | 1.9582 | 13.5476 | 16.9064 | 14.6963 | 15.7577
Nd, Neodymium (ppm) | 38.8761 | 1.0735 | 38.1049 | 39.6473 | 37.925 | 39.8272
Ni, Nickel (ppm) | 43.3695 | 8.2839 | 37.4537 | 49.2854 | 41.8276 | 44.9115
P, Phosphorus (wt.%) | 0.065664 | 0.0048123 | 0.06241 | 0.068919 | 0.060896 | 0.070433
Pb, Lead (ppm) | 63.6088 | 5.3714 | 60.4849 | 66.7327 | 57.6824 | 69.5353
Pr, Praseodymium (ppm) | 10.6446 | 0.35122 | 10.3772 | 10.9121 | 10.4141 | 10.8751
Rb, Rubidium (ppm) | 167.485 | 5.2334 | 162.7968 | 172.1732 | 164.5583 | 170.4117
S, Sulphur (wt.%) | 0.38894 | 0.017138 | 0.37864 | 0.39925 | 0.35832 | 0.41957
Sb, Antimony (ppm) | < 2 | IND | IND | IND | IND | IND
Si, Silicon (wt.%) | 30.5133 | 0.90924 | 29.9276 | 31.0991 | 29.2627 | 31.764
Sm, Samarium (ppm) | 7.3112 | 0.27285 | 7.1091 | 7.5134 | 7.0962 | 7.5263
Sn, Tin (ppm) | 10.018 | 1.5734 | 8.6232 | 11.4128 | 7.7766 | 12.2593
Sr, Strontium (ppm) | 58.3976 | 3.2787 | 55.8496 | 60.9456 | 54.4345 | 62.3607
Ta, Tantalum (ppm) | 1.3276 | 0.15887 | 1.19 | 1.4651 | IND | IND
Tb, Terbium (ppm) | 1.0234 | 0.064183 | 0.97064 | 1.0762 | IND | IND
Th, Thorium (ppm) | 17.6598 | 0.78057 | 16.9865 | 18.3332 | 17.2075 | 18.1121
Ti, Titanium (wt.%) | 0.43924 | 0.019525 | 0.42703 | 0.45144 | 0.4263 | 0.45218
Tl, Thallium (ppm) | 0.88234 | 0.056044 | 0.84596 | 0.91873 | IND | IND
Tm, Thulium (ppm) | 0.51331 | 0.062421 | 0.45705 | 0.56958 | IND | IND
U, Uranium (ppm) | 3.5926 | 0.14638 | 3.4755 | 3.7097 | 3.4866 | 3.6986
V, Vanadium (ppm) | 92.2057 | 5.7646 | 88.6772 | 95.7342 | 88.4012 | 96.0102
W, Tungsten (ppm) | < 5 | IND | IND | IND | IND | IND
Y, Yttrium (ppm) | 31.0675 | 1.0036 | 30.2505 | 31.8845 | 30.2823 | 31.8527
Yb, Ytterbium (ppm) | 3.1742 | 0.26773 | 2.9491 | 3.3992 | 3.0222 | 3.3261
Zn, Zinc (ppm) | 276.6399 | 18.2605 | 264.9296 | 288.3503 | 253.7732 | 299.5067

## Indicative Values
Table 3. Indicative Values for OREAS 922
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 9 | Hg | ppm | < 1 | Ru | ppm | < 0.1
Aqua Regia Digestion
B | ppm | 35.6256 | Pt | ppb | < 5
Pd | ppb | < 10 | Ru | ppm | < 0.005
Infrared Combustion
S | wt.% | 0.39783
Peroxide Fusion ICP
Ag | ppm | 2.8333 | Re | ppm | < 0.1 | Te | ppm | < 6
B | ppm | 42.5 | Sc | ppm | 11.9592 | Zr | ppm | 206.1071
Ge | ppm | 2.3611 | Se | ppm | < 20

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 922
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 0.88812 | 0.10863 | 0.67086 | 1.1054 | 0.56224 | 1.214 | 12.23% | 24.46% | 36.69% | 0.84372 | 0.93253
Al, wt.% | 7.3285 | 0.42619 | 6.4761 | 8.1809 | 6.0499 | 8.607 | 5.815% | 11.63% | 17.45% | 6.9621 | 7.6949
As, ppm | 6.9113 | 0.8228 | 5.2657 | 8.5569 | 4.4428 | 9.3797 | 11.91% | 23.81% | 35.72% | 6.5657 | 7.2568
Ba, ppm | 467.9714 | 19.187 | 429.5974 | 506.3453 | 410.4104 | 525.5323 | 4.1% | 8.2% | 12.3% | 444.5728 | 491.3699
Be, ppm | 2.3983 | 0.23692 | 1.9244 | 2.8721 | 1.6875 | 3.109 | 9.879% | 19.76% | 29.64% | 2.2784 | 2.5182
Bi, ppm | 10.1327 | 1.2686 | 7.5955 | 12.6698 | 6.327 | 13.9384 | 12.52% | 25.04% | 37.56% | 9.626 | 10.6393
Ca, wt.% | 0.47464 | 0.024218 | 0.42621 | 0.52308 | 0.40199 | 0.5473 | 5.102% | 10.2% | 15.31% | 0.45091 | 0.49837
Cd, ppm | 0.29074 | 0.033959 | 0.22282 | 0.35865 | 0.18886 | 0.39261 | 11.68% | 23.36% | 35.04% | 0.2762 | 0.30527
Ce, ppm | 85.7246 | 3.3763 | 78.972 | 92.4773 | 75.5956 | 95.8536 | 3.939% | 7.877% | 11.82% | 81.4384 | 90.0108
Co, ppm | 20.3889 | 1.2138 | 17.9613 | 22.8166 | 16.7474 | 24.0304 | 5.953% | 11.91% | 17.86% | 19.3695 | 21.4084
Cr, ppm | 75.0614 | 9.7439 | 55.5736 | 94.5492 | 45.8297 | 104.2931 | 12.98% | 25.96% | 38.94% | 71.3083 | 78.8145
Cs, ppm | 7.271 | 0.43835 | 6.3943 | 8.1477 | 5.9559 | 8.586 | 6.029% | 12.06% | 18.09% | 6.9074 | 7.6345
Cu, ppm | 2121.9946 | 86.4026 | 1949.1895 | 2294.7998 | 1862.7869 | 2381.2024 | 4.072% | 8.144% | 12.22% | 2015.8949 | 2228.0944
Dy, ppm | 5.3851 | 0.16598 | 5.0532 | 5.7171 | 4.8872 | 5.8831 | 3.082% | 6.165% | 9.247% | 5.1159 | 5.6544
Er, ppm | 3.1056 | 0.15086 | 2.8039 | 3.4073 | 2.653 | 3.5581 | 4.858% | 9.715% | 14.57% | 2.9503 | 3.2608
Eu, ppm | 1.3893 | 0.08788 | 1.2136 | 1.5651 | 1.1257 | 1.653 | 6.325% | 12.65% | 18.98% | 1.3199 | 1.4588
Fe, wt.% | 5.5342 | 0.33164 | 4.8709 | 6.1975 | 4.5393 | 6.5291 | 5.993% | 11.99% | 17.98% | 5.2575 | 5.8109
Ga, ppm | 20.3923 | 1.0181 | 18.3562 | 22.4284 | 17.3381 | 23.4465 | 4.992% | 9.985% | 14.98% | 19.3727 | 21.4119
Gd, ppm | 6.1756 | 0.23248 | 5.7106 | 6.6405 | 5.4781 | 6.873 | 3.764% | 7.529% | 11.29% | 5.8668 | 6.4843
Ge, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Hf, ppm | 3.7765 | 0.28186 | 3.2128 | 4.3402 | 2.9309 | 4.6221 | 7.463% | 14.93% | 22.39% | 3.5877 | 3.9653
Ho, ppm | 1.0433 | 0.038787 | 0.96577 | 1.1209 | 0.92698 | 1.1597 | 3.718% | 7.435% | 11.15% | 0.99117 | 1.0955
In, ppm | 0.30964 | 0.021148 | 0.26734 | 0.35193 | 0.24619 | 0.37308 | 6.83% | 13.66% | 20.49% | 0.29415 | 0.32512
K, wt.% | 2.5406 | 0.063497 | 2.4136 | 2.6676 | 2.3501 | 2.7311 | 2.499% | 4.999% | 7.498% | 2.4136 | 2.6676
La, ppm | 43.3738 | 1.9898 | 39.3943 | 47.3533 | 37.4045 | 49.3431 | 4.587% | 9.175% | 13.76% | 41.2051 | 45.5425
Li, ppm | 30.8012 | 1.7579 | 27.2853 | 34.317 | 25.5274 | 36.0749 | 5.707% | 11.41% | 17.12% | 29.2611 | 32.3412
Lu, ppm | 0.41337 | 0.019729 | 0.37391 | 0.45283 | 0.35418 | 0.47256 | 4.773% | 9.545% | 14.32% | 0.3927 | 0.43404
Mg, wt.% | 1.5805 | 0.082678 | 1.4152 | 1.7459 | 1.3325 | 1.8286 | 5.231% | 10.46% | 15.69% | 1.5015 | 1.6596
Mn, wt.% | 0.083049 | 0.0040952 | 0.074858 | 0.091239 | 0.070763 | 0.095334 | 4.931% | 9.862% | 14.79% | 0.078896 | 0.087201
Mo, ppm | 0.74475 | 0.11976 | 0.50522 | 0.98427 | 0.38546 | 1.104 | 16.08% | 32.16% | 48.24% | 0.70751 | 0.78198
Na, wt.% | 0.44101 | 0.023725 | 0.39356 | 0.48846 | 0.36983 | 0.51218 | 5.38% | 10.76% | 16.14% | 0.41896 | 0.46306
Nb, ppm | 14.9453 | 0.73962 | 13.466 | 16.4245 | 12.7264 | 17.1642 | 4.949% | 9.898% | 14.85% | 14.198 | 15.6925
Nd, ppm | 36.666 | 1.4966 | 33.6728 | 39.6591 | 32.1762 | 41.1557 | 4.082% | 8.163% | 12.25% | 34.8327 | 38.4993
Ni, ppm | 37.8739 | 2.6572 | 32.5595 | 43.1882 | 29.9023 | 45.8454 | 7.016% | 14.03% | 21.05% | 35.9802 | 39.7675
P, wt.% | 0.065825 | 0.0037599 | 0.058305 | 0.073345 | 0.054545 | 0.077104 | 5.712% | 11.42% | 17.14% | 0.062534 | 0.069116
Pb, ppm | 59.2438 | 3.1791 | 52.8856 | 65.602 | 49.7065 | 68.7811 | 5.366% | 10.73% | 16.1% | 56.2816 | 62.206
Pr, ppm | 10.0315 | 0.486 | 9.0595 | 11.0035 | 8.5735 | 11.4895 | 4.845% | 9.69% | 14.53% | 9.5299 | 10.5331
Rb, ppm | 164.3114 | 9.3579 | 145.5955 | 183.0272 | 136.2376 | 192.3852 | 5.695% | 11.39% | 17.09% | 156.0958 | 172.5269
Re, ppb | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.38568 | 0.019089 | 0.3475 | 0.42385 | 0.32841 | 0.44294 | 4.95% | 9.899% | 14.85% | 0.36639 | 0.40496
Sb, ppm | 1.3523 | 0.089181 | 1.174 | 1.5307 | 1.0848 | 1.6199 | 6.595% | 13.19% | 19.78% | 1.2847 | 1.4199
Sc, ppm | 13.1346 | 0.63723 | 11.8601 | 14.409 | 11.2229 | 15.0462 | 4.852% | 9.703% | 14.55% | 12.4778 | 13.7913
Se, ppm | 3.7634 | 0.82245 | 2.1185 | 5.4083 | 1.2961 | 6.2308 | 21.85% | 43.71% | 65.56% | 3.5753 | 3.9516
Sm, ppm | 6.954 | 0.3935 | 6.167 | 7.741 | 5.7735 | 8.1345 | 5.659% | 11.32% | 16.98% | 6.6063 | 7.3017
Sn, ppm | 9.9538 | 0.69023 | 8.5733 | 11.3342 | 7.8831 | 12.0245 | 6.934% | 13.87% | 20.8% | 9.4561 | 10.4515
Sr, ppm | 57.5915 | 3.0484 | 51.4948 | 63.6883 | 48.4464 | 66.7367 | 5.293% | 10.59% | 15.88% | 54.712 | 60.4711
Ta, ppm | 1.2412 | 0.14057 | 0.9601 | 1.5224 | 0.81953 | 1.6629 | 11.32% | 22.65% | 33.97% | 1.1792 | 1.3033
Tb, ppm | 0.9335 | 0.038708 | 0.85609 | 1.0109 | 0.81738 | 1.0496 | 4.147% | 8.293% | 12.44% | 0.88683 | 0.98018
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 17.1647 | 1.0749 | 15.0148 | 19.3146 | 13.9399 | 20.3895 | 6.263% | 12.53% | 18.79% | 16.3065 | 18.0229
Ti, wt.% | 0.42731 | 0.025651 | 0.37601 | 0.47861 | 0.35036 | 0.50427 | 6.003% | 12.01% | 18.01% | 0.40595 | 0.44868
Tl, ppm | 0.84713 | 0.050808 | 0.74551 | 0.94874 | 0.69471 | 0.99955 | 5.998% | 12% | 17.99% | 0.80477 | 0.88949
Tm, ppm | 0.44313 | 0.045908 | 0.35131 | 0.53494 | 0.3054 | 0.58085 | 10.36% | 20.72% | 31.08% | 0.42097 | 0.46528
U, ppm | 3.3055 | 0.17804 | 2.9494 | 3.6616 | 2.7714 | 3.8396 | 5.386% | 10.77% | 16.16% | 3.1402 | 3.4708
V, ppm | 90.8656 | 4.02 | 82.8256 | 98.9055 | 78.8057 | 102.9255 | 4.424% | 8.848% | 13.27% | 86.3223 | 95.4089
W, ppm | 3.8074 | 0.30946 | 3.1885 | 4.4263 | 2.879 | 4.7358 | 8.128% | 16.26% | 24.38% | 3.617 | 3.9978
Y, ppm | 28.9273 | 1.9393 | 25.0487 | 32.8059 | 23.1094 | 34.7452 | 6.704% | 13.41% | 20.11% | 27.4809 | 30.3737
Yb, ppm | 2.7989 | 0.14178 | 2.5153 | 3.0824 | 2.3736 | 3.2242 | 5.065% | 10.13% | 15.2% | 2.659 | 2.9388
Zn, ppm | 267.2858 | 8.2922 | 250.7013 | 283.8702 | 242.4091 | 292.1625 | 3.102% | 6.205% | 9.307% | 253.9215 | 280.6501
Zr, ppm | 127.3306 | 9.2053 | 108.9199 | 145.7413 | 99.7146 | 154.9466 | 7.229% | 14.46% | 21.69% | 120.9641 | 133.6971
Aqua Regia Digestion
Ag, ppm | 0.85127 | 0.077871 | 0.69553 | 1.007 | 0.61766 | 1.0849 | 9.148% | 18.3% | 27.44% | 0.80871 | 0.89383
Al, wt.% | 2.715 | 0.15217 | 2.4107 | 3.0193 | 2.2585 | 3.1715 | 5.605% | 11.21% | 16.81% | 2.5793 | 2.8508
As, ppm | 6.1175 | 0.47662 | 5.1642 | 7.0707 | 4.6876 | 7.5474 | 7.791% | 15.58% | 23.37% | 5.8116 | 6.4234
Au, ppb | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 69.9699 | 12.6767 | 44.6164 | 95.3234 | 31.9397 | 108.0001 | 18.12% | 36.23% | 54.35% | 66.4714 | 73.4684
Be, ppm | 0.65166 | 0.075856 | 0.49995 | 0.80337 | 0.42409 | 0.87922 | 11.64% | 23.28% | 34.92% | 0.61907 | 0.68424
Bi, ppm | 10.3066 | 1.2888 | 7.729 | 12.8841 | 6.4402 | 14.1729 | 12.5% | 25.01% | 37.51% | 9.7912 | 10.8219
Ca, wt.% | 0.3238 | 0.01822 | 0.28736 | 0.36024 | 0.26914 | 0.37846 | 5.627% | 11.25% | 16.88% | 0.30761 | 0.33999
Cd, ppm | 0.27605 | 0.022135 | 0.23178 | 0.32032 | 0.20965 | 0.34246 | 8.019% | 16.04% | 24.06% | 0.26225 | 0.28986
Ce, ppm | 63.2671 | 7.2722 | 48.7227 | 77.8115 | 41.4505 | 85.0836 | 11.49% | 22.99% | 34.48% | 60.1037 | 66.4304
Co, ppm | 19.3507 | 1.1433 | 17.0641 | 21.6373 | 15.9208 | 22.7806 | 5.908% | 11.82% | 17.72% | 18.3832 | 20.3182
Cr, ppm | 40.7392 | 1.4796 | 37.78 | 43.6983 | 36.3005 | 45.1779 | 3.632% | 7.264% | 10.9% | 38.7022 | 42.7761
Cs, ppm | 1.763 | 0.23436 | 1.2943 | 2.2317 | 1.0599 | 2.4661 | 13.29% | 26.59% | 39.88% | 1.6748 | 1.8511
Cu, ppm | 2175.5245 | 82.5636 | 2010.3973 | 2340.6517 | 1927.8338 | 2423.2153 | 3.795% | 7.59% | 11.39% | 2066.7483 | 2284.3007
Dy, ppm | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Er, ppm | < 3 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Eu, ppm | < 1.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe, wt.% | 5.0477 | 0.1892 | 4.6693 | 5.4261 | 4.4801 | 5.6153 | 3.748% | 7.496% | 11.24% | 4.7953 | 5.3001
Ga, ppm | 7.623 | 0.42643 | 6.7701 | 8.4758 | 6.3437 | 8.9023 | 5.594% | 11.19% | 16.78% | 7.2418 | 8.0041
Gd, ppm | 4.4363 | 0.80028 | 2.8357 | 6.0368 | 2.0354 | 6.8371 | 18.04% | 36.08% | 54.12% | 4.2145 | 4.6581
Ge, ppm | 0.102 | 0.011324 | 0.079352 | 0.12465 | 0.068028 | 0.13597 | 11.1% | 22.2% | 33.31% | 0.0969 | 0.1071
Hf, ppm | 0.6109 | 0.05505 | 0.5008 | 0.721 | 0.44575 | 0.77605 | 9.011% | 18.02% | 27.03% | 0.58035 | 0.64144
Hg, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ho, ppm | < 0.7 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
In, ppm | 0.237 | 0.023457 | 0.19008 | 0.28391 | 0.16663 | 0.30737 | 9.898% | 19.8% | 29.69% | 0.22515 | 0.24885
K, wt.% | 0.37562 | 0.057671 | 0.26028 | 0.49096 | 0.2026 | 0.54863 | 15.35% | 30.71% | 46.06% | 0.35684 | 0.3944
La, ppm | 32.4515 | 3.9575 | 24.5365 | 40.3665 | 20.579 | 44.324 | 12.2% | 24.39% | 36.59% | 30.8289 | 34.0741
Li, ppm | 22.7993 | 1.3168 | 20.1656 | 25.4329 | 18.8488 | 26.7497 | 5.776% | 11.55% | 17.33% | 21.6593 | 23.9392
Lu, ppm | < 0.3 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Mg, wt.% | 1.3348 | 0.055882 | 1.223 | 1.4465 | 1.1671 | 1.5024 | 4.187% | 8.373% | 12.56% | 1.268 | 1.4015
Mn, wt.% | 0.073439 | 0.0023838 | 0.068672 | 0.078207 | 0.066288 | 0.080591 | 3.246% | 6.492% | 9.738% | 0.069767 | 0.077111
Mo, ppm | 0.69174 | 0.083194 | 0.52536 | 0.85813 | 0.44216 | 0.94132 | 12.03% | 24.05% | 36.08% | 0.65716 | 0.72633
Na, wt.% | 0.020567 | 0.0035509 | 0.013465 | 0.027668 | 0.009914 | 0.031219 | 17.27% | 34.53% | 51.8% | 0.019538 | 0.021595
Nb, ppm | 0.35274 | 0.032893 | 0.28695 | 0.41852 | 0.25406 | 0.45142 | 9.325% | 18.65% | 27.97% | 0.3351 | 0.37038
Nd, ppm | 27.4845 | 4.9119 | 17.6608 | 37.3083 | 12.749 | 42.2201 | 17.87% | 35.74% | 53.61% | 26.1103 | 28.8588
Ni, ppm | 34.2683 | 1.6413 | 30.9858 | 37.5509 | 29.3445 | 39.1922 | 4.789% | 9.579% | 14.37% | 32.5549 | 35.9818
P, wt.% | 0.063434 | 0.0030407 | 0.057353 | 0.069516 | 0.054312 | 0.072557 | 4.793% | 9.587% | 14.38% | 0.060263 | 0.066606
Pb, ppm | 59.8774 | 3.7707 | 52.3361 | 67.4187 | 48.5654 | 71.1894 | 6.297% | 12.59% | 18.89% | 56.8835 | 62.8713
Pr, ppm | 7.3255 | 0.67381 | 5.9779 | 8.6731 | 5.3041 | 9.3469 | 9.198% | 18.4% | 27.59% | 6.9592 | 7.6918
Rb, ppm | 22.7156 | 3.5791 | 15.5575 | 29.8737 | 11.9784 | 33.4528 | 15.76% | 31.51% | 47.27% | 21.5798 | 23.8514
Re, ppb | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.3863 | 0.024788 | 0.33672 | 0.43587 | 0.31193 | 0.46066 | 6.417% | 12.83% | 19.25% | 0.36698 | 0.40561
Sb, ppm | 0.56872 | 0.081306 | 0.40611 | 0.73133 | 0.32481 | 0.81264 | 14.3% | 28.59% | 42.89% | 0.54029 | 0.59716
Sc, ppm | 3.1472 | 0.46415 | 2.2189 | 4.0755 | 1.7548 | 4.5397 | 14.75% | 29.5% | 44.24% | 2.9899 | 3.3046
Se, ppm | 3.4419 | 0.47235 | 2.4972 | 4.3865 | 2.0248 | 4.8589 | 13.72% | 27.45% | 41.17% | 3.2698 | 3.6139
Sm, ppm | 4.9759 | 0.96679 | 3.0423 | 6.9094 | 2.0755 | 7.8762 | 19.43% | 38.86% | 58.29% | 4.7271 | 5.2247
Sn, ppm | 3.8314 | 0.38839 | 3.0546 | 4.6082 | 2.6662 | 4.9966 | 10.14% | 20.27% | 30.41% | 3.6398 | 4.023
Sr, ppm | 14.9972 | 0.68748 | 13.6223 | 16.3722 | 12.9348 | 17.0597 | 4.584% | 9.168% | 13.75% | 14.2474 | 15.7471
Ta, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.61837 | 0.12003 | 0.3783 | 0.85844 | 0.25826 | 0.97847 | 19.41% | 38.82% | 58.23% | 0.58745 | 0.64928
Te, ppm | < 0.03 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 14.5424 | 1.0835 | 12.3754 | 16.7094 | 11.2919 | 17.7929 | 7.451% | 14.9% | 22.35% | 13.8153 | 15.2695
Ti, wt.% | < 0.15 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 0.14227 | 0.020798 | 0.10067 | 0.18386 | 0.079872 | 0.20466 | 14.62% | 29.24% | 43.86% | 0.13515 | 0.14938
Tm, ppm | < 0.4 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
U, ppm | 1.9777 | 0.25781 | 1.4621 | 2.4933 | 1.2043 | 2.7512 | 13.04% | 26.07% | 39.11% | 1.8788 | 2.0766
V, ppm | 29.3908 | 2.2665 | 24.8579 | 33.9237 | 22.5915 | 36.1902 | 7.711% | 15.42% | 23.13% | 27.9213 | 30.8604
W, ppm | 1.1178 | 0.18546 | 0.74687 | 1.4887 | 0.56141 | 1.6742 | 16.59% | 33.18% | 49.78% | 1.0619 | 1.1737
Y, ppm | 16.0264 | 2.7958 | 10.4348 | 21.6179 | 7.6391 | 24.4136 | 17.44% | 34.89% | 52.33% | 15.225 | 16.8277
Yb, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Zn, ppm | 256.0471 | 12.766 | 230.5152 | 281.579 | 217.7493 | 294.345 | 4.986% | 9.972% | 14.96% | 243.2448 | 268.8495
Zr, ppm | 22.304 | 3.0194 | 16.2651 | 28.3429 | 13.2457 | 31.3623 | 13.54% | 27.08% | 40.61% | 21.1888 | 23.4192
Peroxide Fusion ICP
Al, wt.% | 7.5925 | 0.23038 | 7.1317 | 8.0533 | 6.9014 | 8.2836 | 3.034% | 6.069% | 9.103% | 7.2129 | 7.9721
As, ppm | < 8 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 480.8062 | 12.2083 | 456.3896 | 505.2227 | 444.1814 | 517.431 | 2.539% | 5.078% | 7.617% | 456.7659 | 504.8465
Be, ppm | < 3 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | 10.7943 | 1.0438 | 8.7067 | 12.8819 | 7.6629 | 13.9257 | 9.67% | 19.34% | 29.01% | 10.2545 | 11.334
Ca, wt.% | 0.48635 | 0.039755 | 0.40684 | 0.56586 | 0.36709 | 0.60562 | 8.174% | 16.35% | 24.52% | 0.46204 | 0.51067
Cd, ppm | < 0.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 87.5198 | 5.5026 | 76.5145 | 98.525 | 71.0119 | 104.0277 | 6.287% | 12.57% | 18.86% | 83.1438 | 91.8958
Co, ppm | 20.9039 | 1.369 | 18.166 | 23.6419 | 16.797 | 25.0109 | 6.549% | 13.1% | 19.65% | 19.8587 | 21.9491
Cr, ppm | 90.2367 | 12.1442 | 65.9483 | 114.5251 | 53.8041 | 126.6693 | 13.46% | 26.92% | 40.37% | 85.7249 | 94.7485
Cs, ppm | 7.495 | 0.31471 | 6.8656 | 8.1244 | 6.5509 | 8.4391 | 4.199% | 8.398% | 12.6% | 7.1203 | 7.8698
Cu, ppm | 2214.5991 | 89.4098 | 2035.7795 | 2393.4188 | 1946.3697 | 2482.8286 | 4.037% | 8.075% | 12.11% | 2103.8692 | 2325.3291
Dy, ppm | 5.7482 | 0.21901 | 5.3101 | 6.1862 | 5.0911 | 6.4052 | 3.81% | 7.62% | 11.43% | 5.4608 | 6.0356
Er, ppm | 3.3803 | 0.27383 | 2.8327 | 3.928 | 2.5588 | 4.2018 | 8.101% | 16.2% | 24.3% | 3.2113 | 3.5493
Eu, ppm | 1.5229 | 0.10118 | 1.3206 | 1.7253 | 1.2194 | 1.8265 | 6.644% | 13.29% | 19.93% | 1.4468 | 1.5991
Fe, wt.% | 5.7119 | 0.20938 | 5.2931 | 6.1306 | 5.0837 | 6.34 | 3.666% | 7.331% | 11% | 5.4263 | 5.9975
Ga, ppm | 21.2177 | 1.3895 | 18.4388 | 23.9966 | 17.0493 | 25.386 | 6.549% | 13.1% | 19.65% | 20.1568 | 22.2786
Gd, ppm | 6.9411 | 0.66847 | 5.6042 | 8.278 | 4.9357 | 8.9465 | 9.631% | 19.26% | 28.89% | 6.594 | 7.2882
Hf, ppm | 5.929 | 0.68942 | 4.5502 | 7.3078 | 3.8607 | 7.9973 | 11.63% | 23.26% | 34.88% | 5.6325 | 6.2254
Ho, ppm | 1.2003 | 0.085043 | 1.0302 | 1.3704 | 0.94516 | 1.4554 | 7.085% | 14.17% | 21.26% | 1.1403 | 1.2603
In, ppm | 0.33525 | 0.046286 | 0.24268 | 0.42782 | 0.19639 | 0.47411 | 13.81% | 27.61% | 41.42% | 0.31849 | 0.35201
K, wt.% | 2.5969 | 0.09981 | 2.3973 | 2.7966 | 2.2975 | 2.8964 | 3.843% | 7.687% | 11.53% | 2.4671 | 2.7268
La, ppm | 45.6006 | 3.5029 | 38.5947 | 52.6064 | 35.0918 | 56.1093 | 7.682% | 15.36% | 23.05% | 43.3205 | 47.8806
Li, ppm | 28.8037 | 3.3145 | 22.1747 | 35.4327 | 18.8601 | 38.7473 | 11.51% | 23.01% | 34.52% | 27.3635 | 30.2439
Lu, ppm | 0.48553 | 0.068233 | 0.34907 | 0.622 | 0.28084 | 0.69023 | 14.05% | 28.11% | 42.16% | 0.46126 | 0.50981
Mg, wt.% | 1.6127 | 0.058687 | 1.4953 | 1.73 | 1.4366 | 1.7887 | 3.639% | 7.278% | 10.92% | 1.532 | 1.6933
Mn, wt.% | 0.087587 | 0.0027903 | 0.082006 | 0.093168 | 0.079216 | 0.095958 | 3.186% | 6.372% | 9.557% | 0.083208 | 0.091966
Mo, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Nb, ppm | 15.227 | 1.9582 | 11.3105 | 19.1435 | 9.3523 | 21.1017 | 12.86% | 25.72% | 38.58% | 14.4657 | 15.9884
Nd, ppm | 38.8761 | 1.0735 | 36.7292 | 41.0231 | 35.6558 | 42.0965 | 2.761% | 5.522% | 8.284% | 36.9323 | 40.8199
Ni, ppm | 43.3695 | 8.2839 | 26.8017 | 59.9374 | 18.5177 | 68.2213 | 19.1% | 38.2% | 57.3% | 41.2011 | 45.538
P, wt.% | 0.065664 | 0.0048123 | 0.05604 | 0.075289 | 0.051227 | 0.080101 | 7.329% | 14.66% | 21.99% | 0.062381 | 0.068948
Pb, ppm | 63.6088 | 5.3714 | 52.866 | 74.3517 | 47.4946 | 79.7231 | 8.444% | 16.89% | 25.33% | 60.4284 | 66.7893
Pr, ppm | 10.6446 | 0.35122 | 9.9422 | 11.3471 | 9.591 | 11.6983 | 3.3% | 6.599% | 9.899% | 10.1124 | 11.1769
Rb, ppm | 167.485 | 5.2334 | 157.0182 | 177.9518 | 151.7848 | 183.1852 | 3.125% | 6.249% | 9.374% | 159.1107 | 175.8593
S, wt.% | 0.38894 | 0.017138 | 0.35467 | 0.42322 | 0.33753 | 0.44036 | 4.406% | 8.812% | 13.22% | 0.3695 | 0.40839
Sb, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Si, wt.% | 30.5133 | 0.90924 | 28.6949 | 32.3318 | 27.7856 | 33.241 | 2.98% | 5.96% | 8.939% | 28.9877 | 32.039
Sm, ppm | 7.3112 | 0.27285 | 6.7655 | 7.857 | 6.4927 | 8.1298 | 3.732% | 7.464% | 11.2% | 6.9457 | 7.6768
Sn, ppm | 10.018 | 1.5734 | 6.8712 | 13.1648 | 5.2978 | 14.7382 | 15.71% | 31.41% | 47.12% | 9.5171 | 10.5189
Sr, ppm | 58.3976 | 3.2787 | 51.8403 | 64.9549 | 48.5616 | 68.2336 | 5.614% | 11.23% | 16.84% | 55.4777 | 61.3175
Ta, ppm | 1.3276 | 0.15887 | 1.0098 | 1.6453 | 0.85097 | 1.8042 | 11.97% | 23.93% | 35.9% | 1.2612 | 1.394
Tb, ppm | 1.0234 | 0.064183 | 0.89505 | 1.1518 | 0.83087 | 1.216 | 6.271% | 12.54% | 18.81% | 0.97225 | 1.0746
Th, ppm | 17.6598 | 0.78057 | 16.0987 | 19.221 | 15.3181 | 20.0016 | 4.42% | 8.84% | 13.26% | 16.7768 | 18.5428
Ti, wt.% | 0.43924 | 0.019525 | 0.40019 | 0.47829 | 0.38066 | 0.49782 | 4.445% | 8.891% | 13.34% | 0.41728 | 0.4612
Tl, ppm | 0.88234 | 0.056044 | 0.77025 | 0.99443 | 0.71421 | 1.0505 | 6.352% | 12.7% | 19.06% | 0.83822 | 0.92646
Tm, ppm | 0.51331 | 0.062421 | 0.38847 | 0.63815 | 0.32605 | 0.70057 | 12.16% | 24.32% | 36.48% | 0.48765 | 0.53898
U, ppm | 3.5926 | 0.14638 | 3.2998 | 3.8853 | 3.1535 | 4.0317 | 4.074% | 8.149% | 12.22% | 3.413 | 3.7722
V, ppm | 92.2057 | 5.7646 | 80.6765 | 103.7349 | 74.9119 | 109.4995 | 6.252% | 12.5% | 18.76% | 87.5954 | 96.816
W, ppm | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Y, ppm | 31.0675 | 1.0036 | 29.0602 | 33.0748 | 28.0566 | 34.0784 | 3.231% | 6.461% | 9.692% | 29.5141 | 32.6209
Yb, ppm | 3.1742 | 0.26773 | 2.6387 | 3.7096 | 2.371 | 3.9773 | 8.434% | 16.87% | 25.3% | 3.0155 | 3.3329
Zn, ppm | 276.6399 | 18.2605 | 240.1188 | 313.161 | 221.8583 | 331.4216 | 6.601% | 13.2% | 19.8% | 262.8079 | 290.4719

## 4-Acid
| Analytical results for Ag in OREAS 922 (Certified Value 0.888 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  | 3
3
| 1 | 1 | 0.9 | 0.9 | 0.84 | 1 | 0.89 | 0.85 | 0.85 | 1 | 0.95 | 0.7 | 0.81884 | 1.09 | 1.4 | 0.82 | < 1 | 1 | 0.9 | 0.86 | 0.88 |  | 1
| 1 | 2 | 1 | 0.9 | 0.83 | 1.2 | 0.83 | 0.86 | 0.81 | 1 | 0.87 | 0.7 | 0.8 | 1.35 | 1.3 | 0.882 | < 1 | 0.9 | 1.1 | 0.87 | 0.82 |  | 8
| 1 | 3 | 0.9 | 0.9 | 0.87 | 1 | 1.45 | 0.91 | 0.87 | 1 | 0.77 | 0.7 | 0.77048 | 0.81 | 1.5 | 0.991 | < 1 | 1.1 | 0.9 | 1.24 | 0.89 |  | 16
| 1 | 4 | 0.8 | 1 | 0.83 | 1.1 | 1.74 | 1.8 | 1 | 1 | 1.03 | 0.8 | 0.75 | 0.95 | 1.2 | 1.043 | < 1 | 1 | 0.9 | 0.76 | 0.77 |  | 0.88812
| 1 | 5 | 1 | 0.9 | 0.83 | 1 | 1.02 | 0.84 | 1.08 | 1 | 0.76 | 0.8 | 0.85185 | 1.6 | 1.3 | 0.969 | < 1 | 1 | 1 | 0.87 | 0.73 |  | 7
| 1 | 6 | 0.9 | 0.9 | 0.87 | 1.1 | 1.06 | 0.83 | 0.78 | 1 | 1.24 | 0.6 | 0.8728 | 1.12 | 1.2 | 0.688 | < 1 | 0.8 | 0.8 | 0.84 | 0.88
| Mean |  | 0.91667 | 0.91667 | 0.845 | 1.0667 | 1.165 | 1.015 | 0.89833 | 1 | 0.93667 | 0.71667 | 0.81066 | 1.1533 | 1.3167 | 0.89883 |  | 0.96667 | 0.93333 | 0.90667 | 0.82833
| Median |  | 0.9 | 0.9 | 0.835 | 1.05 | 1.04 | 0.855 | 0.86 | 1 | 0.91 | 0.7 | 0.80942 | 1.105 | 1.3 | 0.9255 |  | 1 | 0.9 | 0.865 | 0.85
| Std Dev. |  | 0.075277 | 0.040825 | 0.019748 | 0.08165 | 0.35546 | 0.38558 | 0.11686 | 0 | 0.18129 | 0.075277 | 0.046972 | 0.28374 | 0.1169 | 0.1305 |  | 0.10328 | 0.10328 | 0.16848 | 0.066758
| Rel.Std.Dev. |  | 8.212% | 4.454% | 2.337% | 7.655% | 30.51% | 37.99% | 13.01% | 0% | 19.35% | 10.5% | 5.794% | 24.6% | 8.879% | 14.52% |  | 10.68% | 11.07% | 18.58% | 8.059%
| PDM3 |  | 3.214% | 3.214% | -4.855% | 20.1% | 31.18% | 14.29% | 1.15% | 12.6% | 5.466% | -19.31% | -8.722% | 29.86% | 48.25% | 1.206% |  | 8.844% | 5.091% | 2.088% | -6.732%
| Z-Score (Absolute) |  | 0.18 | 0.01 | 0.58 | 1.78 | 2.83 | 0.44 | 0.01 | 1.07 | 0.25 | 1.95 | 0.94 | 2.7 | 4.44 | 0.01 | NA | 0.72 | 0.01 | 0.63 | 0.76
| Analytical results for Al in OREAS 922 (Certified Value 7.33 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 7.63 | 6.75 | 7.3471 | 7.93 | 6.37 | 7.56 | 7.34 | 7.59 | 6.79 | 7.8282 | 7.23 | 5.48 | 7.048 | 3.94 | 7.11 | 7.91 | 7.22 | 7.12 |  |  | 1
| 1 | 2 | 7.77 | 6.85 | 7.4765 | 7.88 | 7.05 | 7.39 | 7.72 | 7.54 | 6.4 | 7.6996 | 7.42 | 5.66 | 6.799 | 4.66 | 7.25 | 7.8 | 7.3 | 7.43 |  |  | #N/A
| 1 | 3 | 7.74 | 6.83 | 7.4133 | 7.83 | 7.17 | 7.76 | 7.21 | 7.64 | 6.17 | 7.4493 | 6.72 | 5.63 | 6.723 | 3.75 | 7.33 | 7.75 | 7.17 | 7.26 |  |  | 16
| 1 | 4 | 7.78 | 6.65 | 7.4966 | 7.84 | 7.36 | 7.87 | 7.34 | 7.5 | 6.33 | 7.4967 | 6.71 | 5.71 | 6.92 | 3.7 | 7.05 | 7.91 | 7.38 | 7.17 |  |  | 7.3285
| 1 | 5 | 7.62 | 6.85 | 7.4233 | 7.87 | 7.36 | 6.91 | 7.39 | 7.66 | 6.7 | 7.5749 | 7.36 | 5.29 | 6.963 | 3.37 | 6.95 | 7.81 | 7.3 | 7.27 |  |  | 8
| 1 | 6 | 7.72 | 6.95 | 7.5581 | 8.03 | 8.54 | 7.18 | 7.37 | 7.6 | 6.89 | 7.552 | 7.36 | 5.45 | 6.906 | 3.83 | 7.09 | 7.97 | 7.38 | 7.3
| Mean |  | 7.71 | 6.8133 | 7.4525 | 7.8967 | 7.3083 | 7.445 | 7.395 | 7.5883 | 6.5467 | 7.6001 | 7.1333 | 5.5367 | 6.8932 | 3.875 | 7.13 | 7.8583 | 7.2917 | 7.2583
| Median |  | 7.73 | 6.84 | 7.4499 | 7.875 | 7.265 | 7.475 | 7.355 | 7.595 | 6.55 | 7.5634 | 7.295 | 5.555 | 6.913 | 3.79 | 7.1 | 7.86 | 7.3 | 7.265
| Std Dev. |  | 0.069282 | 0.10231 | 0.073716 | 0.074207 | 0.7055 | 0.36126 | 0.1712 | 0.060139 | 0.28668 | 0.14025 | 0.32995 | 0.15845 | 0.11624 | 0.42983 | 0.13799 | 0.084004 | 0.084479 | 0.10797
| Rel.Std.Dev. |  | 0.8986% | 1.502% | 0.9892% | 0.9397% | 9.653% | 4.852% | 2.315% | 0.7925% | 4.379% | 1.845% | 4.625% | 2.862% | 1.686% | 11.09% | 1.935% | 1.069% | 1.159% | 1.487%
| PDM3 |  | 5.206% | -7.029% | 1.692% | 7.753% | -0.275% | 1.59% | 0.9076% | 3.546% | -10.67% | 3.706% | -2.663% | -24.45% | -5.94% | -47.12% | -2.708% | 7.23% | -0.5024% | -0.9572%
| Z-Score (Absolute) |  | 0.94 | 1.12 | 0.35 | 1.37 | 0.02 | 0.33 | 0.07 | 0.66 | 1.73 | 0.69 | 0.38 | 4.04 | 0.93 | 8.21 | 0.39 | 1.28 | 0.02 | 0.1
| Analytical results for As in OREAS 922 (Certified Value 6.91 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
2
| 1 | 1 | 7 | 17 | 5.9 | 8 | 5.6 | 7.1 | 7.3 | 7 | 5.7 | 6 | 7.2089 | 5.4 | 8 | 6.694 | 6 | 4.5 | 8 | 7 | 6.2 |  | 1
| 1 | 2 | 7 | 17 | 4.5 | 9 | 6.4 | 6.7 | 7.1 | 7 | 6.5 | 8 | 7.4 | 7.6 | 8 | 6.902 | 7 | 5.3 | 6 | 7 | 7.3 |  | #N/A
| 1 | 3 | 7 | 17 | 6.4 | 9 | 6.4 | 7.7 | 6.8 | 7 | 6.7 | 6 | 7.2467 | 7.7 | 8 | 5.981 | 6 | 5.1 | 7 | 6 | 5.5 |  | 16
| 1 | 4 | 7 | 16 | 6.6 | 5 | 6.4 | 6.7 | 6.7 | 7 | 5.6 | 7 | 6.9019 | 6.8 | 8 | 6.833 | 9 | 4.7 | 6 | 7 | 6.7 |  | 6.9113
| 1 | 5 | 8 | 16 | 7 | 6 | 6.4 | 6.5 | 6.5 | 7 | 5.9 | 8 | 7.111 | 7.6 | 8 | 5.81 | 9 | 3.4 | 7 | 7 | 6.9 |  | 9
| 1 | 6 | 7 | 16 | 6.5 | 6 | 6.4 | 6.8 | 6.5 | 7 | 6.8 | 11 | 7.3948 | 6 | 7 | 7.025 | 14 | 4.8 | 9 | 8 | 6.4
| Mean |  | 7.1667 | 16.5 | 6.15 | 7.1667 | 6.2667 | 6.9167 | 6.8167 | 7 | 6.2 | 7.6667 | 7.2106 | 6.85 | 7.8333 | 6.5408 | 8.5 | 4.6333 | 7.1667 | 7 | 6.5
| Median |  | 7 | 16.5 | 6.45 | 7 | 6.4 | 6.75 | 6.75 | 7 | 6.2 | 7.5 | 7.2278 | 7.2 | 8 | 6.7635 | 8 | 4.75 | 7 | 7 | 6.55
| Std Dev. |  | 0.40825 | 0.54772 | 0.88261 | 1.7224 | 0.3266 | 0.43089 | 0.32506 | 0 | 0.52915 | 1.8619 | 0.18779 | 0.96695 | 0.40825 | 0.51404 | 3.0166 | 0.66833 | 1.169 | 0.63246 | 0.6229
| Rel.Std.Dev. |  | 5.696% | 3.32% | 14.35% | 24.03% | 5.212% | 6.23% | 4.769% | 0% | 8.535% | 24.29% | 2.604% | 14.12% | 5.212% | 7.859% | 35.49% | 14.42% | 16.31% | 9.035% | 9.583%
| PDM3 |  | 3.696% | 138.7% | -11.01% | 3.696% | -9.327% | 0.07826% | -1.369% | 1.284% | -10.29% | 10.93% | 4.331% | -0.8863% | 13.34% | -5.36% | 22.99% | -32.96% | 3.696% | 1.284% | -5.951%
| Z-Score (Absolute) |  | 0 | 26.69 | 1.46 | 0.47 | 1.69 | 0.67 | 0.52 | 0 | 2.25 | 0 | 0.59 | 0.42 | 2.81 | 1.29 | 1.12 | 5.96 | 0.47 | 0 | 1.4
| Analytical results for B in OREAS 922 (Indicative Value 9 ppm) | [21 empty] | Indicative
B | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 15 |  | ppm
|  |  | 01: 4A*OES |  | 2
2
| 1 | 1 | 01: 8 |  | 1
| 1 | 2 | 01: 8 |  | 1
| 1 | 3 | 01: 8 |  | 16
| 1 | 4 | 01: 10 |  | 9
| 1 | 5 | 01: 8 |  | 7
| 1 | 6 | 01: 12
| Mean |  | 01: 9
| Median |  | 01: 8
| Std Dev. |  | 01: 1.6733
| Rel.Std.Dev. |  | 01: 18.59%
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 922 (Certified Value 468 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 477 | 499 | 467.7 | 503 | 438 | 480 | 450 | 454 | 523 | 470.8416 | 562.6 | 500 | 447.034 | 296 | 467 | 455 | 451 | 450 |  |  | 1
| 1 | 2 | 480 | 479 | 458.8 | 510 | 472 | 480 | 480 | 457 | 510 | 477.805 | 539.7 | 440 | 432.778 | 352 | 471 | 455 | 464 | 470 |  |  | #N/A
| 1 | 3 | 481 | 468 | 460.5 | 492 | 479 | 500 | 450 | 464 | 484 | 485.467 | 562.5 | 480 | 419.608 | 303 | 462 | 461 | 454 | 460 |  |  | 16
| 1 | 4 | 491 | 445 | 459.7 | 499 | 493 | 500 | 460 | 451 | 510 | 481.887 | 574.4 | 440 | 429.206 | 304 | 468 | 457 | 473 | 450 |  |  | 467.9714
| 1 | 5 | 484 | 486 | 466.1 | 503 | 495 | 440 | 460 | 452 | 525 | 488.7744 | 576.2 | 440 | 446.781 | 301 | 466 | 461 | 443 | 470 |  |  | 10
| 1 | 6 | 485 | 479 | 483.3 | 501 | 485 | 460 | 460 | 465 | 536 | 478.4876 | 609.7 | 460 | 434.592 | 240 | 453 | 459 | 465 | 470
| Mean |  | 483 | 476 | 466.0167 | 501.3333 | 477 | 476.6667 | 460 | 457.1667 | 514.6667 | 480.5438 | 570.85 | 460 | 434.9998 | 299.3333 | 464.5 | 458 | 458.3333 | 461.6667
| Median |  | 482.5 | 479 | 463.3 | 502 | 482 | 480 | 460 | 455.5 | 516.5 | 480.1873 | 568.5 | 450 | 433.685 | 302 | 466.5 | 458 | 459 | 465
| Std Dev. |  | 4.858 | 18.2866 | 9.2083 | 5.8878 | 20.9476 | 23.3809 | 10.9545 | 6.047 | 17.9741 | 6.319 | 23.0593 | 25.2982 | 10.5754 | 35.6183 | 6.3482 | 2.7568 | 10.9484 | 9.8319
| Rel.Std.Dev. |  | 1.006% | 3.842% | 1.976% | 1.174% | 4.392% | 4.905% | 2.381% | 1.323% | 3.492% | 1.315% | 4.039% | 5.5% | 2.431% | 11.9% | 1.367% | 0.6019% | 2.389% | 2.13%
| PDM3 |  | 3.211% | 1.716% | -0.4177% | 7.129% | 1.929% | 1.858% | -1.703% | -2.309% | 9.978% | 2.687% | 21.98% | -1.703% | -7.046% | -36.04% | -0.7418% | -2.131% | -2.06% | -1.347%
| Z-Score (Absolute) |  | 1.03 | 0.66 | 0.05 | 1.99 | 1.12 | 0.69 | 0.19 | 0.34 | 2.69 | 0.9 | 5.65 | 0.19 | 1.5 | 8.02 | 0.05 | 0.29 | 0.27 | 0.1
| Analytical results for Be in OREAS 922 (Certified Value 2.4 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.4 | 3.2 | 2.65 | 3 | 2.4 | 2.44 | 2.42 | 2.4 | 2 | 2.3386 | 4.13 | < 5 | 2.619 | 2 | 2.3 | 2.4 | 2.8 | 2.27 |  |  | 1
| 1 | 2 | 2.3 | 3.2 | 2.57 | 3 | 2.6 | 2.49 | 2.28 | 2.3 | 2 | 2.1998 | 4.72 | < 5 | 2.77 | 3 | 2.3 | 2.4 | 3 | 2.57 |  |  | #N/A
| 1 | 3 | 2.4 | 3.2 | 2.6 | 3 | 2.6 | 2.59 | 2.31 | 2.5 | 2 | 2.2477 | 4.47 | < 5 | 2.632 | 3 | 2.5 | 2.4 | 2.8 | 2.43 |  |  | 16
| 1 | 4 | 2.4 | 3.2 | 2.71 | 3 | 2.5 | 2.54 | 2.32 | 2.5 | 2 | 2.2218 | 4.6 | < 5 | 2.718 | 3 | 2.3 | 2.1 | 3 | 2.28 |  |  | 2.3983
| 1 | 5 | 2.4 | 3.2 | 2.76 | 3 | 2.5 | 2.19 | 2.37 | 2.4 | 2 | 2.1593 | 4.49 | < 5 | 2.713 | 2 | 2.3 | 2.3 | 3.2 | 2.47 |  |  | 11
| 1 | 6 | 2.4 | 3.2 | 2.74 | 3 | 2.5 | 2.33 | 2.37 | 2.4 | 2 | 2.1116 | 4.95 | < 5 | 2.564 | 2 | 2.3 | 2.3 | 3 | 2.37
| Mean |  | 2.3833 | 3.2 | 2.6717 | 3 | 2.5167 | 2.43 | 2.345 | 2.4167 | 2 | 2.2131 | 4.56 |  | 2.6693 | 2.5 | 2.3333 | 2.3167 | 2.9667 | 2.3983
| Median |  | 2.4 | 3.2 | 2.68 | 3 | 2.5 | 2.465 | 2.345 | 2.4 | 2 | 2.2108 | 4.545 |  | 2.6725 | 2.5 | 2.3 | 2.35 | 3 | 2.4
| Std Dev. |  | 0.040825 | 4.8648e-16 | 0.077309 | 0 | 0.075277 | 0.14765 | 0.050892 | 0.075277 | 0 | 0.077965 | 0.27466 |  | 0.076722 | 0.54772 | 0.08165 | 0.1169 | 0.15055 | 0.11566
| Rel.Std.Dev. |  | 1.713% | 1.52e-14% | 2.894% | 0% | 2.991% | 6.076% | 2.17% | 3.115% | 0% | 3.523% | 6.023% |  | 2.874% | 21.91% | 3.499% | 5.046% | 5.075% | 4.822%
| PDM3 |  | -0.6227% | 33.43% | 11.4% | 25.09% | 4.937% | 1.323% | -2.221% | 0.7672% | -16.61% | -7.72% | 90.14% |  | 11.3% | 4.242% | -2.708% | -3.402% | 23.7% | 0.002789%
| Z-Score (Absolute) |  | 0.16 | 3.99 | 1.25 | 2.96 | 0.45 | 0 | 0.44 | 0.07 | 2.23 | 1.12 | 11.05 | NA | 1.24 | 0.36 | 0.67 | 0.59 | 2.78 | 0.16
| Analytical results for Bi in OREAS 922 (Certified Value 10.1 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 1
1
| 1 | 1 | 11.7 | 12 | 10.19 | 10 | 9.59 | 11.5 | 22.5 | 9 | 9.69 | < 5 | 11.1959 | 9.76 | < 5 | 11.098 | 12 | 9.2 | 10.9 | 8.94 | 9.24 |  | 1
| 1 | 2 | 11.9 | 12.3 | 9.68 | 7 | 11.2 | 12.3 | 11.45 | 9.4 | 12.07 | < 5 | 11.3294 | 10.78 | < 5 | 8.367 | 27 | 9.85 | 9.9 | 9.45 | 11.4 |  | #N/A
| 1 | 3 | 11.5 | 11.7 | 9.42 | 10 | 9.99 | 10.2 | 9.25 | 8.6 | 9.69 | < 5 | 11.9672 | 10.1 | 7 | 10.617 | 13 | 10.9 | 9.5 | 8.65 | 9.43 |  | 16
| 1 | 4 | 11.8 | 11.4 | 9.21 | 9 | 8.3 | 11.1 | 8.81 | 9.2 | 8.78 | < 5 | 9.1624 | 10.3 | 6 | 11.543 | 20 | 13.9 | 10.1 | 8.42 | 9.29 |  | 10.1327
| 1 | 5 | 11.6 | 11.9 | 9.86 | 8 | 10.5 | 9.54 | 10.05 | 9.5 | 9.91 | < 5 | 9.9933 | 9.77 | 5 | 8.11 | 21 | 9.88 | 9.4 | 10.4 | 9.4 |  | 12
| 1 | 6 | 11.9 | 12.4 | 10.36 | 8 | 9.03 | 8.76 | 10.15 | 9.5 | 9.35 | < 5 | 10.5856 | 11.06 | 6 | 8.88 | 19 | 13.7 | 10.7 | 9.89 | 10.85
| Mean |  | 11.7333 | 11.95 | 9.7867 | 8.6667 | 9.7683 | 10.5667 | 12.035 | 9.2 | 9.915 |  | 10.7056 | 10.295 | 6 | 9.7692 | 18.6667 | 11.2383 | 10.0833 | 9.2917 | 9.935
| Median |  | 11.75 | 11.95 | 9.77 | 8.5 | 9.79 | 10.65 | 10.1 | 9.3 | 9.69 |  | 10.8908 | 10.2 | 6 | 9.7485 | 19.5 | 10.39 | 10 | 9.195 | 9.415
| Std Dev. |  | 0.1633 | 0.37283 | 0.44162 | 1.2111 | 1.0361 | 1.3124 | 5.2059 | 0.35214 | 1.1272 |  | 1.0124 | 0.53298 | 0.8165 | 1.4927 | 5.5377 | 2.0585 | 0.61455 | 0.76245 | 0.94061
| Rel.Std.Dev. |  | 1.392% | 3.12% | 4.512% | 13.97% | 10.61% | 12.42% | 43.26% | 3.828% | 11.37% |  | 9.456% | 5.177% | 13.61% | 15.28% | 29.67% | 18.32% | 6.095% | 8.206% | 9.468%
| PDM3 |  | 15.8% | 17.94% | -3.415% | -14.47% | -3.596% | 4.283% | 18.77% | -9.205% | -2.148% |  | 5.654% | 1.602% | -40.79% | -3.587% | 84.22% | 10.91% | -0.487% | -8.3% | -1.951%
| Z-Score (Absolute) |  | 2.04 | 2.27 | 0.08 | 1.31 | 0.1 | 0.77 | 0.08 | 0.72 | 0.41 | NA | 0.92 | 0.47 | 4.21 | 0.1 | 9.6 | 1.5 | 0.24 | 0.62 | 0.24
| Analytical results for Ca in OREAS 922 (Certified Value 0.475 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.459 | 0.48 | 0.4654 | 0.55 | 0.44 | 0.49 | 0.48 | 0.5 | 0.45 | 0.51284 | 0.45 | 0.47 | 0.491 | 0.21 | 0.47 | 0.51 | 0.44 | 0.5 |  |  | 1
| 1 | 2 | 0.466 | 0.47 | 0.4696 | 0.54 | 0.48 | 0.48 | 0.5 | 0.5 | 0.44 | 0.50351 | 0.47 | 0.46 | 0.48 | 0.28 | 0.47 | 0.5 | 0.45 | 0.5 |  |  | #N/A
| 1 | 3 | 0.452 | 0.45 | 0.4663 | 0.54 | 0.48 | 0.51 | 0.47 | 0.51 | 0.42 | 0.50506 | 0.41 | 0.49 | 0.48 | 0.22 | 0.46 | 0.5 | 0.44 | 0.49 |  |  | 16
| 1 | 4 | 0.467 | 0.44 | 0.4599 | 0.54 | 0.5 | 0.51 | 0.48 | 0.5 | 0.43 | 0.52238 | 0.42 | 0.46 | 0.484 | 0.22 | 0.46 | 0.51 | 0.45 | 0.5 |  |  | 0.47464
| 1 | 5 | 0.464 | 0.46 | 0.4632 | 0.54 | 0.5 | 0.45 | 0.48 | 0.5 | 0.46 | 0.49966 | 0.46 | 0.46 | 0.492 | 0.2 | 0.46 | 0.5 | 0.48 | 0.5 |  |  | 13
| 1 | 6 | 0.469 | 0.46 | 0.4668 | 0.54 | 0.51 | 0.47 | 0.47 | 0.5 | 0.47 | 0.499 | 0.46 | 0.46 | 0.48 | 0.17 | 0.45 | 0.51 | 0.45 | 0.5
| Mean |  | 0.46283 | 0.46 | 0.4652 | 0.54167 | 0.485 | 0.485 | 0.48 | 0.50167 | 0.445 | 0.50707 | 0.445 | 0.46667 | 0.4845 | 0.21667 | 0.46167 | 0.505 | 0.45167 | 0.49833
| Median |  | 0.465 | 0.46 | 0.46585 | 0.54 | 0.49 | 0.485 | 0.48 | 0.5 | 0.445 | 0.50428 | 0.455 | 0.46 | 0.482 | 0.215 | 0.46 | 0.505 | 0.45 | 0.5
| Std Dev. |  | 0.0063061 | 0.014142 | 0.0033245 | 0.0040825 | 0.0251 | 0.023452 | 0.010954 | 0.0040825 | 0.018708 | 0.008993 | 0.02429 | 0.012111 | 0.005648 | 0.036148 | 0.0075277 | 0.0054772 | 0.01472 | 0.0040825
| Rel.Std.Dev. |  | 1.362% | 3.074% | 0.7146% | 0.7537% | 5.175% | 4.835% | 2.282% | 0.8138% | 4.204% | 1.773% | 5.458% | 2.595% | 1.166% | 16.68% | 1.631% | 1.085% | 3.259% | 0.8192%
| PDM3 |  | -2.488% | -3.085% | -1.989% | 14.12% | 2.182% | 2.182% | 1.129% | 5.694% | -6.245% | 6.833% | -6.245% | -1.68% | 2.077% | -54.35% | -2.734% | 6.396% | -4.841% | 4.991%
| Z-Score (Absolute) |  | 0.34 | 0.44 | 0.26 | 2.43 | 0.44 | 0.44 | 0.26 | 1.02 | 0.97 | 1.21 | 0.97 | 0.37 | 0.42 | 9 | 0.38 | 1.14 | 0.94 | 0.91
| Analytical results for Cd in OREAS 922 (Certified Value 0.29 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.33 | 0.41 | 0.24 | 0.3 | 0.3 | 0.26 | 0.25 | < 0.5 | < 0.4 | 0.27107 | 0.35 | < 0.5 | 0.315 | < 4 | 0.3 | 0.3 | 0.29 | 0.27 |  |  | 1
| 1 | 2 | 0.31 | 0.42 | 0.24 | 0.3 | 0.3 | 0.27 | 0.22 | < 0.5 | < 0.4 | 0.27748 | 0.33 | < 0.5 | 0.353 | < 4 | 0.3 | 0.4 | 0.29 | 0.24 |  |  | #N/A
| 1 | 3 | 0.33 | 0.45 | 0.23 | 0.3 | 0.3 | 0.28 | 0.25 | < 0.5 | < 0.4 | 0.28639 | 0.32 | < 0.5 | 0.303 | < 4 | 0.3 | < 0.2 | 0.28 | 0.23 |  |  | 16
| 1 | 4 | 0.34 | 0.44 | 0.22 | 0.3 | 0.3 | 0.27 | 0.26 | < 0.5 | < 0.4 | 0.29522 | 0.34 | < 0.5 | 0.345 | < 4 | 0.3 | 0.3 | 0.27 | 0.28 |  |  | 0.29074
| 1 | 5 | 0.33 | 0.42 | 0.26 | 0.3 | 0.3 | 0.27 | 0.26 | < 0.5 | < 0.4 | 0.30033 | 0.37 | < 0.5 | 0.334 | < 4 | 0.3 | < 0.2 | 0.3 | 0.2 |  |  | 14
| 1 | 6 | 0.33 | 0.41 | 0.24 | 0.3 | 0.3 | 0.25 | 0.29 | < 0.5 | < 0.4 | 0.2859 | 0.42 | < 0.5 | 0.307 | < 4 | 0.3 | 0.3 | 0.28 | 0.27
| Mean |  | 0.32833 | 0.425 | 0.23833 | 0.3 | 0.3 | 0.26667 | 0.255 |  |  | 0.28606 | 0.355 |  | 0.32617 |  | 0.3 | 0.325 | 0.285 | 0.24833
| Median |  | 0.33 | 0.42 | 0.24 | 0.3 | 0.3 | 0.27 | 0.255 |  |  | 0.28614 | 0.345 |  | 0.3245 |  | 0.3 | 0.3 | 0.285 | 0.255
| Std Dev. |  | 0.0098319 | 0.016432 | 0.013292 | 0 | 0 | 0.010328 | 0.022583 |  |  | 0.010824 | 0.036194 |  | 0.020808 |  | 0 | 0.05 | 0.010488 | 0.030605
| Rel.Std.Dev. |  | 2.994% | 3.866% | 5.577% | 0% | 0% | 3.873% | 8.856% |  |  | 3.784% | 10.2% |  | 6.38% |  | 0% | 15.38% | 3.68% | 12.32%
| PDM3 |  | 12.93% | 46.18% | -18.02% | 3.187% | 3.187% | -8.279% | -12.29% |  |  | -1.607% | 22.1% |  | 12.19% |  | 3.187% | 11.79% | -1.973% | -14.58%
| Z-Score (Absolute) |  | 0.74 | 2.9 | 1.43 | 0 | 0 | 0.77 | 1.04 | NA | NA | 0.32 | 0.97 | NA | 0.61 | NA | 0 | 0 | 0.35 | 1.2
| Analytical results for Ce in OREAS 922 (Certified Value 86 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [6 empty] | 0
1
| 1 | 1 | 112 | 85.36 | 86.9 | 80.1 | 87.9 | 89.8 | 85.8 | 91.69 | 83.3 | 79.898 | 85.2 | 93.1 | 82.3 | 84 | [6 empty] | 1
| 1 | 2 | 112 | 86.43 | 88.8 | 84.1 | 88.5 | 89.2 | 84.5 | 91.7439 | 81.16 | 81.781 | 88.2 | 95 | 82.3 | 88.3 | [6 empty] | #N/A
| 1 | 3 | 110 | 87.61 | 86.6 | 84.7 | 90.9 | 88.5 | 85.4 | 93.399 | 84.91 | 80.248 | 84.3 | 94.5 | 81 | 82.4 | [6 empty] | 16
| 1 | 4 | 110 | 87.85 | 85.7 | 82.3 | 90.3 | 89 | 85.7 | 92.9072 | 86.01 | 83.371 | 84.3 | 93.8 | 83.7 | 83.6 | [6 empty] | 85.7246
| 1 | 5 | 110 | 89.72 | 87 | 81.3 | 81.9 | 84.1 | 84.6 | 92.1583 | 85.68 | 81.531 | 83.5 | 95.2 | 87.3 | 86.1 | [6 empty] | 15
| 1 | 6 | 111 | 85.69 | 88 | 82.8 | 86.1 | 88.1 | 84.4 | 92.9394 | 85.05 | 80.975 | 82.8 | 94.2 | 82.9 | 85.6
| Mean |  | 110.8333 | 87.11 | 87.1667 | 82.55 | 87.6 | 88.1167 | 85.0667 | 92.473 | 84.3517 | 81.3007 | 84.7167 | 94.3 | 83.25 | 85
| Median |  | 110.5 | 87.02 | 86.95 | 82.55 | 88.2 | 88.75 | 85 | 92.5328 | 84.98 | 81.253 | 84.3 | 94.35 | 82.6 | 84.8
| Std Dev. |  | 0.98319 | 1.6218 | 1.0893 | 1.7155 | 3.2809 | 2.0527 | 0.6377 | 0.70815 | 1.8225 | 1.2451 | 1.8904 | 0.77974 | 2.1723 | 2.1062
| Rel.Std.Dev. |  | 0.8871% | 1.862% | 1.25% | 2.078% | 3.745% | 2.33% | 0.7497% | 0.7658% | 2.161% | 1.532% | 2.231% | 0.8269% | 2.609% | 2.478%
| PDM3 |  | 29.29% | 1.616% | 1.682% | -3.703% | 2.188% | 2.79% | -0.7675% | 7.872% | -1.602% | -5.161% | -1.176% | 10% | -2.887% | -0.8453%
| Z-Score (Absolute) |  | 7.3 | 0.3 | 0.32 | 1.04 | 0.45 | 0.84 | 0.3 | 1.88 | 0.51 | 1.41 | 0.4 | 2.42 | 1.08 | 0.32
| Analytical results for Co in OREAS 922 (Certified Value 20.4 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 20 | 23.3 | 20.5 | 21.2 | 20.4 | 19.8 | 19.2 | 21.6 | 25 | 19.8 | 17 | 20.9534 | 18 | 23 | 18.54 | 24 | 21 | 21 | 20 | 19.3 | 1
| 1 | 2 | 20 | 23.2 | 20.3 | 21.7 | 21 | 19.4 | 19.1 | 21.8 | 25 | 19.9 | 18 | 19.9888 | 20 | 21 | 18.68 | 25 | 23 | 21 | 20.5 | 20.4 | #N/A
| 1 | 3 | 21 | 23.1 | 20.2 | 21 | 21.5 | 20.5 | 18.4 | 21.8 | 25 | 19.9 | 17 | 20.0203 | 21 | 21 | 18.02 | 24 | 21 | 21 | 19.6 | 18.4 | 16
| 1 | 4 | 20 | 23 | 20.1 | 21.3 | 21 | 20.1 | 18.5 | 22.1 | 25 | 20 | 18 | 20.0694 | 20 | 21 | 18.51 | 25 | 21 | 20 | 20.4 | 18.9 | 20.3889
| 1 | 5 | 21 | 23 | 20 | 21.6 | 20.6 | 17.9 | 18.6 | 22.2 | 20 | 20 | 18 | 19.887 | 20 | 20 | 18.35 | 24 | 21 | 21 | 21.5 | 18.9 | 16
| 1 | 6 | 21 | 23 | 20 | 21.4 | 20.8 | 18.9 | 18.3 | 21.6 | 25 | 20.1 | 18 | 19.752 | 20 | 21 | 18.22 | 26 | 21 | 21 | 20.7 | 18.9
| Mean |  | 20.5 | 23.1 | 20.1833 | 21.3667 | 20.8833 | 19.4333 | 18.6833 | 21.85 | 24.1667 | 19.95 | 17.6667 | 20.1118 | 19.8333 | 21.1667 | 18.3867 | 24.6667 | 21.3333 | 20.8333 | 20.45 | 19.1333
| Median |  | 20.5 | 23.05 | 20.15 | 21.35 | 20.9 | 19.6 | 18.55 | 21.8 | 25 | 19.95 | 18 | 20.0045 | 20 | 21 | 18.43 | 24.5 | 21 | 21 | 20.45 | 18.9
| Std Dev. |  | 0.54772 | 0.12649 | 0.19408 | 0.2582 | 0.38166 | 0.9331 | 0.37639 | 0.251 | 2.0412 | 0.10488 | 0.5164 | 0.42747 | 0.98319 | 0.98319 | 0.24014 | 0.8165 | 0.8165 | 0.40825 | 0.6473 | 0.68313
| Rel.Std.Dev. |  | 2.672% | 0.5476% | 0.9616% | 1.208% | 1.828% | 4.802% | 2.015% | 1.149% | 8.447% | 0.5257% | 2.923% | 2.125% | 4.957% | 4.645% | 1.306% | 3.31% | 3.827% | 1.96% | 3.165% | 3.57%
| PDM3 |  | 0.5448% | 13.3% | -1.008% | 4.795% | 2.425% | -4.687% | -8.365% | 7.166% | 18.53% | -2.153% | -13.35% | -1.359% | -2.725% | 3.814% | -9.82% | 20.98% | 4.632% | 2.18% | 0.2995% | -6.158%
| Z-Score (Absolute) |  | 0.02 | 2.5 | 0.28 | 0.85 | 0.39 | 0.99 | 1.71 | 1.31 | 4.31 | 0.5 | 2.67 | 0.35 | 0.26 | 0.31 | 1.99 | 3.99 | 0.5 | 0.34 | 0.02 | 1.52
| Analytical results for Cr in OREAS 922 (Certified Value 75 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 82 | 66 | 63 | 73.4 | 71.2 | 78 | 71 | 60 | 69 | 89.094 | 77 | 90 | 67 | 95 | 86 | 66 | 77 |  |  |  | 1
| 1 | 2 | 86 | 67 | 61 | 75.3 | 79.4 | 78 | 74 | 70 | 70 | 86.568 | 70 | 96 | 77 | 74 | 87 | 70 | 77 |  |  |  | #N/A
| 1 | 3 | 88 | 67 | 62 | 75.4 | 75.4 | 82 | 71 | 70 | 64 | 88.3843 | 71 | 93 | 70 | 73 | 88 | 64 | 76 |  |  |  | 16
| 1 | 4 | 85 | 62 | 61 | 79.3 | 77.2 | 79 | 71 | 60 | 69 | 88.2762 | 76 | 97 | 70 | 80 | 85 | 67 | 76 |  |  |  | 75.0614
| 1 | 5 | 84 | 75 | 60 | 75.9 | 82.9 | 72 | 73 | 50 | 70 | 86.5516 | 81 | 90 | 70 | 77 | 85 | 71 | 77 |  |  |  | 17
| 1 | 6 | 82 | 55 | 62 | 76.8 | 111 | 77 | 70 | 50 | 71 | 87.9703 | 80 | 93 | 67 | 80 | 84 | 67 | 86
| Mean |  | 84.5 | 65.3333 | 61.5 | 76.0167 | 82.85 | 77.6667 | 71.6667 | 60 | 68.8333 | 87.8074 | 75.8333 | 93.1667 | 70.1667 | 79.8333 | 85.8333 | 67.5 | 78.1667
| Median |  | 84.5 | 66.5 | 61.5 | 75.65 | 78.3 | 78 | 71 | 60 | 69.5 | 88.1233 | 76.5 | 93 | 70 | 78.5 | 85.5 | 67 | 77
| Std Dev. |  | 2.3452 | 6.5929 | 1.0488 | 1.957 | 14.3351 | 3.266 | 1.5055 | 8.9443 | 2.4833 | 1.0342 | 4.535 | 2.9269 | 3.656 | 7.9854 | 1.472 | 2.5884 | 3.8687
| Rel.Std.Dev. |  | 2.775% | 10.09% | 1.705% | 2.574% | 17.3% | 4.205% | 2.101% | 14.91% | 3.608% | 1.178% | 5.98% | 3.142% | 5.211% | 10% | 1.715% | 3.835% | 4.949%
| PDM3 |  | 12.57% | -12.96% | -18.07% | 1.273% | 10.38% | 3.471% | -4.523% | -20.07% | -8.297% | 16.98% | 1.028% | 24.12% | -6.521% | 6.357% | 14.35% | -10.07% | 4.137%
| Z-Score (Absolute) |  | 0.79 | 1 | 1.36 | 0 | 0.11 | 0.15 | 0.41 | 1.5 | 0.58 | 1.1 | 0.02 | 1.6 | 0.67 | 0.07 | 0.92 | 0.8 | 0.05
| Analytical results for Cs in OREAS 922 (Certified Value 7.27 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 7.7 | 8.64 | 7.31 | 7.4 | 6.55 | 7.33 | 7.08 | 7.1 | 7.0214 | 7.86 | 6.574 | 6.86 | 7.3 | 8 | 7.2 | [5 empty] | 1
| 1 | 2 | 7.8 | 8.6 | 7.28 | 7.49 | 6.64 | 7.37 | 7.2 | 7.2 | 6.9479 | 7.65 | 6.718 | 6.91 | 7.5 | 8 | 7.4 | [5 empty] | #N/A
| 1 | 3 | 7.9 | 8.53 | 7.32 | 7.4 | 6.76 | 7.44 | 7.24 | 7 | 6.9305 | 7.91 | 6.557 | 6.81 | 7.4 | 7 | 7.11 | [5 empty] | 16
| 1 | 4 | 7.7 | 8.55 | 7.26 | 7.35 | 6.63 | 7.61 | 7.04 | 7 | 7.1146 | 8.15 | 6.828 | 6.76 | 7.4 | 8 | 7.13 | [5 empty] | 7.271
| 1 | 5 | 7.9 | 8.52 | 7.47 | 7.42 | 6.66 | 6.79 | 7.11 | 7 | 6.9801 | 8.06 | 6.804 | 6.81 | 7.2 | 8 | 7.32 | [5 empty] | 18
| 1 | 6 | 7.7 | 8.57 | 7.52 | 7.46 | 6.71 | 7.03 | 7.01 | 7.1 | 6.9316 | 8.37 | 6.526 | 6.7 | 7.8 | 8 | 7.24
| Mean |  | 7.7833 | 8.5683 | 7.36 | 7.42 | 6.6583 | 7.2617 | 7.1133 | 7.0667 | 6.9877 | 8 | 6.6678 | 6.8083 | 7.4333 | 7.8333 | 7.2333
| Median |  | 7.75 | 8.56 | 7.315 | 7.41 | 6.65 | 7.35 | 7.095 | 7.05 | 6.964 | 7.985 | 6.646 | 6.81 | 7.4 | 8 | 7.22
| Std Dev. |  | 0.098319 | 0.04535 | 0.10789 | 0.049396 | 0.071949 | 0.29856 | 0.090259 | 0.08165 | 0.071143 | 0.25028 | 0.1326 | 0.073598 | 0.20656 | 0.40825 | 0.11165
| Rel.Std.Dev. |  | 1.263% | 0.5293% | 1.466% | 0.6657% | 1.081% | 4.111% | 1.269% | 1.155% | 1.018% | 3.128% | 1.989% | 1.081% | 2.779% | 5.212% | 1.544%
| PDM3 |  | 7.046% | 17.84% | 1.224% | 2.049% | -8.426% | -0.1282% | -2.168% | -2.81% | -3.896% | 10.03% | -8.295% | -6.363% | 2.233% | 7.734% | -0.5179%
| Z-Score (Absolute) |  | 1.28 | 3.22 | 0.24 | 0.39 | 1.48 | 0 | 0.37 | 0.48 | 0.67 | 1.82 | 1.46 | 1.12 | 0.42 | 1.82 | 0.07
| Analytical results for Cu in OREAS 922 (Certified Value 2122 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 2270 | 2136 | 2143 | 2170 | 1890 | 2170 | 2100 | 2190 | 2102 | 2079 | 2214.6969 | 2091 | 2386 | 2018.18 | 2034 | 2100 | 2073 | 2040 | 2170 |  | 1
| 1 | 2 | 2270 | 2125 | 2139 | 2220 | 1990 | 2140 | 2200 | 2190 | 2103 | 2153 | 2189.78 | 2185 | 2378 | 1923.02 | 2185 | 2100 | 2154 | 2080 | 2220 |  | #N/A
| 1 | 3 | 2270 | 2074 | 2195 | 2160 | 2090 | 2270 | 2130 | 2190 | 2061 | 1903 | 2200.08 | 1959 | 2388 | 1945.41 | 2083 | 2110 | 2088 | 2030 | 2160 |  | 16
| 1 | 4 | 2250 | 2080 | 2147 | 2150 | 2180 | 2240 | 2110 | 2240 | 2157 | 2055 | 2215.53 | 1955 | 2307 | 2016.32 | 2062 | 2090 | 2064 | 2090 | 2140 |  | 2121.9946
| 1 | 5 | 2210 | 2058 | 2127 | 2190 | 2090 | 1980 | 2150 | 2210 | 2009 | 2175 | 2208.0677 | 2127 | 2271 | 1957.16 | 1967 | 2150 | 2024 | 2210 | 2210 |  | 19
| 1 | 6 | 2220 | 2078 | 2201 | 2170 | 2190 | 2050 | 2110 | 2220 | 2129 | 2212 | 2181.2066 | 2155 | 2345 | 1962.97 | 2211 | 2060 | 2121 | 2080 | 2200
| Mean |  | 2248.3333 | 2091.8333 | 2158.6667 | 2176.6667 | 2071.6667 | 2141.6667 | 2133.3333 | 2206.6667 | 2093.5 | 2096.1667 | 2201.5602 | 2078.6667 | 2345.8333 | 1970.51 | 2090.3333 | 2101.6667 | 2087.3333 | 2088.3333 | 2183.3333
| Median |  | 2260 | 2079 | 2145 | 2170 | 2090 | 2155 | 2120 | 2200 | 2102.5 | 2116 | 2204.0738 | 2109 | 2361.5 | 1960.065 | 2072.5 | 2100 | 2080.5 | 2080 | 2185
| Std Dev. |  | 27.1416 | 31.1282 | 31.2517 | 25.0333 | 114.9638 | 110.8903 | 37.238 | 20.6559 | 52.2523 | 111.5032 | 13.8947 | 99.2324 | 48.0309 | 38.7075 | 92.4806 | 29.2689 | 45.4738 | 64.3169 | 31.4113
| Rel.Std.Dev. |  | 1.207% | 1.488% | 1.448% | 1.15% | 5.549% | 5.178% | 1.746% | 0.9361% | 2.496% | 5.319% | 0.6311% | 4.774% | 2.047% | 1.964% | 4.424% | 1.393% | 2.179% | 3.08% | 1.439%
| PDM3 |  | 5.954% | -1.421% | 1.728% | 2.576% | -2.372% | 0.9271% | 0.5343% | 3.99% | -1.343% | -1.217% | 3.75% | -2.042% | 10.55% | -7.139% | -1.492% | -0.958% | -1.633% | -1.586% | 2.891%
| Z-Score (Absolute) |  | 2.63 | 0.18 | 1.02 | 1.34 | 0.54 | 0.72 | 0.57 | 1.88 | 0.15 | 0.1 | 1.79 | 0.41 | 4.37 | 2.35 | 0.2 | 0 | 0.26 | 0.67 | 1.46
| Analytical results for Dy in OREAS 922 (Certified Value 5.39 ppm) | [21 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 10: 4A*MS |  | 2
3
| 1 | 1 | 01: 6.81; 02: 5.25; 03: 6; 04: 5.1; 05: 5.15; 06: 5.624; 07: 5.36; 09: 5.3; 10: 5.3 |  | 1
| 1 | 2 | 01: 6.79; 02: 5.27; 03: 6.2; 04: 5.3; 05: 5.3; 06: 5.5722; 07: 5.35; 09: 5.5; 10: 5.5 |  | #N/A
| 1 | 3 | 01: 6.74; 02: 5.32; 03: 6; 04: 5.4; 05: 5.2; 06: 5.5523; 07: 5.78; 09: 5.4; 10: 5.3 |  | 16
| 1 | 4 | 01: 6.73; 02: 5.27; 03: 5.9; 04: 5.3; 05: 5.3; 06: 5.4745; 07: 5.75; 09: 5.3; 10: 5.3 |  | 5.3851
| 1 | 5 | 01: 6.73; 02: 5.36; 03: 6; 04: 5.3; 05: 5.15; 06: 5.5664; 07: 5.8; 09: 5.5; 10: 5.2 |  | 20
| 1 | 6 | 01: 6.77; 02: 5.31; 03: 6.1; 04: 5.4; 05: 5.3; 06: 5.4367; 07: 5.63; 09: 5.3; 10: 5.4
| Mean |  | 01: 6.7617; 02: 5.2967; 03: 6.0333; 04: 5.3; 05: 5.2333; 06: 5.5377; 07: 5.6117; 09: 5.3833; 10: 5.3333
| Median |  | 01: 6.755; 02: 5.29; 03: 6; 04: 5.3; 05: 5.25; 06: 5.5593; 07: 5.69; 09: 5.35; 10: 5.3
| Std Dev. |  | 01: 0.033714; 02: 0.040825; 03: 0.10328; 04: 0.10954; 05: 0.075277; 06: 0.069086; 07: 0.2074; 09: 0.098319; 10: 0.10328
| Rel.Std.Dev. |  | 01: 0.4986%; 02: 0.7708%; 03: 1.712%; 04: 2.067%; 05: 1.438%; 06: 1.248%; 07: 3.696%; 09: 1.826%; 10: 1.936%
| PDM3 |  | 01: 25.56%; 02: -1.643%; 03: 12.04%; 04: -1.581%; 05: -2.819%; 06: 2.832%; 07: 4.206%; 09: -0.03357%; 10: -0.9621%
| Z-Score (Absolute) |  | 01: 6.2; 02: 0.39; 03: 2.77; 04: 0.37; 05: 0.67; 06: 0.69; 07: 1.03; 09: 0; 10: 0.22
| Analytical results for Er in OREAS 922 (Certified Value 3.11 ppm) | [21 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 10: 4A*MS |  | 2
3
| 1 | 1 | 01: 4.02; 02: 3.08; 03: 3.4; 04: 2.8; 05: 3.1; 06: 3.2642; 07: 3.06; 09: 3; 10: 3 |  | 1
| 1 | 2 | 01: 4; 02: 3.02; 03: 3.4; 04: 3; 05: 3.05; 06: 3.2675; 07: 3.04; 09: 3.1; 10: 2.9 |  | #N/A
| 1 | 3 | 01: 3.96; 02: 3.14; 03: 3.4; 04: 3; 05: 3.15; 06: 3.1996; 07: 3.2; 09: 3; 10: 3.1 |  | 16
| 1 | 4 | 01: 3.97; 02: 3.09; 03: 3.4; 04: 2.9; 05: 3.2; 06: 3.1371; 07: 3.19; 09: 3; 10: 3.1 |  | 3.1056
| 1 | 5 | 01: 3.98; 02: 3.13; 03: 3.4; 04: 2.9; 05: 2.95; 06: 3.1583; 07: 3.16; 09: 3; 10: 3 |  | 21
| 1 | 6 | 01: 3.99; 02: 3.02; 03: 3.4; 04: 2.9; 05: 3.15; 06: 3.1305; 07: 3.18; 09: 3; 10: 3
| Mean |  | 01: 3.9867; 02: 3.08; 03: 3.4; 04: 2.9167; 05: 3.1; 06: 3.1929; 07: 3.1383; 09: 3.0167; 10: 3.0167
| Median |  | 01: 3.985; 02: 3.085; 03: 3.4; 04: 2.9; 05: 3.125; 06: 3.179; 07: 3.17; 09: 3; 10: 3
| Std Dev. |  | 01: 0.021602; 02: 0.051769; 03: 0; 04: 0.075277; 05: 0.089443; 06: 0.061475; 07: 0.069976; 09: 0.040825; 10: 0.075277
| Rel.Std.Dev. |  | 01: 0.5419%; 02: 1.681%; 03: 0%; 04: 2.581%; 05: 2.885%; 06: 1.925%; 07: 2.23%; 09: 1.353%; 10: 2.495%
| PDM3 |  | 01: 28.37%; 02: -0.8233%; 03: 9.481%; 04: -6.083%; 05: -0.1792%; 06: 2.811%; 07: 1.055%; 09: -2.863%; 10: -2.863%
| Z-Score (Absolute) |  | 01: 6.44; 02: 0.15; 03: 2.18; 04: 1.33; 05: 0; 06: 0.67; 07: 0.28; 09: 0.73; 10: 0.61
| Analytical results for Eu in OREAS 922 (Certified Value 1.39 ppm) | [21 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 10: 4A*MS |  | 2
3
| 1 | 1 | 01: 1.86; 02: 1.32; 03: 1.48; 04: 1.24; 05: 1.45; 06: 1.4524; 07: 1.33; 09: 1.31; 10: 1.5 |  | 1
| 1 | 2 | 01: 1.86; 02: 1.35; 03: 1.48; 04: 1.25; 05: 1.4; 06: 1.5005; 07: 1.36; 09: 1.34; 10: 1.5 |  | #N/A
| 1 | 3 | 01: 1.85; 02: 1.35; 03: 1.47; 04: 1.25; 05: 1.4; 06: 1.4812; 07: 1.41; 09: 1.32; 10: 1.4 |  | 16
| 1 | 4 | 01: 1.84; 02: 1.34; 03: 1.45; 04: 1.23; 05: 1.4; 06: 1.4219; 07: 1.44; 09: 1.3; 10: 1.5 |  | 1.3893
| 1 | 5 | 01: 1.84; 02: 1.31; 03: 1.46; 04: 1.23; 05: 1.4; 06: 1.4397; 07: 1.41; 09: 1.29; 10: 1.5 |  | 22
| 1 | 6 | 01: 1.85; 02: 1.33; 03: 1.51; 04: 1.25; 05: 1.5; 06: 1.4124; 07: 1.44; 09: 1.27; 10: 1.5
| Mean |  | 01: 1.85; 02: 1.3333; 03: 1.475; 04: 1.2417; 05: 1.425; 06: 1.4513; 07: 1.3983; 09: 1.305; 10: 1.4833
| Median |  | 01: 1.85; 02: 1.335; 03: 1.475; 04: 1.245; 05: 1.4; 06: 1.446; 07: 1.41; 09: 1.305; 10: 1.5
| Std Dev. |  | 01: 0.0089443; 02: 0.01633; 03: 0.020736; 04: 0.0098319; 05: 0.041833; 06: 0.034144; 07: 0.04446; 09: 0.02429; 10: 0.040825
| Rel.Std.Dev. |  | 01: 0.4835%; 02: 1.225%; 03: 1.406%; 04: 0.7918%; 05: 2.936%; 06: 2.353%; 07: 3.179%; 09: 1.861%; 10: 2.752%
| PDM3 |  | 01: 33.16%; 02: -4.031%; 03: 6.166%; 04: -10.63%; 05: 2.567%; 06: 4.463%; 07: 0.6477%; 09: -6.07%; 10: 6.766%
| Z-Score (Absolute) |  | 01: 3.87; 02: 0.67; 03: 0.57; 04: 1.48; 05: 0; 06: 0.36; 07: 0.1; 09: 0.92; 10: 0.79
| Analytical results for Fe in OREAS 922 (Certified Value 5.53 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 5.8 | 5.41 | 5.58 | 5.91 | 4.88 | 5.47 | 5.29 | 5.82 | 5.98 | 5.54 | 5.4998 | 6.11 | 5.078 | 4.73 | 5.64 | 5.87 | 5.13 | 5.4 |  |  | 1
| 1 | 2 | 5.82 | 5.52 | 5.69 | 5.97 | 5.31 | 5.41 | 5.58 | 5.9 | 5.93 | 5.35 | 5.56 | 6.03 | 4.965 | 5.13 | 5.73 | 5.83 | 5.19 | 5.64 |  |  | #N/A
| 1 | 3 | 5.74 | 5.29 | 5.59 | 5.81 | 5.4 | 5.62 | 5.25 | 5.82 | 5.88 | 5.14 | 5.54 | 6.15 | 4.888 | 4.83 | 5.63 | 5.78 | 5.1 | 5.53 |  |  | 16
| 1 | 4 | 5.77 | 5.3 | 5.52 | 5.86 | 5.57 | 5.7 | 5.34 | 5.94 | 5.9 | 5.34 | 5.54 | 5.96 | 5.049 | 4.79 | 5.67 | 5.91 | 5.25 | 5.44 |  |  | 5.5342
| 1 | 5 | 5.73 | 5.35 | 5.56 | 5.9 | 5.58 | 5 | 5.36 | 5.9 | 5.76 | 5.56 | 5.53 | 5.8 | 5.079 | 4.59 | 5.65 | 5.82 | 5.55 | 5.52 |  |  | 23
| 1 | 6 | 5.78 | 5.42 | 5.69 | 6.03 | 5.65 | 5.21 | 5.34 | 5.89 | 5.99 | 5.61 | 5.52 | 6.06 | 5.044 | 5.12 | 5.49 | 5.8 | 5.24 | 5.52
| Mean |  | 5.7733 | 5.3817 | 5.605 | 5.9133 | 5.3983 | 5.4017 | 5.36 | 5.8783 | 5.9067 | 5.4233 | 5.5316 | 6.0183 | 5.0172 | 4.865 | 5.635 | 5.835 | 5.2433 | 5.5083
| Median |  | 5.775 | 5.38 | 5.585 | 5.905 | 5.485 | 5.44 | 5.34 | 5.895 | 5.915 | 5.445 | 5.535 | 6.045 | 5.0465 | 4.81 | 5.645 | 5.825 | 5.215 | 5.52
| Std Dev. |  | 0.034448 | 0.086583 | 0.070071 | 0.078145 | 0.28365 | 0.26057 | 0.11507 | 0.048339 | 0.083825 | 0.17874 | 0.020461 | 0.12545 | 0.075714 | 0.21723 | 0.079436 | 0.047645 | 0.16145 | 0.083046
| Rel.Std.Dev. |  | 0.5967% | 1.609% | 1.25% | 1.322% | 5.254% | 4.824% | 2.147% | 0.8223% | 1.419% | 3.296% | 0.3699% | 2.084% | 1.509% | 4.465% | 1.41% | 0.8165% | 3.079% | 1.508%
| PDM3 |  | 4.321% | -2.757% | 1.279% | 6.85% | -2.455% | -2.395% | -3.148% | 6.218% | 6.73% | -2.004% | -0.04663% | 8.748% | -9.343% | -12.09% | 1.821% | 5.435% | -5.256% | -0.4678%
| Z-Score (Absolute) |  | 0.75 | 0.41 | 0.25 | 1.16 | 0.36 | 0.35 | 0.6 | 1.06 | 1.14 | 0.29 | 0.03 | 1.47 | 1.41 | 1.93 | 0.34 | 0.93 | 1 | 0.03
| Analytical results for Ga in OREAS 922 (Certified Value 20.4 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 21.4 | 23.3 | 19.51 | 19 | 21.1 | 20 | 19.45 | 20.6 | 19.7358 | 21.76 | 23 | 20.8 | 19.5 | 21.6 | 20.6 | 19.7 |  |  |  |  | 1
| 1 | 2 | 22 | 23.2 | 19.46 | 19.6 | 21.9 | 20 | 20.1 | 20.6 | 19.7772 | 22.21 | 22 | 21.42 | 20.3 | 21.9 | 20.8 | 20.2 |  |  |  |  | #N/A
| 1 | 3 | 21 | 23.1 | 19.66 | 18.8 | 21.4 | 20.3 | 19.6 | 19.2 | 19.3616 | 22.58 | 22 | 20.62 | 20.7 | 21.5 | 20.3 | 19.1 |  |  |  |  | 16
| 1 | 4 | 21.1 | 23.1 | 19.54 | 19.2 | 20.9 | 20.6 | 19.7 | 19.2 | 19.7221 | 23.32 | 23 | 21.15 | 19.9 | 21.7 | 20.8 | 19.4 |  |  |  |  | 20.3923
| 1 | 5 | 22 | 23 | 19.71 | 19.3 | 20.3 | 18.25 | 19.6 | 20.8 | 19.6093 | 23.4 | 21 | 20.97 | 19.1 | 21.6 | 22 | 19.7 |  |  |  |  | 24
| 1 | 6 | 21.5 | 23 | 20.35 | 20.1 | 20.6 | 19.15 | 19.4 | 20 | 19.6517 | 24.57 | 22 | 20.53 | 19 | 21.9 | 21 | 20.4
| Mean |  | 21.5 | 23.1167 | 19.705 | 19.3333 | 21.0333 | 19.7167 | 19.6417 | 20.0667 | 19.643 | 22.9733 | 22.1667 | 20.915 | 19.75 | 21.7 | 20.9167 | 19.75
| Median |  | 21.45 | 23.1 | 19.6 | 19.25 | 21 | 20 | 19.6 | 20.3 | 19.6869 | 22.95 | 22 | 20.885 | 19.7 | 21.65 | 20.8 | 19.7
| Std Dev. |  | 0.42895 | 0.1169 | 0.32965 | 0.46332 | 0.57155 | 0.86641 | 0.24983 | 0.72296 | 0.15052 | 1.0061 | 0.75277 | 0.33519 | 0.67454 | 0.16733 | 0.58109 | 0.48477
| Rel.Std.Dev. |  | 1.995% | 0.5057% | 1.673% | 2.396% | 2.717% | 4.394% | 1.272% | 3.603% | 0.7663% | 4.379% | 3.396% | 1.603% | 3.415% | 0.7711% | 2.778% | 2.455%
| PDM3 |  | 5.432% | 13.36% | -3.37% | -5.193% | 3.143% | -3.313% | -3.681% | -1.597% | -3.675% | 12.66% | 8.701% | 2.563% | -3.15% | 6.413% | 2.571% | -3.15%
| Z-Score (Absolute) |  | 1.02 | 2.49 | 0.73 | 0.96 | 0.59 | 0.61 | 0.67 | 0.29 | 0.67 | 2.36 | 1.62 | 0.48 | 0.58 | 1.2 | 0.29 | 0.58
| Analytical results for Gd in OREAS 922 (Certified Value 6.18 ppm) | [21 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 13: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 13: 19 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 01: 7.76; 02: 6.18; 03: 6.4; 04: 5.5; 05: 6; 06: 6.1701; 07: 6; 09: 5.8; 13: 6.3 |  |  |  |  | 1
| 1 | 2 | 01: 7.75; 02: 6.25; 03: 6.5; 04: 5.4; 05: 6.2; 06: 6.2794; 07: 5.71; 09: 5.8; 13: 6.6 |  |  |  |  | #N/A
| 1 | 3 | 01: 7.67; 02: 6.09; 03: 6.3; 04: 5.6; 05: 6.2; 06: 6.1578; 07: 6.09; 09: 5.8; 13: 6.5 |  |  |  |  | 16
| 1 | 4 | 01: 7.67; 02: 6.25; 03: 6.3; 04: 5.4; 05: 6.4; 06: 6.163; 07: 6.03; 09: 5.8; 13: 6.5 |  |  |  |  | 6.1756
| 1 | 5 | 01: 7.65; 02: 6.1; 03: 6.3; 04: 5.4; 05: 5.8; 06: 6.3087; 07: 6.17; 09: 5.9; 13: 6.5 |  |  |  |  | 25
| 1 | 6 | 01: 7.72; 02: 5.94; 03: 6.3; 04: 5.5; 05: 6.4; 06: 6.1246; 07: 6.01; 09: 5.8; 13: 6.6
| Mean |  | 01: 7.7033; 02: 6.135; 03: 6.35; 04: 5.4667; 05: 6.1667; 06: 6.2006; 07: 6.0017; 09: 5.8167; 13: 6.5
| Median |  | 01: 7.695; 02: 6.14; 03: 6.3; 04: 5.45; 05: 6.2; 06: 6.1666; 07: 6.02; 09: 5.8; 13: 6.5
| Std Dev. |  | 01: 0.046332; 02: 0.11811; 03: 0.083666; 04: 0.08165; 05: 0.23381; 06: 0.074629; 07: 0.15626; 09: 0.040825; 13: 0.10954
| Rel.Std.Dev. |  | 01: 0.6015%; 02: 1.925%; 03: 1.318%; 04: 1.494%; 05: 3.791%; 06: 1.204%; 07: 2.604%; 09: 0.7019%; 13: 1.685%
| PDM3 |  | 01: 24.74%; 02: -0.6569%; 03: 2.825%; 04: -11.48%; 05: -0.1441%; 06: 0.4057%; 07: -2.816%; 09: -5.812%; 13: 5.254%
| Z-Score (Absolute) |  | 01: 5.65; 02: 0.12; 03: 0.67; 04: 2.57; 05: 0; 06: 0.12; 07: 0.39; 09: 1.29; 13: 1.23
| Analytical results for Ge in OREAS 922 (Certified Value < 2 ppm) | [21 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 13 | 14 | 18 | 19 | 22 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 2 | 1.9 | 0.7 | 0.7 | 0.19 | 0.21 | 1.7783 | 1.3 | 0.5 | < 0.5 | 0.17 | [9 empty] | 1
| 1 | 2 | 2.4 | 1.9 | 1 | 0.7 | 0.18 | 0.19 | 1.7281 | 1.54 | 0.7 | < 0.5 | 0.16 | [9 empty] | 10
| 1 | 3 | 2.5 | 2 | 1.1 | 0.7 | 0.18 | 0.19 | 2.0468 | 1.31 | 0.8 | < 0.5 | 0.19 | [9 empty] | 16
| 1 | 4 | 2.3 | 2 | 0.7 | 0.7 | 0.19 | 0.19 | 1.8226 | 1.58 | 0.6 | < 0.5 | 0.19 | [9 empty] | < 2
| 1 | 5 | 2.6 | 1.8 | 0.8 | 0.6 | 0.18 | 0.19 | 1.889 | 1.63 | 0.7 | < 0.5 | 0.15 | [9 empty] | 26
| 1 | 6 | 2.7 | 1.7 | 0.9 | 0.6 | 0.16 | 0.29 | 1.6592 | 1.77 | 0.5 | < 0.5 | 0.18
| Mean |  | 2.4167 | 1.8833 | 0.86667 | 0.66667 | 0.18 | 0.21 | 1.8207 | 1.5217 | 0.63333 |  | 0.17333
| Median |  | 2.45 | 1.9 | 0.85 | 0.7 | 0.18 | 0.19 | 1.8005 | 1.56 | 0.65 |  | 0.175
| Std Dev. |  | 0.24833 | 0.1169 | 0.1633 | 0.05164 | 0.010954 | 0.04 | 0.13586 | 0.18498 | 0.12111 |  | 0.01633
| Rel.Std.Dev. |  | 10.28% | 6.207% | 18.84% | 7.746% | 6.086% | 19.05% | 7.462% | 12.16% | 19.12% |  | 9.421%
PDM3
| Z-Score (Absolute) |  | 1.89 | 1.28 | 0.11 | 0.11 | 0.67 | 0.65 | 1.2 | 0.86 | 0.15 | NA | 0.68
| Analytical results for Hf in OREAS 922 (Certified Value 3.78 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4 | 4.45 | 3.3 | 3.7 | 3.6 | 4 | 3.7 | 4.2 | 3.783 | 3.6 | 3.72 | 3.2 | 3.9 | 3.74 | 4.2 | [5 empty] | 1
| 1 | 2 | 4.1 | 4.42 | 3.4 | 4.3 | 3.5 | 3.9 | 3.6 | 4.2 | 3.803 | 3.41 | 3.776 | 3.6 | 3.7 | 3.62 | 3.9 | [5 empty] | #N/A
| 1 | 3 | 3.9 | 4.39 | 3.15 | 4.1 | 3.7 | 4 | 3.7 | 4.4 | 3.7249 | 3.59 | 3.65 | 3.7 | 3.8 | 3.66 | 3.8 | [5 empty] | 16
| 1 | 4 | 4.1 | 4.38 | 3.37 | 4.2 | 3.7 | 4 | 3.7 | 4.2 | 3.6388 | 3.56 | 3.772 | 3.2 | 3.8 | 3.82 | 3.8 | [5 empty] | 3.7765
| 1 | 5 | 4.3 | 4.41 | 3.41 | 3.9 | 3.7 | 3.6 | 3.6 | 4.4 | 3.6781 | 3.71 | 3.947 | 3 | 3.8 | 3.81 | 4.2 | [5 empty] | 27
| 1 | 6 | 4.2 | 4.43 | 3.46 | 4 | 3.7 | 3.7 | 3.7 | 4.4 | 3.6759 | 3.48 | 3.741 | 2 | 3.7 | 3.67 | 3.8
| Mean |  | 4.1 | 4.4133 | 3.3483 | 4.0333 | 3.65 | 3.8667 | 3.6667 | 4.3 | 3.7173 | 3.5583 | 3.7677 | 3.1167 | 3.7833 | 3.72 | 3.95
| Median |  | 4.1 | 4.415 | 3.385 | 4.05 | 3.7 | 3.95 | 3.7 | 4.3 | 3.7015 | 3.575 | 3.7565 | 3.2 | 3.8 | 3.705 | 3.85
| Std Dev. |  | 0.14142 | 0.02582 | 0.11053 | 0.21602 | 0.083666 | 0.17512 | 0.05164 | 0.10954 | 0.065028 | 0.10381 | 0.099073 | 0.608 | 0.075277 | 0.083187 | 0.19748
| Rel.Std.Dev. |  | 3.449% | 0.585% | 3.301% | 5.356% | 2.292% | 4.529% | 1.408% | 2.548% | 1.749% | 2.917% | 2.63% | 19.51% | 1.99% | 2.236% | 5%
| PDM3 |  | 8.566% | 16.86% | -11.34% | 6.8% | -3.35% | 2.387% | -2.909% | 13.86% | -1.568% | -5.777% | -0.2345% | -17.47% | 0.1804% | -1.497% | 4.594%
| Z-Score (Absolute) |  | 1.23 | 2.39 | 1.4 | 0.98 | 0.32 | 0.37 | 0.37 | 1.97 | 0.19 | 0.77 | 0 | 1.58 | 0.06 | 0.18 | 0.67
| Analytical results for Hg in OREAS 922 (Indicative Value < 1 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab |  |  |  |  | 1
| No | No | 01: 04; 02: 05; 03: 15; 04: 18 |  |  |  |  | ppm
|  |  | 01: 4A*OES; 02: 4A*OES; 03: 4A*OES; 04: 4A*OES |  |  |  |  | 2
2
| 1 | 1 | 01: < 1; 02: < 1; 03: < 10; 04: < 1 |  |  |  |  | 1
| 1 | 2 | 01: < 1; 02: < 1; 03: < 10; 04: < 1 |  |  |  |  | 1
| 1 | 3 | 01: < 1; 02: < 1; 03: < 10; 04: < 1 |  |  |  |  | 16
| 1 | 4 | 01: < 1; 02: 1; 03: < 10; 04: < 1 |  |  |  |  | < 1
| 1 | 5 | 01: < 1; 02: < 1; 03: < 10; 04: < 1 |  |  |  |  | 7
| 1 | 6 | 01: < 1; 02: < 1; 03: < 10; 04: < 1
Mean | bar: 02
Median | bar: 02
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 04
| Analytical results for Ho in OREAS 922 (Certified Value 1.04 ppm) | [21 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 13: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 13: 19 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 01: 1.27; 02: 1.04; 03: 1.2; 04: 1; 05: 1.04; 06: 1.0446; 07: 1.04; 09: 1; 13: 1 |  |  |  |  | 1
| 1 | 2 | 01: 1.27; 02: 1.06; 03: 1.2; 04: 1; 05: 1.08; 06: 1.0151; 07: 1.02; 09: 1.1; 13: 1.1 |  |  |  |  | #N/A
| 1 | 3 | 01: 1.26; 02: 1.05; 03: 1.2; 04: 1; 05: 1.06; 06: 1.0504; 07: 1.08; 09: 1.1; 13: 1 |  |  |  |  | 16
| 1 | 4 | 01: 1.27; 02: 1.08; 03: 1.2; 04: 1; 05: 1.08; 06: 1.0591; 07: 1.06; 09: 1; 13: 1 |  |  |  |  | 1.0433
| 1 | 5 | 01: 1.26; 02: 1.06; 03: 1.2; 04: 1; 05: 0.98; 06: 1.0666; 07: 1.07; 09: 1.1; 13: 1 |  |  |  |  | 28
| 1 | 6 | 01: 1.26; 02: 1.1; 03: 1.2; 04: 1; 05: 1.04; 06: 1.0345; 07: 1.11; 09: 1; 13: 1.1
| Mean |  | 01: 1.265; 02: 1.065; 03: 1.2; 04: 1; 05: 1.0467; 06: 1.0451; 07: 1.0633; 09: 1.05; 13: 1.0333
| Median |  | 01: 1.265; 02: 1.06; 03: 1.2; 04: 1; 05: 1.05; 06: 1.0475; 07: 1.065; 09: 1.05; 13: 1
| Std Dev. |  | 01: 0.0054772; 02: 0.021679; 03: 0; 04: 0; 05: 0.037238; 06: 0.018427; 07: 0.031411; 09: 0.054772; 13: 0.05164
| Rel.Std.Dev. |  | 01: 0.433%; 02: 2.036%; 03: 0%; 04: 0%; 05: 3.558%; 06: 1.763%; 07: 2.954%; 09: 5.216%; 13: 4.997%
| PDM3 |  | 01: 21.25%; 02: 2.076%; 03: 15.02%; 04: -4.154%; 05: 0.3188%; 06: 0.1642%; 07: 1.916%; 09: 0.6382%; 13: -0.9592%
| Z-Score (Absolute) |  | 01: 9.67; 02: 0.67; 03: 6.74; 04: 2.25; 05: 0.15; 06: 0.22; 07: 0.6; 09: 0; 13: 0.75
| Analytical results for In in OREAS 922 (Certified Value 0.31 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.36 | 0.42 | 0.308 | 0.3 | 0.3 | 0.327 | 0.288 | 0.3 | 0.29165 | 0.416 | < 10 | 0.311 | 0.3 | 0.3 | 0.33 | 0.303 |  |  |  |  | 1
| 1 | 2 | 0.34 | 0.41 | 0.313 | 0.3 | 0.3 | 0.328 | 0.294 | 0.28 | 0.2771 | 0.388 | < 10 | 0.317 | 0.3 | 0.3 | 0.35 | 0.328 |  |  |  |  | #N/A
| 1 | 3 | 0.33 | 0.41 | 0.322 | 0.3 | 0.3 | 0.335 | 0.3 | 0.3 | 0.27271 | 0.408 | < 10 | 0.312 | 0.3 | 0.3 | 0.33 | 0.335 |  |  |  |  | 16
| 1 | 4 | 0.34 | 0.41 | 0.309 | 0.3 | 0.3 | 0.338 | 0.287 | 0.32 | 0.29256 | 0.398 | < 10 | 0.32 | 0.3 | 0.3 | 0.35 | 0.326 |  |  |  |  | 0.30964
| 1 | 5 | 0.34 | 0.41 | 0.316 | 0.3 | 0.3 | 0.298 | 0.295 | 0.26 | 0.28598 | 0.4 | < 10 | 0.322 | 0.3 | 0.3 | 0.36 | 0.314 |  |  |  |  | 29
| 1 | 6 | 0.36 | 0.41 | 0.32 | 0.3 | 0.3 | 0.316 | 0.288 | 0.28 | 0.26456 | 0.441 | < 10 | 0.31 | 0.3 | 0.3 | 0.34 | 0.317
| Mean |  | 0.345 | 0.41167 | 0.31467 | 0.3 | 0.3 | 0.32367 | 0.292 | 0.29 | 0.28076 | 0.4085 |  | 0.31533 | 0.3 | 0.3 | 0.34333 | 0.3205
| Median |  | 0.34 | 0.41 | 0.3145 | 0.3 | 0.3 | 0.3275 | 0.291 | 0.29 | 0.28154 | 0.404 |  | 0.3145 | 0.3 | 0.3 | 0.345 | 0.3215
| Std Dev. |  | 0.012247 | 0.0040825 | 0.0057155 | 0 | 0 | 0.014706 | 0.0051769 | 0.020976 | 0.011187 | 0.018523 |  | 0.0050465 | 0 | 0 | 0.012111 | 0.011467
| Rel.Std.Dev. |  | 3.55% | 0.9917% | 1.816% | 0% | 0% | 4.544% | 1.773% | 7.233% | 3.985% | 4.534% |  | 1.6% | 0% | 0% | 3.527% | 3.578%
| PDM3 |  | 11.42% | 32.95% | 1.625% | -3.112% | -3.112% | 4.532% | -5.696% | -6.341% | -9.326% | 31.93% |  | 1.84% | -3.112% | -3.112% | 10.88% | 3.509%
| Z-Score (Absolute) |  | 1.39 | 4.46 | 0 | 0.67 | 0.67 | 0.41 | 1.04 | 1.13 | 1.56 | 4.31 | NA | 0.03 | 0.67 | 0.67 | 1.32 | 0.27
| Analytical results for K in OREAS 922 (Certified Value 2.54 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.55 | 2.55 | 2.4633 | 2.51 | 2.73 | 2.55 | 2.53 | 2.54 | 2.58 | 2.6147 | 2.4521 | 2.62 | 2.589 | 0.27 | 2.55 | 2.61 | 2.44 | 2.58 |  |  | 1
| 1 | 2 | 2.54 | 2.6 | 2.4895 | 2.47 | 2.66 | 2.53 | 2.67 | 2.58 | 2.52 | 2.5831 | 2.4738 | 2.56 | 2.538 | < 0.01 | 2.56 | 2.58 | 2.46 | 2.6 |  |  | #N/A
| 1 | 3 | 2.57 | 2.52 | 2.4497 | 2.38 | 2.74 | 2.68 | 2.53 | 2.52 | 1.65 | 2.5379 | 2.4257 | 2.58 | 2.516 | 0.18 | 2.54 | 2.56 | 2.48 | 2.51 |  |  | 16
| 1 | 4 | 2.52 | 2.54 | 2.4313 | 2.66 | 2.7 | 2.66 | 2.55 | 2.55 | 2.17 | 2.6072 | 2.4451 | 2.57 | 2.601 | 0.03 | 2.5 | 2.61 | 2.43 | 2.53 |  |  | 2.5406
| 1 | 5 | 2.6 | 2.56 | 2.4585 | 2.56 | 2.63 | 2.36 | 2.57 | 2.57 | 2.53 | 2.6078 | 2.5256 | 2.54 | 2.623 | 0.1 | 2.51 | 2.57 | 2.37 | 2.56 |  |  | 30
| 1 | 6 | 2.61 | 2.61 | 2.5177 | 2.41 | 2.77 | 2.47 | 2.52 | 2.56 | 2.65 | 2.6162 | 2.5235 | 2.56 | 2.595 | 0.25 | 2.48 | 2.63 | 2.46 | 2.59
| Mean |  | 2.565 | 2.5633 | 2.4683 | 2.4983 | 2.705 | 2.5417 | 2.5617 | 2.5533 | 2.35 | 2.5945 | 2.4743 | 2.5717 | 2.577 | 0.166 | 2.5233 | 2.5933 | 2.44 | 2.5617
| Median |  | 2.56 | 2.555 | 2.4609 | 2.49 | 2.715 | 2.54 | 2.54 | 2.555 | 2.525 | 2.6075 | 2.463 | 2.565 | 2.592 | 0.18 | 2.525 | 2.595 | 2.45 | 2.57
| Std Dev. |  | 0.035071 | 0.035024 | 0.030743 | 0.10265 | 0.05244 | 0.11957 | 0.056006 | 0.021602 | 0.38121 | 0.03015 | 0.041857 | 0.027142 | 0.040993 | 0.10114 | 0.031411 | 0.027325 | 0.038471 | 0.035449
| Rel.Std.Dev. |  | 1.367% | 1.366% | 1.246% | 4.109% | 1.939% | 4.704% | 2.186% | 0.846% | 16.22% | 1.162% | 1.692% | 1.055% | 1.591% | 60.93% | 1.245% | 1.054% | 1.577% | 1.384%
| PDM3 |  | 0.9599% | 0.8943% | -2.845% | -1.664% | 6.47% | 0.0415% | 0.8287% | 0.5007% | -7.503% | 2.121% | -2.61% | 1.222% | 1.432% | -93.47% | -0.6801% | 2.075% | -3.96% | 0.8287%
| Z-Score (Absolute) |  | 0.19 | 0.15 | 2.24 | 1.49 | 3.71 | 0.4 | 0.44 | 0.1 | 0.31 | 0.93 | 2.09 | 0.36 | 0.49 | 60.1 | 0.86 | 0.9 | 2.6 | 0.1
| Analytical results for La in OREAS 922 (Certified Value 43.4 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 53.3 | 44.48 | 44.8 | 40.7 | 45.8 | 43.1 | 42.9 | 41 | 46.5035 | 41.45 | 24 | 42.04 | 43.5 | 46 | 44.2 | 41.8 |  |  |  |  | 1
| 1 | 2 | 53.3 | 44.32 | 45.5 | 42.2 | 45.7 | 42.6 | 41.6 | 40 | 47.3713 | 40.5 | 23 | 41.94 | 43.8 | 45 | 43.5 | 43.7 |  |  |  |  | #N/A
| 1 | 3 | 52.8 | 44.28 | 44.8 | 42.9 | 46.7 | 42.3 | 41.5 | 36 | 47.0061 | 42.36 | 23 | 41.27 | 42.7 | 44 | 43.4 | 39.4 |  |  |  |  | 16
| 1 | 4 | 52.8 | 44.6 | 44.6 | 41.3 | 47 | 42.2 | 42.8 | 40 | 47.003 | 41.72 | 22 | 40.79 | 42.3 | 45 | 45 | 40.8 |  |  |  |  | 43.3738
| 1 | 5 | 52.6 | 44.15 | 44.9 | 41.7 | 42.5 | 39.7 | 41.2 | 43 | 47.4058 | 42.54 | 21 | 40.08 | 42.9 | 45 | 46.5 | 42.1 |  |  |  |  | 31
| 1 | 6 | 53 | 45.23 | 46.1 | 41.9 | 44.7 | 42.3 | 42.1 | 44 | 46.8583 | 42.02 | 22 | 42.18 | 42.5 | 45 | 44.6 | 42.1
| Mean |  | 52.9667 | 44.51 | 45.1167 | 41.7833 | 45.4 | 42.0333 | 42.0167 | 40.6667 | 47.0247 | 41.765 | 22.5 | 41.3833 | 42.95 | 45 | 44.5333 | 41.65
| Median |  | 52.9 | 44.4 | 44.85 | 41.8 | 45.75 | 42.3 | 41.85 | 40.5 | 47.0045 | 41.87 | 22.5 | 41.605 | 42.8 | 45 | 44.4 | 41.95
| Std Dev. |  | 0.28752 | 0.38626 | 0.57067 | 0.75476 | 1.6371 | 1.1894 | 0.70828 | 2.8048 | 0.33632 | 0.73772 | 1.0488 | 0.82938 | 0.58566 | 0.63246 | 1.1448 | 1.4433
| Rel.Std.Dev. |  | 0.5428% | 0.8678% | 1.265% | 1.806% | 3.606% | 2.83% | 1.686% | 6.897% | 0.7152% | 1.766% | 4.661% | 2.004% | 1.364% | 1.405% | 2.571% | 3.465%
| PDM3 |  | 22.12% | 2.62% | 4.018% | -3.667% | 4.672% | -3.09% | -3.129% | -6.241% | 8.417% | -3.709% | -48.13% | -4.589% | -0.9771% | 3.749% | 2.673% | -3.974%
| Z-Score (Absolute) |  | 4.42 | 0.77 | 1.03 | 0.41 | 1.15 | 0.1 | 0.31 | 0.49 | 1.85 | 0.41 | 8.72 | 0.58 | 0.1 | 0.98 | 0.78 | 0.46
| Analytical results for Li in OREAS 922 (Certified Value 30.8 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 33.5 | 30 | 29.2 | 30 | 26 | 31.5 | 30.4 | 30 | 29.2941 | 112.2 | 34 | 30.043 | 20 | 30.1 | 30 | 33 | 30.9 |  |  |  | 1
| 1 | 2 | 33.9 | 31 | 29 | 30 | 28 | 31.4 | 31 | 29 | 28.9022 | 104.8 | 34 | 28.188 | 28 | 30 | 31 | 33 | 32.2 |  |  |  | #N/A
| 1 | 3 | 33 | 31 | 29.3 | 29 | 29 | 31.7 | 30.9 | 30 | 28.1368 | 107.5 | 35 | 27.996 | 24 | 30.3 | 31 | 32 | 29.8 |  |  |  | 16
| 1 | 4 | 33.4 | 30 | 28.6 | 31 | 29 | 32.1 | 30.7 | 29.5 | 28.8936 | 111.4 | 34 | 28.937 | 25 | 29.9 | 31 | 34 | 30.9 |  |  |  | 30.8012
| 1 | 5 | 33.6 | 32 | 30.8 | 30 | 29 | 28.3 | 30.7 | 30 | 28.736 | 107.5 | 33 | 28.902 | 23 | 29.8 | 32 | 35 | 31.1 |  |  |  | 32
| 1 | 6 | 33.4 | 32 | 30 | 30 | 30 | 29.5 | 30.4 | 31 | 28.4477 | 126 | 34 | 30.287 | 21 | 29.6 | 31 | 34 | 31
| Mean |  | 33.4667 | 31 | 29.4833 | 30 | 28.5 | 30.75 | 30.6833 | 29.9167 | 28.7351 | 111.5667 | 34 | 29.0588 | 23.5 | 29.95 | 31 | 33.5 | 30.9833
| Median |  | 33.45 | 31 | 29.25 | 30 | 29 | 31.45 | 30.7 | 30 | 28.8148 | 109.45 | 34 | 28.9195 | 23.5 | 29.95 | 31 | 33.5 | 30.95
| Std Dev. |  | 0.29439 | 0.89443 | 0.79099 | 0.63246 | 1.3784 | 1.5017 | 0.24833 | 0.66458 | 0.40151 | 7.5844 | 0.63246 | 0.93855 | 2.881 | 0.2429 | 0.63246 | 1.0488 | 0.76267
| Rel.Std.Dev. |  | 0.8797% | 2.885% | 2.683% | 2.108% | 4.837% | 4.883% | 0.8093% | 2.221% | 1.397% | 6.798% | 1.86% | 3.23% | 12.26% | 0.811% | 2.04% | 3.131% | 2.462%
| PDM3 |  | 8.654% | 0.6456% | -4.278% | -2.601% | -7.471% | -0.1661% | -0.3825% | -2.872% | -6.708% | 262.2% | 10.39% | -5.657% | -23.7% | -2.763% | 0.6456% | 8.762% | 0.5915%
| Z-Score (Absolute) |  | 1.56 | 0.18 | 0.67 | 0.38 | 0.95 | 0.31 | 0 | 0.43 | 1.09 | 43.83 | 1.86 | 0.91 | 4.04 | 0.41 | 0.18 | 1.58 | 0.17
| Analytical results for Lu in OREAS 922 (Certified Value 0.41 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.51 | 0.381 | 0.4 | 0.4 | 0.44 | 0.42352 | 0.613 | 0.411 | 0.4 | 0.4 | 0.46 | [9 empty] | 1
| 1 | 2 | 0.5 | 0.385 | 0.4 | 0.4 | 0.44 | 0.43964 | 0.561 | 0.42 | 0.4 | 0.4 | 0.45 | [9 empty] | #N/A
| 1 | 3 | 0.5 | 0.388 | 0.4 | 0.4 | 0.42 | 0.42948 | 0.613 | 0.419 | 0.4 | 0.4 | 0.45 | [9 empty] | 16
| 1 | 4 | 0.5 | 0.389 | 0.4 | 0.4 | 0.44 | 0.41276 | 0.626 | 0.435 | 0.4 | 0.4 | 0.46 | [9 empty] | 0.41337
| 1 | 5 | 0.5 | 0.405 | 0.4 | 0.4 | 0.4 | 0.42559 | 0.618 | 0.428 | 0.4 | 0.4 | 0.48 | [9 empty] | 33
| 1 | 6 | 0.5 | 0.41 | 0.4 | 0.4 | 0.42 | 0.42203 | 0.605 | 0.414 | 0.4 | 0.4 | 0.45
| Mean |  | 0.50167 | 0.393 | 0.4 | 0.4 | 0.42667 | 0.4255 | 0.606 | 0.42117 | 0.4 | 0.4 | 0.45833
| Median |  | 0.5 | 0.3885 | 0.4 | 0.4 | 0.43 | 0.42455 | 0.613 | 0.4195 | 0.4 | 0.4 | 0.455
| Std Dev. |  | 0.0040825 | 0.011679 | 6.0809e-17 | 6.0809e-17 | 0.01633 | 0.0088784 | 0.0231 | 0.0089312 | 6.0809e-17 | 6.0809e-17 | 0.01169
| Rel.Std.Dev. |  | 0.8138% | 2.972% | 1.52e-14% | 1.52e-14% | 3.827% | 2.087% | 3.812% | 2.121% | 1.52e-14% | 1.52e-14% | 2.551%
| PDM3 |  | 21.36% | -4.928% | -3.235% | -3.235% | 3.216% | 2.935% | 46.6% | 1.886% | -3.235% | -3.235% | 10.88%
| Z-Score (Absolute) |  | 2.56 | 0.9 | 0.67 | 0.67 | 0.18 | 0.14 | 6.17 | 0 | 0.67 | 0.67 | 1.05
| Analytical results for Mg in OREAS 922 (Certified Value 1.58 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.66 | 1.6 | 1.6088 | 1.71 | 1.4 | 1.55 | 1.47 | 1.67 | 1.56 | 1.642 | 1.42 | 1.67 | 1.605 | 1.33 | 1.62 | 1.54 | 1.5 | 1.61 |  |  | 1
| 1 | 2 | 1.68 | 1.62 | 1.6514 | 1.69 | 1.51 | 1.52 | 1.55 | 1.69 | 1.53 | 1.6189 | 1.47 | 1.54 | 1.521 | 1.44 | 1.64 | 1.52 | 1.52 | 1.68 |  |  | #N/A
| 1 | 3 | 1.66 | 1.56 | 1.6259 | 1.66 | 1.54 | 1.59 | 1.45 | 1.69 | 1.46 | 1.6361 | 1.31 | 1.62 | 1.507 | 1.28 | 1.61 | 1.53 | 1.49 | 1.64 |  |  | 16
| 1 | 4 | 1.67 | 1.56 | 1.6084 | 1.69 | 1.59 | 1.61 | 1.48 | 1.68 | 1.54 | 1.6644 | 1.32 | 1.66 | 1.533 | 1.31 | 1.63 | 1.54 | 1.53 | 1.62 |  |  | 1.5805
| 1 | 5 | 1.66 | 1.58 | 1.6153 | 1.7 | 1.59 | 1.41 | 1.49 | 1.7 | 1.57 | 1.6492 | 1.43 | 1.5 | 1.58 | 1.22 | 1.62 | 1.53 | 1.62 | 1.66 |  |  | 34
| 1 | 6 | 1.67 | 1.61 | 1.6483 | 1.7 | 1.74 | 1.47 | 1.48 | 1.67 | 1.61 | 1.6382 | 1.45 | 1.57 | 1.537 | 1.55 | 1.58 | 1.55 | 1.53 | 1.64
| Mean |  | 1.6667 | 1.5883 | 1.6264 | 1.6917 | 1.5617 | 1.525 | 1.4867 | 1.6833 | 1.545 | 1.6415 | 1.4 | 1.5933 | 1.5472 | 1.355 | 1.6167 | 1.535 | 1.5317 | 1.6417
| Median |  | 1.665 | 1.59 | 1.6206 | 1.695 | 1.565 | 1.535 | 1.48 | 1.685 | 1.55 | 1.6401 | 1.425 | 1.595 | 1.535 | 1.32 | 1.62 | 1.535 | 1.525 | 1.64
| Std Dev. |  | 0.008165 | 0.025626 | 0.019297 | 0.017224 | 0.11197 | 0.075299 | 0.033862 | 0.012111 | 0.0501 | 0.015041 | 0.068118 | 0.06802 | 0.037483 | 0.11979 | 0.020656 | 0.010488 | 0.046224 | 0.025626
| Rel.Std.Dev. |  | 0.4899% | 1.613% | 1.187% | 1.018% | 7.17% | 4.938% | 2.278% | 0.7194% | 3.243% | 0.9163% | 4.866% | 4.269% | 2.423% | 8.841% | 1.278% | 0.6833% | 3.018% | 1.561%
| PDM3 |  | 5.449% | 0.4926% | 2.898% | 7.03% | -1.195% | -3.514% | -5.94% | 6.503% | -2.249% | 3.854% | -11.42% | 0.809% | -2.112% | -14.27% | 2.285% | -2.882% | -3.093% | 3.867%
| Z-Score (Absolute) |  | 1.1 | 0.16 | 0.62 | 1.4 | 0.16 | 0.6 | 1.21 | 1.3 | 0.36 | 0.8 | 1.88 | 0.22 | 0.33 | 2.64 | 0.5 | 0.48 | 0.73 | 0.8
| Analytical results for Mn in OREAS 922 (Certified Value 0.083 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.0835 | 0.0801 | 0.0859 | 0.0875 | 0.0731 | 0.0825 | 0.0819 | 0.086 | 0.081 | 0.082407 | 0.0851 | 0.096 | 0.081486 | 0.0732 | 0.0861 | 0.08435 | 0.076 | 0.0851 |  |  | 1
| 1 | 2 | 0.0842 | 0.081 | 0.0881 | 0.0855 | 0.0802 | 0.082 | 0.0851 | 0.0874 | 0.0798 | 0.083965 | 0.0874 | 0.09 | 0.079409 | 0.0799 | 0.0909 | 0.08513 | 0.0773 | 0.0847 |  |  | #N/A
| 1 | 3 | 0.0847 | 0.0797 | 0.0856 | 0.0856 | 0.0811 | 0.0865 | 0.0813 | 0.0878 | 0.0769 | 0.081247 | 0.0794 | 0.092 | 0.07903 | 0.0756 | 0.0874 | 0.08472 | 0.0759 | 0.0832 |  |  | 16
| 1 | 4 | 0.0837 | 0.0797 | 0.0857 | 0.0852 | 0.084 | 0.0856 | 0.0822 | 0.0852 | 0.0804 | 0.08293 | 0.0792 | 0.094 | 0.080442 | 0.0749 | 0.0849 | 0.08491 | 0.0781 | 0.0834 |  |  | 0.083049
| 1 | 5 | 0.0818 | 0.0797 | 0.0848 | 0.0876 | 0.0846 | 0.0755 | 0.0832 | 0.0854 | 0.0822 | 0.083922 | 0.0867 | 0.088 | 0.080565 | 0.0717 | 0.0853 | 0.08556 | 0.0827 | 0.0843 |  |  | 35
| 1 | 6 | 0.0829 | 0.0811 | 0.0856 | 0.0869 | 0.0847 | 0.0794 | 0.0813 | 0.087 | 0.0839 | 0.083165 | 0.0893 | 0.09 | 0.080623 | 0.0765 | 0.0827 | 0.08627 | 0.0779 | 0.085
| Mean |  | 0.083467 | 0.080217 | 0.08595 | 0.086383 | 0.081283 | 0.081917 | 0.0825 | 0.086467 | 0.0807 | 0.082939 | 0.084517 | 0.091667 | 0.080259 | 0.0753 | 0.086217 | 0.085157 | 0.077983 | 0.084283
| Median |  | 0.0836 | 0.0799 | 0.08565 | 0.08625 | 0.08255 | 0.08225 | 0.08205 | 0.0865 | 0.0807 | 0.083047 | 0.0859 | 0.091 | 0.080503 | 0.07525 | 0.0857 | 0.08502 | 0.0776 | 0.0845
| Std Dev. |  | 0.0010211 | 0.00066458 | 0.0011185 | 0.0010759 | 0.0044323 | 0.0040603 | 0.0014546 | 0.0010857 | 0.002363 | 0.0010215 | 0.0042602 | 0.0029439 | 0.00089436 | 0.0028362 | 0.0027672 | 0.00067899 | 0.0024903 | 0.00081343
| Rel.Std.Dev. |  | 1.223% | 0.8285% | 1.301% | 1.246% | 5.453% | 4.957% | 1.763% | 1.256% | 2.928% | 1.232% | 5.041% | 3.212% | 1.114% | 3.767% | 3.21% | 0.7973% | 3.193% | 0.9651%
| PDM3 |  | 0.5034% | -3.41% | 3.494% | 4.015% | -2.126% | -1.363% | -0.6606% | 4.116% | -2.828% | -0.1315% | 1.768% | 10.38% | -3.359% | -9.33% | 3.815% | 2.538% | -6.099% | 1.487%
| Z-Score (Absolute) |  | 0.07 | 0.77 | 0.71 | 0.82 | 0.49 | 0.33 | 0.31 | 0.84 | 0.64 | 0.07 | 0.34 | 1.95 | 0.76 | 2.03 | 0.77 | 0.5 | 1.58 | 0.28
| Analytical results for Mo in OREAS 922 (Certified Value 0.74 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.8 | 1 | 0.4 | 0.6 | 0.7 | 0.81 | 0.5 | < 2 | 0.66019 | < 2 | 2 | 0.898 | < 1 | < 1 | 0.9 | 0.79 | 0.68 |  |  |  | 1
| 1 | 2 | 0.7 | 0.8 | 0.7 | 0.6 | 0.68 | 0.78 | 1 | < 2 | 0.61557 | < 2 | < 2 | 0.928 | < 1 | < 1 | 0.9 | 0.8 | 0.71 |  |  |  | #N/A
| 1 | 3 | 0.7 | 0.8 | 0.6 | 0.8 | 0.73 | 0.75 | 0.5 | < 2 | 0.70062 | < 2 | < 2 | 0.936 | < 1 | < 1 | 0.7 | 0.74 | 0.77 |  |  |  | 16
| 1 | 4 | 0.8 | 0.7 | 0.8 | 1 | 0.85 | 0.74 | 1 | < 2 | 0.64757 | < 2 | < 2 | 1.043 | < 1 | < 1 | 0.8 | 0.84 | 0.7 |  |  |  | 0.74475
| 1 | 5 | 0.8 | 0.7 | 0.8 | 0.7 | 0.73 | 0.81 | 0.5 | < 2 | 0.62372 | < 2 | < 2 | 0.984 | < 1 | < 1 | 0.8 | 0.76 | 0.71 |  |  |  | 36
| 1 | 6 | 0.8 | 0.8 | 0.7 | 0.8 | 0.72 | 0.77 | 1 | < 2 | 0.94348 | < 2 | < 2 | 0.919 | < 1 | < 1 | 0.8 | 0.81 | 1.07
| Mean |  | 0.76667 | 0.8 | 0.66667 | 0.75 | 0.735 | 0.77667 | 0.75 |  | 0.69853 |  | 2 | 0.95133 |  |  | 0.81667 | 0.79 | 0.77333
| Median |  | 0.8 | 0.8 | 0.7 | 0.75 | 0.725 | 0.775 | 0.75 |  | 0.65388 |  | 2 | 0.932 |  |  | 0.8 | 0.795 | 0.71
| Std Dev. |  | 0.05164 | 0.10954 | 0.15055 | 0.15166 | 0.059582 | 0.029439 | 0.27386 |  | 0.12373 |  |  | 0.053178 |  |  | 0.075277 | 0.035777 | 0.14841
| Rel.Std.Dev. |  | 6.736% | 13.69% | 22.58% | 20.22% | 8.106% | 3.79% | 36.51% |  | 17.71% |  |  | 5.59% |  |  | 9.218% | 4.529% | 19.19%
| PDM3 |  | 2.943% | 7.419% | -10.48% | 0.7055% | -1.309% | 4.286% | 0.7055% |  | -6.206% |  | 168.5% | 27.74% |  |  | 9.657% | 6.076% | 3.839%
| Z-Score (Absolute) |  | 0 | 0.45 | 1.35 | 0.22 | 0.74 | 0.13 | 0.22 | NA | 1.58 | NA | 16.63 | 2.24 | NA | NA | 0.67 | 0.31 | 0.71
| Analytical results for Na in OREAS 922 (Certified Value 0.441 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.465 | 0.44 | 0.4537 | 0.45 | 0.41 | 0.45 | 0.42 | 0.46 | 0.41 | 0.46479 | 0.4 | 0.55 | 0.434 | 0.44 | 0.49 | 0.42 | 0.43 |  |  |  | 1
| 1 | 2 | 0.462 | 0.45 | 0.4578 | 0.45 | 0.43 | 0.44 | 0.44 | 0.47 | 0.4 | 0.4543 | 0.42 | 0.52 | 0.428 | 0.44 | 0.48 | 0.42 | 0.45 |  |  |  | #N/A
| 1 | 3 | 0.461 | 0.43 | 0.4455 | 0.44 | 0.44 | 0.46 | 0.42 | 0.46 | 0.39 | 0.45703 | 0.38 | 0.51 | 0.425 | 0.43 | 0.48 | 0.41 | 0.44 |  |  |  | 16
| 1 | 4 | 0.462 | 0.43 | 0.446 | 0.48 | 0.45 | 0.47 | 0.42 | 0.46 | 0.4 | 0.46276 | 0.37 | 0.52 | 0.434 | 0.44 | 0.49 | 0.42 | 0.43 |  |  |  | 0.44101
| 1 | 5 | 0.465 | 0.44 | 0.4498 | 0.46 | 0.45 | 0.41 | 0.43 | 0.46 | 0.41 | 0.46065 | 0.41 | 0.5 | 0.442 | 0.44 | 0.48 | 0.45 | 0.44 |  |  |  | 37
| 1 | 6 | 0.471 | 0.45 | 0.4585 | 0.45 | 0.45 | 0.43 | 0.42 | 0.46 | 0.42 | 0.45679 | 0.41 | 0.52 | 0.44 | 0.42 | 0.49 | 0.44 | 0.44
| Mean |  | 0.46433 | 0.44 | 0.45188 | 0.455 | 0.43833 | 0.44333 | 0.425 | 0.46167 | 0.405 | 0.45938 | 0.39833 | 0.52 | 0.43383 | 0.435 | 0.485 | 0.42667 | 0.43833
| Median |  | 0.4635 | 0.44 | 0.45175 | 0.45 | 0.445 | 0.445 | 0.42 | 0.46 | 0.405 | 0.45884 | 0.405 | 0.52 | 0.434 | 0.44 | 0.485 | 0.42 | 0.44
| Std Dev. |  | 0.0036697 | 0.0089443 | 0.0056884 | 0.013784 | 0.016021 | 0.021602 | 0.0083666 | 0.0040825 | 0.010488 | 0.0040078 | 0.019408 | 0.016733 | 0.0065853 | 0.0083666 | 0.0054772 | 0.015055 | 0.0075277
| Rel.Std.Dev. |  | 0.7903% | 2.033% | 1.259% | 3.029% | 3.655% | 4.873% | 1.969% | 0.8843% | 2.59% | 0.8724% | 4.872% | 3.218% | 1.518% | 1.923% | 1.129% | 3.529% | 1.717%
| PDM3 |  | 5.289% | -0.2282% | 2.466% | 3.173% | -0.6061% | 0.5277% | -3.629% | 4.685% | -8.165% | 4.167% | -9.676% | 17.91% | -1.627% | -1.362% | 9.976% | -3.252% | -0.6061%
| Z-Score (Absolute) |  | 1.23 | 0 | 0.6 | 0.51 | 0.08 | 0.17 | 0.91 | 1.1 | 1.77 | 0.98 | 2.11 | 3.74 | 0.31 | 0.1 | 2.28 | 0.67 | 0.08
| Analytical results for Nb in OREAS 922 (Certified Value 14.9 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.5 | 39.3 | 14.55 | 2.4 | 2.1 | 15.5 | 15.5 | 14 | 14 | 13.9878 | 14.74 | 15 | 17.647 | 2.3 | 15.4 | 15.9 | 15.1 |  |  |  | 1
| 1 | 2 | 14.8 | 39.1 | 14.54 | 10.6 | 2.9 | 15.5 | 15.4 | 14 | 14 | 13.8259 | 14.6 | 15 | 17.582 | 5.8 | 15.4 | 16.2 | 15.7 |  |  |  | #N/A
| 1 | 3 | 14.6 | 39.1 | 14.47 | 10.9 | 7.4 | 16.1 | 15.5 | 13.5 | 14 | 13.9888 | 15.17 | 15 | 17.213 | 6.6 | 15.5 | 15.7 | 14.9 |  |  |  | 16
| 1 | 4 | 14 | 39.3 | 14.47 | 8.9 | 11.8 | 15.9 | 15.4 | 13.5 | 14 | 14.0686 | 15.63 | 15 | 17.438 | 1.7 | 15.6 | 16 | 15.5 |  |  |  | 14.9453
| 1 | 5 | 14.6 | 39.2 | 14.23 | 5 | 12.1 | 14.2 | 15.6 | 13.5 | 15 | 14.3561 | 15.58 | 14 | 17.754 | 3.4 | 15.7 | 17.1 | 15.1 |  |  |  | 38
| 1 | 6 | 14.8 | 39.1 | 14.85 | 9.8 | 14.2 | 14.8 | 15.7 | 14 | 15 | 13.9031 | 15.94 | 15 | 17.151 | 0.8 | 16 | 16.2 | 15.7
| Mean |  | 14.55 | 39.1833 | 14.5183 | 7.9333 | 8.4167 | 15.3333 | 15.5167 | 13.75 | 14.3333 | 14.0217 | 15.2767 | 14.8333 | 17.4642 | 3.4333 | 15.6 | 16.1833 | 15.3333
| Median |  | 14.6 | 39.15 | 14.505 | 9.35 | 9.6 | 15.5 | 15.5 | 13.75 | 14 | 13.9883 | 15.375 | 15 | 17.51 | 2.85 | 15.55 | 16.1 | 15.3
| Std Dev. |  | 0.29496 | 0.098319 | 0.19964 | 3.4512 | 5.0957 | 0.71181 | 0.1169 | 0.27386 | 0.5164 | 0.18365 | 0.53181 | 0.40825 | 0.24214 | 2.3175 | 0.22804 | 0.48751 | 0.34448
| Rel.Std.Dev. |  | 2.027% | 0.2509% | 1.375% | 43.5% | 60.54% | 4.642% | 0.7534% | 1.992% | 3.603% | 1.31% | 3.481% | 2.752% | 1.386% | 67.5% | 1.462% | 3.012% | 2.247%
| PDM3 |  | -2.645% | 162.2% | -2.857% | -46.92% | -43.68% | 2.596% | 3.823% | -7.998% | -4.095% | -6.18% | 2.217% | -0.7491% | 16.85% | -77.03% | 4.381% | 8.284% | 2.596%
| Z-Score (Absolute) |  | 0.34 | 24.46 | 0.49 | 7.15 | 6.66 | 0.34 | 0.52 | 1.26 | 0.67 | 0.99 | 0.28 | 0 | 2.49 | 11.7 | 0.61 | 1.01 | 0.34
| Analytical results for Nd in OREAS 922 (Certified Value 36.7 ppm) | [21 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 11: 19 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 11: 4A*MS |  |  | 1
2
| 1 | 1 | 01: 47.8; 02: 35.65; 03: 38.7; 04: 33.8; 05: 35.7; 06: 38.3467; 07: 37.17; 09: 35.9; 11: 36.5 |  |  | 1
| 1 | 2 | 01: 47.6; 02: 36.09; 03: 38.8; 04: 35.2; 05: 35.3; 06: 38.0959; 07: 36.22; 09: 37.1; 11: 37.9 |  |  | #N/A
| 1 | 3 | 01: 47.3; 02: 36.1; 03: 37.6; 04: 35; 05: 35.2; 06: 37.9881; 07: 38.04; 09: 36.9; 11: 35.4 |  |  | 16
| 1 | 4 | 01: 47.2; 02: 36.39; 03: 38; 04: 34.1; 05: 35.9; 06: 38.819; 07: 38.36; 09: 36.2; 11: 36.3 |  |  | 36.666
| 1 | 5 | 01: 47.1; 02: 36.14; 03: 38; 04: 33.8; 05: 34.9; 06: 38.4151; 07: 38.78; 09: 35.9; 11: 37.5 |  |  | 39
| 1 | 6 | 01: 47.4; 02: 36.27; 03: 38.8; 04: 33.9; 05: 35.8; 06: 37.9313; 07: 37.06; 09: 34.9; 11: 39.1
| Mean |  | 01: 47.4; 02: 36.1067; 03: 38.3167; 04: 34.3; 05: 35.4667; 06: 38.266; 07: 37.605; 09: 36.15; 11: 37.1167
| Median |  | 01: 47.35; 02: 36.12; 03: 38.35; 04: 34; 05: 35.5; 06: 38.2213; 07: 37.605; 09: 36.05; 11: 37
| Std Dev. |  | 01: 0.26077; 02: 0.25177; 03: 0.51543; 04: 0.63246; 05: 0.39328; 06: 0.33248; 07: 0.95335; 09: 0.79436; 11: 1.3182
| Rel.Std.Dev. |  | 01: 0.5501%; 02: 0.6973%; 03: 1.345%; 04: 1.844%; 05: 1.109%; 06: 0.8689%; 07: 2.535%; 09: 2.197%; 11: 3.552%
| PDM3 |  | 01: 29.28%; 02: -1.525%; 03: 4.502%; 04: -6.453%; 05: -3.271%; 06: 4.364%; 07: 2.561%; 09: -1.407%; 11: 1.229%
| Z-Score (Absolute) |  | 01: 6.03; 02: 0.59; 03: 0.7; 04: 1.65; 05: 0.97; 06: 0.67; 07: 0.29; 09: 0.57; 11: 0
| Analytical results for Ni in OREAS 922 (Certified Value 37.9 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 37 | 42.5 | 36 | 43.2 | 32 | 36.8 | 37.5 | 44 | 38 | 37 | 40 | 41 | 34.1 | 39 | 37.7 | 38 | 34.4 | 36.5 |  |  | 1
| 1 | 2 | 35 | 42.4 | 37 | 42.8 | 37 | 36.6 | 39.4 | 42 | 38 | 37 | 35 | 41 | 34.39 | 38 | 37.5 | 38 | 36.5 | 36.9 |  |  | #N/A
| 1 | 3 | 36 | 42.2 | 36 | 42.7 | 35 | 37 | 41.8 | 42 | 37 | 37.7986 | 34 | 40 | 33 | 46 | 52.3 | 38 | 34.6 | 36 |  |  | 16
| 1 | 4 | 37 | 42.1 | 35 | 43.2 | 38 | 37.8 | 37.7 | 38 | 39 | 37.6085 | 38 | 40 | 33.65 | 40 | 37.3 | 38 | 35.6 | 36.2 |  |  | 37.8739
| 1 | 5 | 35 | 42.1 | 36 | 43.7 | 37 | 33.3 | 38 | 38 | 40 | 36.2556 | 47 | 39 | 34.23 | 39 | 38.5 | 39 | 37.6 | 36.4 |  |  | 40
| 1 | 6 | 35 | 42.1 | 36 | 43 | 37 | 36.2 | 37.5 | 40 | 40 | 37.0835 | 39 | 41 | 33.06 | 34 | 36.9 | 38 | 34.8 | 41.4
| Mean |  | 35.8333 | 42.2333 | 36 | 43.1 | 36 | 36.2833 | 38.65 | 40.6667 | 38.6667 | 37.1244 | 38.8333 | 40.3333 | 33.7383 | 39.3333 | 40.0333 | 38.1667 | 35.5833 | 37.2333
| Median |  | 35.5 | 42.15 | 36 | 43.1 | 37 | 36.7 | 37.85 | 41 | 38.5 | 37.0418 | 38.5 | 40.5 | 33.875 | 39 | 37.6 | 38 | 35.2 | 36.45
| Std Dev. |  | 0.98319 | 0.17512 | 0.63246 | 0.35777 | 2.1909 | 1.5549 | 1.7003 | 2.4221 | 1.2111 | 0.54338 | 4.6224 | 0.8165 | 0.60171 | 3.8816 | 6.0328 | 0.40825 | 1.2561 | 2.0637
| Rel.Std.Dev. |  | 2.744% | 0.4146% | 1.757% | 0.8301% | 6.086% | 4.285% | 4.399% | 5.956% | 3.132% | 1.464% | 11.9% | 2.024% | 1.783% | 9.868% | 15.07% | 1.07% | 3.53% | 5.542%
| PDM3 |  | -5.388% | 11.51% | -4.948% | 13.8% | -4.948% | -4.2% | 2.049% | 7.374% | 2.093% | -1.979% | 2.533% | 6.494% | -10.92% | 3.854% | 5.702% | 0.7731% | -6.048% | -1.691%
| Z-Score (Absolute) |  | 0.88 | 2.74 | 0.79 | 3.23 | 0.33 | 0.63 | 0.36 | 1.85 | 0.72 | 0.15 | 0.11 | 1.67 | 2.07 | 0.35 | 0.11 | 0.44 | 1.02 | 0.56
| Analytical results for P in OREAS 922 (Certified Value 0.066 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.061 | 0.0649 | 0.0597 | 0.064 | 0.053 | 0.07 | 0.063 | 0.07 | 0.075 | 0.06882 | 0.0619 | 0.07 | 0.069421 | 0.064 | 0.063 | 0.0695 | 0.065 | 0.069 |  |  | 1
| 1 | 2 | 0.062 | 0.0687 | 0.0639 | 0.064 | 0.059 | 0.068 | 0.067 | 0.075 | 0.077 | 0.071 | 0.061 | 0.07 | 0.063621 | 0.0699 | 0.065 | 0.0698 | 0.067 | 0.072 |  |  | #N/A
| 1 | 3 | 0.062 | 0.0646 | 0.0605 | 0.063 | 0.06 | 0.071 | 0.063 | 0.07 | 0.074 | 0.068 | 0.0617 | 0.07 | 0.062557 | 0.0656 | 0.063 | 0.0679 | 0.062 | 0.069 |  |  | 16
| 1 | 4 | 0.062 | 0.065 | 0.059 | 0.067 | 0.062 | 0.072 | 0.065 | 0.07 | 0.078 | 0.067 | 0.0602 | 0.068 | 0.071584 | 0.0651 | 0.065 | 0.0691 | 0.065 | 0.069 |  |  | 0.065825
| 1 | 5 | 0.062 | 0.0656 | 0.0612 | 0.065 | 0.062 | 0.064 | 0.065 | 0.07 | 0.076 | 0.065884 | 0.0648 | 0.066 | 0.067353 | 0.0626 | 0.06 | 0.0681 | 0.068 | 0.072 |  |  | 41
| 1 | 6 | 0.062 | 0.0699 | 0.0618 | 0.065 | 0.065 | 0.065 | 0.064 | 0.075 | 0.078 | 0.069 | 0.0645 | 0.068 | 0.062989 | 0.068 | 0.058 | 0.0718 | 0.067 | 0.07
| Mean |  | 0.061833 | 0.06645 | 0.061017 | 0.064667 | 0.060167 | 0.068333 | 0.0645 | 0.071667 | 0.076333 | 0.068284 | 0.06235 | 0.068667 | 0.066254 | 0.065867 | 0.062333 | 0.069367 | 0.065667 | 0.070167
| Median |  | 0.062 | 0.0653 | 0.06085 | 0.0645 | 0.061 | 0.069 | 0.0645 | 0.07 | 0.0765 | 0.06841 | 0.0618 | 0.069 | 0.065487 | 0.06535 | 0.063 | 0.0693 | 0.066 | 0.0695
| Std Dev. |  | 0.00040825 | 0.0022634 | 0.0017337 | 0.0013663 | 0.0040702 | 0.003266 | 0.0015166 | 0.002582 | 0.001633 | 0.00177 | 0.0018812 | 0.001633 | 0.0037658 | 0.0026696 | 0.0028048 | 0.0014109 | 0.0021602 | 0.001472
| Rel.Std.Dev. |  | 0.6602% | 3.406% | 2.841% | 2.113% | 6.765% | 4.779% | 2.351% | 3.603% | 2.139% | 2.592% | 3.017% | 2.378% | 5.684% | 4.053% | 4.5% | 2.034% | 3.29% | 2.098%
| PDM3 |  | -6.064% | 0.9498% | -7.304% | -1.759% | -8.596% | 3.811% | -2.013% | 8.875% | 15.96% | 3.736% | -5.279% | 4.317% | 0.6523% | 0.06361% | -5.304% | 5.381% | -0.2402% | 6.596%
| Z-Score (Absolute) |  | 0.96 | 0.09 | 1.15 | 0.32 | 1.02 | 0.52 | 0.36 | 1.28 | 2.34 | 0.51 | 0.85 | 0.59 | 0.04 | 0.04 | 0.85 | 0.75 | 0.09 | 0.94
| Analytical results for Pb in OREAS 922 (Certified Value 59 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 59 | 57.1 | 63.1 | 58 | 52.9 | 61.2 | 62 | 61 | 62.1 | 61 | 61.0717 | 57 | 79 | 57.68 | 56 | 59.8 | 59 | 58.2 | 60.4 |  | 1
| 1 | 2 | 62 | 56.6 | 61.4 | 58 | 57 | 69.2 | 60.7 | 60 | 61.6 | 60 | 60.0111 | 58 | 77 | 60.675 | 63 | 57.7 | 59 | 58.6 | 62.6 |  | #N/A
| 1 | 3 | 58 | 57.4 | 59.5 | 57 | 54.8 | 65.5 | 62.1 | 64 | 62.7 | 59 | 59.298 | 53 | 77 | 56.642 | 54 | 60.4 | 55 | 54.2 | 54.3 |  | 16
| 1 | 4 | 60 | 56.7 | 59.4 | 56 | 53.4 | 70 | 60.5 | 60 | 65.2 | 60 | 61 | 54 | 78 | 58.071 | 59 | 58.2 | 57 | 61.2 | 62.1 |  | 59.2438
| 1 | 5 | 62 | 56 | 59.7 | 58 | 52 | 59.9 | 58.6 | 59 | 63.5 | 64 | 61.9243 | 56 | 78 | 59.527 | 51 | 60.3 | 59 | 64.6 | 59.6 |  | 42
| 1 | 6 | 61 | 56.4 | 61.4 | 55 | 55 | 57.4 | 60.3 | 61 | 67.3 | 63 | 62.4641 | 55 | 78 | 60.724 | 58 | 57.2 | 59 | 58.1 | 57.9
| Mean |  | 60.3333 | 56.7 | 60.75 | 57 | 54.1833 | 63.8667 | 60.7 | 60.8333 | 63.7333 | 61.1667 | 60.9615 | 55.5 | 77.8333 | 58.8865 | 56.8333 | 58.9333 | 58 | 59.15 | 59.4833
| Median |  | 60.5 | 56.65 | 60.55 | 57.5 | 54.1 | 63.35 | 60.6 | 60.5 | 63.1 | 60.5 | 61.0358 | 55.5 | 78 | 58.799 | 57 | 59 | 59 | 58.4 | 60
| Std Dev. |  | 1.633 | 0.498 | 1.4734 | 1.2649 | 1.7893 | 5.1651 | 1.2853 | 1.7224 | 2.1547 | 1.9408 | 1.1721 | 1.8708 | 0.75277 | 1.6819 | 4.1673 | 1.4024 | 1.6733 | 3.4847 | 3.059
| Rel.Std.Dev. |  | 2.707% | 0.8783% | 2.425% | 2.219% | 3.302% | 8.087% | 2.117% | 2.831% | 3.381% | 3.173% | 1.923% | 3.371% | 0.9672% | 2.856% | 7.333% | 2.38% | 2.885% | 5.891% | 5.143%
| PDM3 |  | 1.839% | -4.294% | 2.542% | -3.787% | -8.542% | 7.803% | 2.458% | 2.683% | 7.578% | 3.246% | 2.899% | -6.319% | 31.38% | -0.6031% | -4.069% | -0.524% | -2.099% | -0.1583% | 0.4044%
| Z-Score (Absolute) |  | 0.39 | 1.27 | 0.58 | 1.13 | 2.42 | 1.44 | 0.56 | 0.62 | 1.94 | 0.77 | 0.67 | 1.82 | 8.37 | 0.27 | 1.21 | 0.25 | 0.4 | 0.15 | 0
| Analytical results for Pr in OREAS 922 (Certified Value 10 ppm) | [21 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 10: 4A*MS |  | 1
2
| 1 | 1 | 01: 13.2; 02: 9.713; 03: 10.8; 04: 9.2; 05: 9.9; 06: 10.2269; 07: 9.795; 09: 9.7; 10: 10.4 |  | 1
| 1 | 2 | 01: 13.2; 02: 9.693; 03: 11.1; 04: 9.6; 05: 9.8; 06: 10.2121; 07: 9.566; 09: 10; 10: 10.3 |  | #N/A
| 1 | 3 | 01: 13.1; 02: 9.67; 03: 10.8; 04: 9.6; 05: 10; 06: 10.4572; 07: 9.894; 09: 9.8; 10: 10.1 |  | 16
| 1 | 4 | 01: 13.1; 02: 9.727; 03: 10.8; 04: 9.3; 05: 10; 06: 10.418; 07: 10.114; 09: 9.6; 10: 10.3 |  | 10.0315
| 1 | 5 | 01: 13.1; 02: 9.61; 03: 10.9; 04: 9.3; 05: 9.6; 06: 10.3332; 07: 10.168; 09: 9.5; 10: 10.5 |  | 43
| 1 | 6 | 01: 13.1; 02: 9.89; 03: 11; 04: 9.3; 05: 10.3; 06: 10.2307; 07: 10.193; 09: 9.4; 10: 10.7
| Mean |  | 01: 13.1333; 02: 9.7172; 03: 10.9; 04: 9.3833; 05: 9.9333; 06: 10.313; 07: 9.955; 09: 9.6667; 10: 10.3833
| Median |  | 01: 13.1; 02: 9.703; 03: 10.85; 04: 9.3; 05: 9.95; 06: 10.2819; 07: 10.004; 09: 9.65; 10: 10.35
| Std Dev. |  | 01: 0.05164; 02: 0.094105; 03: 0.12649; 04: 0.17224; 05: 0.23381; 06: 0.10638; 07: 0.24817; 09: 0.21602; 10: 0.20412
| Rel.Std.Dev. |  | 01: 0.3932%; 02: 0.9684%; 03: 1.16%; 04: 1.836%; 05: 2.354%; 06: 1.032%; 07: 2.493%; 09: 2.235%; 10: 1.966%
| PDM3 |  | 01: 30.92%; 02: -3.133%; 03: 8.658%; 04: -6.461%; 05: -0.9784%; 06: 2.807%; 07: -0.7624%; 09: -3.637%; 10: 3.507%
| Z-Score (Absolute) |  | 01: 5.99; 02: 0.45; 03: 1.78; 04: 1.08; 05: 0.04; 06: 0.67; 07: 0; 09: 0.54; 10: 0.81
| Analytical results for Rb in OREAS 922 (Certified Value 164 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 0
0
| 1 | 1 | 158 | 170 | 171.95 | 147 | 151 | 167 | 155.5 | 199 | 164.0039 | 178.12 | 167.01 | 157 | 169 | 122 | 170 | [5 empty] | 1
| 1 | 2 | 155 | 170 | 171.87 | 156 | 156 | 164.5 | 160 | 201 | 165.6604 | 175.2 | 170.7 | 158 | 165.3 | 135 | 177 | [5 empty] | #N/A
| 1 | 3 | 157 | 168 | 173.55 | 145 | 155 | 167 | 154.5 | 204 | 164.9647 | 183.44 | 170.38 | 158 | 161.6 | 123 | 171.5 | [5 empty] | 16
| 1 | 4 | 158 | 168 | 175.37 | 159 | 155 | 173.5 | 153.5 | 202 | 164.4143 | 184.69 | 175.84 | 152 | 163.6 | 137 | 173.5 | [5 empty] | 164.3114
| 1 | 5 | 157 | 168 | 170.7 | 156 | 151 | 152 | 147 | 186 | 165.1184 | 187.7 | 176.66 | 155 | 169.7 | 144 | 171.5 | [5 empty] | 44
| 1 | 6 | 156 | 169 | 173.59 | 152 | 160 | 160.5 | 156 | 207 | 164.6055 | 183.03 | 170.12 | 155 | 165.8 | 128 | 171.5
| Mean |  | 156.8333 | 168.8333 | 172.8383 | 152.5 | 154.6667 | 164.0833 | 154.4167 | 199.8333 | 164.7945 | 182.03 | 171.785 | 155.8333 | 165.8333 | 131.5 | 172.5
| Median |  | 157 | 168.5 | 172.75 | 154 | 155 | 165.75 | 155 | 201.5 | 164.7851 | 183.235 | 170.54 | 156 | 165.55 | 131.5 | 171.5
| Std Dev. |  | 1.169 | 0.98319 | 1.6601 | 5.5408 | 3.3862 | 7.2761 | 4.2593 | 7.3052 | 0.58119 | 4.564 | 3.713 | 2.3166 | 3.104 | 8.6429 | 2.4698
| Rel.Std.Dev. |  | 0.7454% | 0.5823% | 0.9605% | 3.633% | 2.189% | 4.434% | 2.758% | 3.656% | 0.3527% | 2.507% | 2.161% | 1.487% | 1.872% | 6.573% | 1.432%
| PDM3 |  | -4.551% | 2.752% | 5.19% | -7.188% | -5.87% | -0.1388% | -6.022% | 21.62% | 0.2941% | 10.78% | 4.548% | -5.16% | 0.9263% | -19.97% | 4.984%
| Z-Score (Absolute) |  | 0.67 | 0.34 | 0.67 | 1.03 | 0.85 | 0.06 | 0.87 | 3.17 | 0 | 1.44 | 0.59 | 0.75 | 0.09 | 2.79 | 0.57
| Analytical results for Re in OREAS 922 (Certified Value < 2 ppb) | [21 empty] | Certified
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab |  | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 05; 05: 06; 06: 07; 07: 09; 09: 13; 10: 18; 11: 19; 13: 22 |  | ppb
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 09: 4A*OES/MS; 10: 4A*MS; 11: 4A*MS; 13: 4A*MS |  | 0
0
| 1 | 1 | 01: < 50; 02: < 2; 03: < 1; 04: < 1; 05: < 2; 06: < 2; 07: < 100; 09: < 2; 10: < 1; 11: < 10; 13: 3 |  | 1
| 1 | 2 | 01: < 50; 02: < 2; 03: < 1; 04: < 1; 05: < 2; 06: < 2; 07: < 100; 09: < 2; 10: 1; 11: < 10; 13: 2 |  | 8
| 1 | 3 | 01: < 50; 02: < 2; 03: 1; 04: 1; 05: < 2; 06: < 2; 07: < 100; 09: < 2; 10: < 1; 11: 10; 13: 3 |  | 16
| 1 | 4 | 01: < 50; 02: < 2; 03: 1; 04: 1; 05: < 2; 06: < 2; 07: < 100; 09: < 2; 10: 1; 11: < 10; 13: < 2 |  | < 2
| 1 | 5 | 01: < 50; 02: < 2; 03: < 1; 04: < 1; 05: < 2; 06: < 2; 07: < 100; 09: 2.0503; 10: < 1; 11: < 10; 13: 2 |  | 45
| 1 | 6 | 01: < 50; 02: < 2; 03: 3; 04: < 1; 05: < 2; 06: < 2; 07: < 100; 09: < 2; 10: < 1; 11: < 10; 13: 2
| Mean |  | 03: 1.6667; 04: 1; 09: 2.0503; 10: 1; 11: 10; 13: 2.4
| Median |  | 03: 1; 04: 1; 09: 2.0503; 10: 1; 11: 10; 13: 2
| Std Dev. |  | 03: 1.1547; 04: 0; 10: 0; 13: 0.54772
| Rel.Std.Dev. |  | 03: 69.28%; 04: 0%; 10: 0%; 13: 22.82%
PDM3
| Z-Score (Absolute) |  | 01: NA; 02: NA; 03: 0.07; 04: 0.67; 05: NA; 06: NA; 07: NA; 09: 0.07; 10: 0.67; 11: 8.62; 13: 2.02
| Analytical results for Ru in OREAS 922 (Indicative Value < 0.1 ppm) | [21 empty] | Indicative
Ru | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 19 |  | ppm
|  |  | 01: 4A*MS |  | 2
2
| 1 | 1 | 01: < 0.1 |  | 1
| 1 | 2 | 01: < 0.1 |  | 4
| 1 | 3 | 01: < 0.1 |  | 16
| 1 | 4 | 01: < 0.1 |  | < 0.1
| 1 | 5 | 01: < 0.1 |  | 7
| 1 | 6 | 01: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for S in OREAS 922 (Certified Value 0.386 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.402 | 0.4 | 0.3822 | 0.4 | 0.33 | 0.38 | 0.35 | 0.41 | 0.3623 | 0.4 | 0.386 | 0.41 | 0.4 | 0.314 | 0.38 | 0.4 | 0.39 | 0.37 |  |  | 1
| 1 | 2 | 0.401 | 0.43 | 0.3823 | 0.4 | 0.35 | 0.37 | 0.38 | 0.39 | 0.3656 | 0.4 | 0.387 | 0.4 | 0.38 | 0.319 | 0.39 | 0.4 | 0.39 | 0.39 |  |  | 6
| 1 | 3 | 0.416 | 0.41 | 0.3876 | 0.39 | 0.36 | 0.39 | 0.36 | 0.41 | 0.3558 | 0.3 | 0.383 | 0.37 | 0.39 | 0.293 | 0.38 | 0.39 | 0.38 | 0.38 |  |  | 16
| 1 | 4 | 0.407 | 0.39 | 0.4018 | 0.39 | 0.37 | 0.4 | 0.35 | 0.39 | 0.3654 | 0.4 | 0.377 | 0.36 | 0.42 | 0.327 | 0.39 | 0.4 | 0.39 | 0.38 |  |  | 0.38568
| 1 | 5 | 0.397 | 0.43 | 0.3674 | 0.39 | 0.37 | 0.35 | 0.36 | 0.39 | 0.3523 | 0.4 | 0.389 | 0.4 | 0.37 | 0.331 | 0.39 | 0.4 | 0.42 | 0.38 |  |  | 46
| 1 | 6 | 0.4 | 0.44 | 0.3916 | 0.4 | 0.38 | 0.36 | 0.35 | 0.38 | 0.3572 | 0.4 | 0.385 | 0.39 | 0.38 | 0.317 | 0.38 | 0.4 | 0.39 | 0.38
| Mean |  | 0.40383 | 0.41667 | 0.38548 | 0.395 | 0.36 | 0.375 | 0.35833 | 0.395 | 0.35977 | 0.38333 | 0.3845 | 0.38833 | 0.39 | 0.31683 | 0.385 | 0.39833 | 0.39333 | 0.38
| Median |  | 0.4015 | 0.42 | 0.38495 | 0.395 | 0.365 | 0.375 | 0.355 | 0.39 | 0.35975 | 0.4 | 0.3855 | 0.395 | 0.385 | 0.318 | 0.385 | 0.4 | 0.39 | 0.38
| Std Dev. |  | 0.0067946 | 0.019664 | 0.011455 | 0.0054772 | 0.017889 | 0.018708 | 0.01169 | 0.012247 | 0.0054818 | 0.040825 | 0.0041833 | 0.019408 | 0.017889 | 0.013303 | 0.0054772 | 0.0040825 | 0.013663 | 0.0063246
| Rel.Std.Dev. |  | 1.683% | 4.719% | 2.972% | 1.387% | 4.969% | 4.989% | 3.262% | 3.101% | 1.524% | 10.65% | 1.088% | 4.998% | 4.587% | 4.199% | 1.423% | 1.025% | 3.474% | 1.664%
| PDM3 |  | 4.708% | 8.036% | -0.04982% | 2.418% | -6.657% | -2.768% | -7.089% | 2.418% | -6.718% | -0.6073% | -0.3048% | 0.6891% | 1.121% | -17.85% | -0.1751% | 3.282% | 1.986% | -1.472%
| Z-Score (Absolute) |  | 1 | 2.03 | 0.09 | 0.56 | 1.82 | 0.8 | 2.22 | 0.56 | 1.83 | 0.9 | 0.15 | 0.11 | 0.22 | 4.75 | 0.12 | 0.79 | 0.09 | 0.46
| Analytical results for Sb in OREAS 922 (Certified Value 1.35 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 1.3 | 1.66 | 1.36 | 0.6 | 0.7 | 1.34 | 1.21 | 1.4 | 1.34 | < 5 | 1.3208 | 1.32 | < 5 | 1.537 | < 5 | 0.4 | 1.3 | 1.41 | 1.29 |  | 1
| 1 | 2 | 1.3 | 1.65 | 1.26 | 1 | 0.8 | 1.37 | 1.22 | 1.4 | 1.24 | < 5 | 1.3523 | 1.34 | < 5 | 1.536 | < 5 | 0.8 | 1.4 | 1.43 | 1.46 |  | #N/A
| 1 | 3 | 1.3 | 1.72 | 1.36 | 1.1 | 0.9 | 1.36 | 1.2 | 1.4 | 1.34 | < 5 | 1.3066 | 1.34 | < 5 | 1.486 | 7 | 0.8 | 1.4 | 1.33 | 1.17 |  | 16
| 1 | 4 | 1.4 | 1.69 | 1.27 | 1 | 1.1 | 1.41 | 1.24 | 1.4 | 1.36 | < 5 | 1.2912 | 1.53 | < 5 | 1.492 | < 5 | 0.3 | 1.4 | 1.38 | 1.34 |  | 1.3523
| 1 | 5 | 1.3 | 1.64 | 1.26 | 1 | 1.1 | 1.3 | 1.26 | 1.4 | 1.31 | < 5 | 1.2633 | 1.42 | < 5 | 1.517 | < 5 | 0.7 | 1.3 | 1.48 | 1.32 |  | 47
| 1 | 6 | 1.3 | 1.67 | 1.28 | 1.1 | 1.1 | 1.3 | 1.15 | 1.6 | 1.39 | < 5 | 1.2611 | 1.51 | < 5 | 1.486 | < 5 | 0.3 | 1.4 | 1.5 | 1.31
| Mean |  | 1.3167 | 1.6717 | 1.2983 | 0.96667 | 0.95 | 1.3467 | 1.2133 | 1.4333 | 1.33 |  | 1.2992 | 1.41 |  | 1.509 | 7 | 0.55 | 1.3667 | 1.4217 | 1.315
| Median |  | 1.3 | 1.665 | 1.275 | 1 | 1 | 1.35 | 1.215 | 1.4 | 1.34 |  | 1.2989 | 1.38 |  | 1.5045 | 7 | 0.55 | 1.4 | 1.42 | 1.315
| Std Dev. |  | 0.040825 | 0.029269 | 0.048339 | 0.18619 | 0.17607 | 0.04274 | 0.037771 | 0.08165 | 0.051381 |  | 0.035081 | 0.092087 |  | 0.024183 |  | 0.2429 | 0.05164 | 0.063061 | 0.093113
| Rel.Std.Dev. |  | 3.101% | 1.751% | 3.723% | 19.26% | 18.53% | 3.174% | 3.113% | 5.696% | 3.863% |  | 2.7% | 6.531% |  | 1.603% |  | 44.16% | 3.779% | 4.436% | 7.081%
| PDM3 |  | -2.637% | 23.61% | -3.992% | -28.52% | -29.75% | -0.4183% | -10.28% | 5.99% | -1.651% |  | -3.928% | 4.265% |  | 11.59% | 417.6% | -59.33% | 1.061% | 5.128% | -2.76%
| Z-Score (Absolute) |  | 0.5 | 3.46 | 0.51 | 3.27 | 4.22 | 0 | 1.42 | 0.57 | 0.01 | NA | 0.51 | 0.67 | NA | 1.73 | 60.19 | 8.48 | 0.21 | 0.8 | 0.34
| Analytical results for Sc in OREAS 922 (Certified Value 13.1 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 12 | 17.8 | 12 | 16 | 13 | 13 | 13 | 12 | 18 | 13.012 | 15.8 | 10 | 13.157 | 14 | 13 | 13.1 | 13 |  |  |  | 1
| 1 | 2 | 13 | 17.7 | 13 | 16 | 14 | 13.1 | 13.2 | 13 | 18 | 13.09 | 15.9 | 9 | 13.184 | 14 | 13 | 13.4 | 13.4 |  |  |  | 9
| 1 | 3 | 12 | 17.6 | 13 | 15 | 14 | 13.3 | 12.9 | 13 | 17 | 12.9542 | 16.5 | 10 | 12.739 | 14 | 13 | 13.1 | 12.9 |  |  |  | 16
| 1 | 4 | 13 | 17.7 | 13 | 15 | 15 | 13.4 | 13.1 | 13 | 18 | 13.0054 | 16.4 | 10 | 13.047 | 14 | 13 | 13.7 | 13 |  |  |  | 13.1346
| 1 | 5 | 12 | 17.5 | 13 | 16 | 15 | 12 | 12.9 | 12 | 19 | 13.0526 | 17.3 | 9 | 13.102 | 14 | 13 | 14.5 | 13.1 |  |  |  | 48
| 1 | 6 | 12 | 17.5 | 13 | 16 | 17 | 12.7 | 12.9 | 12 | 18 | 12.7165 | 16.5 | 9 | 12.729 | 14 | 14 | 13.9 | 13.1
| Mean |  | 12.3333 | 17.6333 | 12.8333 | 15.6667 | 14.6667 | 12.9167 | 13 | 12.5 | 18 | 12.9718 | 16.4 | 9.5 | 12.993 | 14 | 13.1667 | 13.6167 | 13.0833
| Median |  | 12 | 17.65 | 13 | 16 | 14.5 | 13.05 | 12.95 | 12.5 | 18 | 13.0087 | 16.45 | 9.5 | 13.0745 | 14 | 13 | 13.55 | 13.05
| Std Dev. |  | 0.5164 | 0.12111 | 0.40825 | 0.5164 | 1.3663 | 0.51153 | 0.12649 | 0.54772 | 0.63246 | 0.13323 | 0.53666 | 0.54772 | 0.2061 | 0 | 0.40825 | 0.53821 | 0.17224
| Rel.Std.Dev. |  | 4.187% | 0.6868% | 3.181% | 3.296% | 9.315% | 3.96% | 0.973% | 4.382% | 3.514% | 1.027% | 3.272% | 5.766% | 1.586% | 0% | 3.101% | 3.953% | 1.316%
| PDM3 |  | -6.1% | 34.25% | -2.293% | 19.28% | 11.66% | -1.659% | -1.025% | -4.831% | 37.04% | -1.239% | 24.86% | -27.67% | -1.078% | 6.589% | 0.2444% | 3.67% | -0.3901%
| Z-Score (Absolute) |  | 0.73 | 5.07 | 0 | 2.92 | 1.31 | 0.09 | 0 | 0.55 | 5.47 | 0.03 | 3.72 | 3.83 | 0.01 | 1.09 | 0 | 0.67 | 0.09
| Analytical results for Se in OREAS 922 (Certified Value 3.76 ppm) | [21 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 5 | < 2 | 3.1 | 4.2 | 3.5 | 4 | 4 | < 5 | 4.1 | 3.34 | < 2 | 2.7 | < 5 | 3 | 4 | [5 empty] | 1
| 1 | 2 | 5 | < 2 | 2.6 | 4.3 | 3.8 | 5 | 4 | < 5 | 3.1 | 3.1138 | 3.239 | 2.9 | < 5 | 4 | 5 | [5 empty] | 10
| 1 | 3 | 5 | < 2 | 2.7 | 4.8 | 3.9 | 4 | 4 | < 5 | 3.7 | 3.6905 | 2.017 | 2.8 | < 5 | 3 | 4 | [5 empty] | 16
| 1 | 4 | 5 | < 2 | 2.7 | 4.1 | 3.9 | 6 | 4 | < 5 | 3.7 | 3.5522 | 2.834 | 3.1 | < 5 | 3 | 5 | [5 empty] | 3.7634
| 1 | 5 | 4 | < 2 | 2.5 | 4.3 | 4 | 4 | 4 | < 5 | 3.3 | 3.4429 | < 2 | 2 | < 5 | 4 | 4 | [5 empty] | 49
| 1 | 6 | 5 | < 2 | 2.9 | 4.2 | 3.9 | 5 | 4 | < 5 | 4 | 3.4 | 3.515 | 2.4 | < 5 | 4 | 5
| Mean |  | 4.8333 |  | 2.75 | 4.3167 | 3.8333 | 4.6667 | 4 |  | 3.65 | 3.4232 | 2.9013 | 2.65 |  | 3.5 | 4.5
| Median |  | 5 |  | 2.7 | 4.25 | 3.9 | 4.5 | 4 |  | 3.7 | 3.4214 | 3.0365 | 2.75 |  | 3.5 | 4.5
| Std Dev. |  | 0.40825 |  | 0.21679 | 0.24833 | 0.17512 | 0.8165 | 0 |  | 0.38859 | 0.19579 | 0.65248 | 0.3937 |  | 0.54772 | 0.54772
| Rel.Std.Dev. |  | 8.447% |  | 7.883% | 5.753% | 4.568% | 17.5% | 0% |  | 10.65% | 5.719% | 22.49% | 14.86% |  | 15.65% | 12.17%
| PDM3 |  | 28.43% |  | -26.93% | 14.7% | 1.857% | 24% | 6.286% |  | -3.014% | -9.04% | -22.91% | -29.59% |  | -7% | 19.57%
| Z-Score (Absolute) |  | 1.41 | NA | 1.18 | 0.51 | 0.14 | 1.03 | 0.26 | NA | 0.14 | 0.41 | 1.01 | 1.3 | NA | 0.32 | 0.84
| Analytical results for Sm in OREAS 922 (Certified Value 6.95 ppm) | [21 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 10: 4A*MS |  | 2
3
| 1 | 1 | 01: 9; 02: 6.76; 03: 7.5; 04: 6.3; 05: 6.75; 06: 7.5945; 07: 7.24; 09: 6.4; 10: 7.1 |  | 1
| 1 | 2 | 01: 9; 02: 6.72; 03: 7.4; 04: 6.4; 05: 6.8; 06: 7.2896; 07: 7.12; 09: 6.6; 10: 7.3 |  | #N/A
| 1 | 3 | 01: 8.9; 02: 6.77; 03: 7.2; 04: 6.5; 05: 6.9; 06: 7.5321; 07: 7.29; 09: 6.6; 10: 6.7 |  | 16
| 1 | 4 | 01: 8.9; 02: 6.96; 03: 7.2; 04: 6.2; 05: 6.9; 06: 7.4031; 07: 7.43; 09: 6.6; 10: 7 |  | 6.954
| 1 | 5 | 01: 8.9; 02: 6.92; 03: 7.2; 04: 6.2; 05: 6.8; 06: 7.3933; 07: 7.44; 09: 6.7; 10: 7 |  | 50
| 1 | 6 | 01: 8.9; 02: 6.91; 03: 7.4; 04: 6.3; 05: 6.95; 06: 7.3511; 07: 7.35; 09: 6.4; 10: 7.8
| Mean |  | 01: 8.9333; 02: 6.84; 03: 7.3167; 04: 6.3167; 05: 6.85; 06: 7.4273; 07: 7.3117; 09: 6.55; 10: 7.15
| Median |  | 01: 8.9; 02: 6.84; 03: 7.3; 04: 6.3; 05: 6.85; 06: 7.3982; 07: 7.32; 09: 6.6; 10: 7.05
| Std Dev. |  | 01: 0.05164; 02: 0.10139; 03: 0.13292; 04: 0.1169; 05: 0.07746; 06: 0.11443; 07: 0.12189; 09: 0.12247; 10: 0.37283
| Rel.Std.Dev. |  | 01: 0.5781%; 02: 1.482%; 03: 1.817%; 04: 1.851%; 05: 1.131%; 06: 1.541%; 07: 1.667%; 09: 1.87%; 10: 5.214%
| PDM3 |  | 01: 28.46%; 02: -1.64%; 03: 5.215%; 04: -9.165%; 05: -1.496%; 06: 6.806%; 07: 5.143%; 09: -5.81%; 10: 2.818%
| Z-Score (Absolute) |  | 01: 4.35; 02: 0.41; 03: 0.67; 04: 1.6; 05: 0.39; 06: 0.93; 07: 0.66; 09: 1.07; 10: 0
| Analytical results for Sn in OREAS 922 (Certified Value 9.95 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 9.4 | 13.1 | 10.3 | 10 | 9 | 9.8 | 8.7 | 10 | 9.6 | 12 | 9.9183 | 13.1 | 12 | 10.205 | 13 | 10 | 9.6 | 9.9 | 10.6 |  | 1
| 1 | 2 | 9 | 13.1 | 10.3 | 10 | 10 | 9.7 | 8.9 | 9 | 9.1 | 12 | 9.4478 | 12.4 | 10 | 10.384 | < 10 | 11 | 9.5 | 10.1 | 10.7 |  | #N/A
| 1 | 3 | 9.1 | 13 | 10.2 | 10 | 10 | 9.7 | 9.1 | 10 | 9.3 | 12 | 9.5 | 12.8 | 11 | 9.904 | 12 | 10 | 9.6 | 9.8 | 10.4 |  | 16
| 1 | 4 | 8.9 | 13.2 | 10.2 | 10 | 9 | 10.1 | 9 | 10 | 9.4 | 12 | 9.7374 | 13.2 | 10 | 9.92 | 11 | 9 | 9.6 | 10.2 | 10.6 |  | 9.9538
| 1 | 5 | 9.2 | 13 | 10.3 | 10 | 9 | 8.8 | 9.2 | 10 | 9.2 | 12 | 9.8228 | 13.4 | 11 | 10.056 | 12 | 12 | 9.6 | 10.6 | 10.4 |  | 51
| 1 | 6 | 9 | 13 | 10.3 | 10 | 10 | 9.5 | 8.8 | 10 | 9.8 | 12 | 9.6843 | 14.5 | 10 | 9.624 | 10 | 10 | 9.8 | 10.3 | 10.7
| Mean |  | 9.1 | 13.0667 | 10.2667 | 10 | 9.5 | 9.6 | 8.95 | 9.8333 | 9.4 | 12 | 9.6851 | 13.2333 | 10.6667 | 10.0155 | 11.6 | 10.3333 | 9.6167 | 10.15 | 10.5667
| Median |  | 9.05 | 13.05 | 10.3 | 10 | 9.5 | 9.7 | 8.95 | 10 | 9.35 | 12 | 9.7109 | 13.15 | 10.5 | 9.988 | 12 | 10 | 9.6 | 10.15 | 10.6
| Std Dev. |  | 0.17889 | 0.08165 | 0.05164 | 0 | 0.54772 | 0.43818 | 0.18708 | 0.40825 | 0.26077 | 0 | 0.18256 | 0.71181 | 0.8165 | 0.26388 | 1.1402 | 1.0328 | 0.098319 | 0.2881 | 0.13663
| Rel.Std.Dev. |  | 1.966% | 0.6249% | 0.503% | 0% | 5.766% | 4.564% | 2.09% | 4.152% | 2.774% | 0% | 1.885% | 5.379% | 7.655% | 2.635% | 9.829% | 9.995% | 1.022% | 2.838% | 1.293%
| PDM3 |  | -8.578% | 31.27% | 3.143% | 0.4643% | -4.559% | -3.554% | -10.08% | -1.21% | -5.564% | 20.56% | -2.699% | 32.95% | 7.162% | 0.62% | 16.54% | 3.813% | -3.387% | 1.971% | 6.157%
| Z-Score (Absolute) |  | 1.2 | 3.99 | 0.33 | 0.02 | 0.67 | 0.33 | 1.39 | 0.02 | 0.81 | 2.6 | 0.43 | 4.21 | 0.85 | 0 | 1.61 | 0.42 | 0.52 | 0.18 | 0.72
| Analytical results for Sr in OREAS 922 (Certified Value 58 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 61.6 | 52.7 | 58.52 | 58.3 | 50 | 59.7 | 56 | 57 | 63 | 57.6333 | 66.41 | 77 | 53.674 | 76 | 58 | 56 | 55.6 | 61.6 |  |  | 1
| 1 | 2 | 61.9 | 54 | 57.72 | 59.1 | 54 | 59 | 58.4 | 55.5 | 63 | 58.137 | 63.69 | 75 | 53.109 | 84 | 59 | 56 | 56.1 | 63.9 |  |  | #N/A
| 1 | 3 | 61.1 | 51.5 | 57.31 | 57.8 | 57 | 60.6 | 56.3 | 56.5 | 57 | 58.527 | 64.93 | 77 | 52.899 | 78 | 57 | 57 | 54.9 | 59.2 |  |  | 16
| 1 | 4 | 59.6 | 51.3 | 56.75 | 57.3 | 58 | 61.6 | 55.9 | 57.5 | 58 | 58.9365 | 64.43 | 76 | 53.826 | 77 | 58 | 57 | 56.3 | 60.6 |  |  | 57.5915
| 1 | 5 | 62.6 | 53.1 | 57.08 | 58.5 | 58 | 54.7 | 56.2 | 56.5 | 64 | 58.8 | 67.62 | 72 | 54.522 | 75 | 59 | 57 | 59.8 | 62.1 |  |  | 52
| 1 | 6 | 62.4 | 53 | 57.47 | 58.6 | 62 | 57 | 55.4 | 57 | 65 | 58.6845 | 70.37 | 74 | 54.16 | 68 | 57 | 57 | 56.7 | 62.8
| Mean |  | 61.5333 | 52.6 | 57.475 | 58.2667 | 56.5 | 58.7667 | 56.3667 | 56.6667 | 61.6667 | 58.4531 | 66.2417 | 75.1667 | 53.6983 | 76.3333 | 58 | 56.6667 | 56.5667 | 61.7
| Median |  | 61.75 | 52.85 | 57.39 | 58.4 | 57.5 | 59.35 | 56.1 | 56.75 | 63 | 58.6058 | 65.67 | 75.5 | 53.75 | 76.5 | 58 | 57 | 56.2 | 61.85
| Std Dev. |  | 1.0912 | 1.0276 | 0.61007 | 0.63456 | 4.0866 | 2.5288 | 1.0443 | 0.68313 | 3.3267 | 0.48682 | 2.4692 | 1.9408 | 0.61558 | 5.164 | 0.89443 | 0.5164 | 1.7014 | 1.6541
| Rel.Std.Dev. |  | 1.773% | 1.954% | 1.061% | 1.089% | 7.233% | 4.303% | 1.853% | 1.206% | 5.395% | 0.8328% | 3.728% | 2.582% | 1.146% | 6.765% | 1.542% | 0.9113% | 3.008% | 2.681%
| PDM3 |  | 6.844% | -8.667% | -0.2024% | 1.172% | -1.895% | 2.04% | -2.127% | -1.606% | 7.076% | 1.496% | 15.02% | 30.52% | -6.76% | 32.54% | 0.7092% | -1.606% | -1.78% | 7.134%
| Z-Score (Absolute) |  | 1.05 | 1.7 | 0.2 | 0.04 | 0.5 | 0.19 | 0.67 | 0.45 | 1.09 | 0.1 | 2.49 | 5.24 | 1.36 | 5.6 | 0.04 | 0.45 | 0.68 | 1.1
| Analytical results for Ta in OREAS 922 (Certified Value 1.24 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 0.76 | 4.06 | 1.55 | 0.1 | < 0.1 | 1.21 | 1.18 | 1.2 | 1.0374 | 1.15 | < 5 | 1.544 | < 0.1 | 1.5 | 1.29 | 1.26 |  |  |  |  | 1
| 1 | 2 | 0.75 | 4.13 | 1.15 | 0.7 | < 0.1 | 1.2 | 1.19 | 1.2 | 1.1031 | 1.15 | < 5 | 1.532 | < 0.1 | 1.5 | 1.29 | 1.27 |  |  |  |  | #N/A
| 1 | 3 | 0.74 | 4.05 | 0.98 | 0.8 | 0.5 | 1.22 | 1.22 | 1.1 | 0.98565 | 1.19 | < 5 | 1.471 | < 0.1 | 1.4 | 1.29 | 1.23 |  |  |  |  | 16
| 1 | 4 | 0.72 | 4.08 | 1.01 | 0.1 | 0.4 | 1.21 | 1.2 | 1.2 | 1.2102 | 1.2 | < 5 | 1.511 | < 0.1 | 1.3 | 1.28 | 1.26 |  |  |  |  | 1.2412
| 1 | 5 | 0.76 | 4.09 | 1.02 | < 0.1 | 0.4 | 1.11 | 1.18 | 1.2 | 1.0855 | 1.29 | < 5 | 1.502 | < 0.1 | 1.4 | 1.36 | 1.22 |  |  |  |  | 53
| 1 | 6 | 0.76 | 4.15 | 1.06 | 0.5 | 0.8 | 1.17 | 1.2 | 1.2 | 1.2362 | 1.23 | < 5 | 1.514 | < 0.1 | 1.4 | 1.28 | 1.23
| Mean |  | 0.74833 | 4.0933 | 1.1283 | 0.44 | 0.525 | 1.1867 | 1.195 | 1.1833 | 1.1097 | 1.2017 |  | 1.5123 |  | 1.4167 | 1.2983 | 1.245
| Median |  | 0.755 | 4.085 | 1.04 | 0.5 | 0.45 | 1.205 | 1.195 | 1.2 | 1.0943 | 1.195 |  | 1.5125 |  | 1.4 | 1.29 | 1.245
| Std Dev. |  | 0.016021 | 0.039328 | 0.21479 | 0.32863 | 0.1893 | 0.041312 | 0.015166 | 0.040825 | 0.097317 | 0.053072 |  | 0.025335 |  | 0.075277 | 0.030605 | 0.020736
| Rel.Std.Dev. |  | 2.141% | 0.9608% | 19.04% | 74.69% | 36.06% | 3.481% | 1.269% | 3.45% | 8.77% | 4.417% |  | 1.675% |  | 5.314% | 2.357% | 1.666%
| PDM3 |  | -39.71% | 229.8% | -9.096% | -64.55% | -57.7% | -4.396% | -3.725% | -4.665% | -10.6% | -3.188% |  | 21.84% |  | 14.13% | 4.6% | 0.3034%
| Z-Score (Absolute) |  | 2.46 | 15.73 | 0.85 | 4.14 | 3.68 | 0.01 | 0.03 | 0 | 0.5 | 0 | NA | 1.69 | NA | 1.17 | 0.46 | 0.24
| Analytical results for Tb in OREAS 922 (Certified Value 0.93 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 1.15 | 0.926 | 1 | 0.8 | 0.92 | 0.925 | 0.885 | 0.987 | 0.9 | 0.9 | 1.07 | [9 empty] | 1
| 1 | 2 | 1.14 | 0.95 | 1 | 0.8 | 0.92 | 0.93372 | 0.85 | 0.989 | 0.9 | 0.9 | 1.07 | [9 empty] | #N/A
| 1 | 3 | 1.14 | 0.94 | 1 | 0.9 | 0.92 | 0.91374 | 0.909 | 0.967 | 0.9 | 0.9 | 1.06 | [9 empty] | 16
| 1 | 4 | 1.14 | 0.957 | 1 | 0.8 | 0.9 | 0.93552 | 0.902 | 0.978 | 0.9 | 0.9 | 1.08 | [9 empty] | 0.9335
| 1 | 5 | 1.13 | 0.949 | 1 | 0.8 | 0.88 | 0.93496 | 0.927 | 0.967 | 0.9 | 0.9 | 1.13 | [9 empty] | 54
| 1 | 6 | 1.14 | 0.978 | 1 | 0.8 | 0.92 | 0.93517 | 0.948 | 0.961 | 0.9 | 0.9 | 1.07
| Mean |  | 1.14 | 0.95 | 1 | 0.81667 | 0.91 | 0.92968 | 0.9035 | 0.97483 | 0.9 | 0.9 | 1.08
| Median |  | 1.14 | 0.9495 | 1 | 0.8 | 0.92 | 0.93434 | 0.9055 | 0.9725 | 0.9 | 0.9 | 1.07
| Std Dev. |  | 0.0063246 | 0.017378 | 0 | 0.040825 | 0.016733 | 0.008766 | 0.033969 | 0.0116 | 0 | 0 | 0.025298
| Rel.Std.Dev. |  | 0.5548% | 1.829% | 0% | 4.999% | 1.839% | 0.9429% | 3.76% | 1.19% | 0% | 0% | 2.342%
| PDM3 |  | 22.12% | 1.767% | 7.123% | -12.52% | -2.518% | -0.4089% | -3.214% | 4.428% | -3.589% | -3.589% | 15.69%
| Z-Score (Absolute) |  | 4.78 | 0.46 | 1.6 | 2.95 | 0.45 | 0 | 0.59 | 1.03 | 0.67 | 0.67 | 3.19
| Analytical results for Te in OREAS 922 (Certified Value < 0.05 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.1 | 1.55 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.2 | < 0.05 | 0.3 | < 0.05 | < 0.1 | < 1 | < 0.05 | < 0.05 | [5 empty] | 1
| 1 | 2 | < 0.1 | 1.53 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.2 | < 0.05 | 0.19 | 0.115 | < 0.1 | < 1 | < 0.05 | 0.06 | [5 empty] | 8
| 1 | 3 | < 0.1 | 1.58 | < 0.05 | < 0.1 | < 0.1 | 0.05 | < 0.05 | < 0.2 | < 0.05 | 0.09 | 0.075 | < 0.1 | < 1 | < 0.05 | < 0.05 | [5 empty] | 16
| 1 | 4 | < 0.1 | 1.59 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.2 | < 0.05 | 0.05 | < 0.05 | < 0.1 | < 1 | 0.05 | 0.06 | [5 empty] | < 0.05
| 1 | 5 | < 0.1 | 1.61 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | 0.1 | < 0.2 | < 0.05 | 0.34 | < 0.05 | < 0.1 | < 1 | < 0.05 | < 0.05 | [5 empty] | 55
| 1 | 6 | < 0.1 | 1.53 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.2 | < 0.05 | 0.25 | 0.145 | < 0.1 | < 1 | < 0.05 | < 0.05
| Mean |  | 02: 1.565; 06: 0.05; 07: 0.1; 14: 0.20333; 15: 0.11167; 19: 0.05; 21: 0.06
| Median |  | 02: 1.565; 06: 0.05; 07: 0.1; 14: 0.22; 15: 0.115; 19: 0.05; 21: 0.06
| Std Dev. |  | 02: 0.033317; 14: 0.11553; 15: 0.035119; 21: 0
| Rel.Std.Dev. |  | 02: 2.129%; 14: 56.82%; 15: 31.45%; 21: 0%
PDM3
| Z-Score (Absolute) |  | NA | 88.33 | NA | NA | NA | 0.67 | 0.67 | NA | NA | 9.63 | 0.94 | NA | NA | 0.67 | 0
| Analytical results for Th in OREAS 922 (Certified Value 17.2 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 18.9 | 21.2 | 16.72 | 18.6 | 14.7 | 17.4 | 17.3 | 16.1 | 18 | 16.615 | 17.24 | 17.451 | 15.6 | 18.15 | 18.2 | 17.5 |  |  |  |  | 1
| 1 | 2 | 18.3 | 21.1 | 16.35 | 19.2 | 15.3 | 17.3 | 17.2 | 16.8 | 17 | 16.6231 | 16.97 | 17.179 | 16 | 18.05 | 18.1 | 18.1 |  |  |  |  | #N/A
| 1 | 3 | 18.4 | 21 | 16.26 | 18.7 | 15.7 | 17.3 | 16.8 | 15.7 | 16 | 16.6121 | 17.43 | 16.776 | 15.9 | 17.22 | 17.9 | 16.7 |  |  |  |  | 16
| 1 | 4 | 18 | 21 | 16.53 | 18.7 | 14.9 | 17.5 | 17.1 | 16 | 18 | 16.4383 | 17.4 | 17.877 | 15.7 | 17.74 | 18.4 | 17.5 |  |  |  |  | 17.1647
| 1 | 5 | 18.7 | 20.9 | 16.56 | 18.5 | 15.1 | 15.6 | 16.2 | 15.6 | 19 | 16.5975 | 17.73 | 17.42 | 15.8 | 17.61 | 19.5 | 17.4 |  |  |  |  | 56
| 1 | 6 | 18.6 | 20.9 | 17.33 | 18.6 | 15.4 | 16.3 | 17 | 16.2 | 19 | 16.4993 | 18.01 | 17.024 | 15.2 | 18.71 | 18.1 | 17.2
| Mean |  | 18.4833 | 21.0167 | 16.625 | 18.7167 | 15.1833 | 16.9 | 16.9333 | 16.0667 | 17.8333 | 16.5642 | 17.4633 | 17.2878 | 15.7 | 17.9133 | 18.3667 | 17.4
| Median |  | 18.5 | 21 | 16.545 | 18.65 | 15.2 | 17.3 | 17.05 | 16.05 | 18 | 16.6048 | 17.415 | 17.2995 | 15.75 | 17.895 | 18.15 | 17.45
| Std Dev. |  | 0.31885 | 0.1169 | 0.38161 | 0.24833 | 0.36009 | 0.77201 | 0.39833 | 0.4274 | 1.169 | 0.076843 | 0.3656 | 0.38331 | 0.28284 | 0.51228 | 0.5785 | 0.45607
| Rel.Std.Dev. |  | 1.725% | 0.5562% | 2.295% | 1.327% | 2.372% | 4.568% | 2.352% | 2.66% | 6.555% | 0.4639% | 2.094% | 2.217% | 1.802% | 2.86% | 3.15% | 2.621%
| PDM3 |  | 7.682% | 22.44% | -3.144% | 9.042% | -11.54% | -1.542% | -1.348% | -6.397% | 3.895% | -3.498% | 1.74% | 0.7174% | -8.533% | 4.362% | 7.003% | 1.371%
| Z-Score (Absolute) |  | 1.03 | 3.31 | 0.65 | 1.24 | 1.94 | 0.17 | 0.37 | 1.15 | 0.44 | 0.7 | 0.11 | 0.05 | 1.48 | 0.51 | 0.72 | 0.05
| Analytical results for Ti in OREAS 922 (Certified Value 0.427 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.457 | 0.4 | 0.434 | 0.36 | 0.32 | 0.428 | 0.404 | 0.44 | 0.43 | 0.44764 | 0.48 | 0.398 | 0.3423 | 0.35 | 0.44 | 0.4 | 0.424 |  |  |  | 1
| 1 | 2 | 0.461 | 0.42 | 0.5652 | 0.36 | 0.38 | 0.412 | 0.426 | 0.44 | 0.41 | 0.44722 | 0.46 | 0.436 | 0.3713 | 0.44 | 0.43 | 0.4 | 0.434 |  |  |  | #N/A
| 1 | 3 | 0.459 | 0.4 | 0.4397 | 0.37 | 0.43 | 0.434 | 0.403 | 0.43 | 0.41 | 0.43397 | 0.46 | 0.415 | 0.352 | 0.43 | 0.44 | 0.39 | 0.422 |  |  |  | 16
| 1 | 4 | 0.454 | 0.4 | 0.4519 | 0.47 | 0.44 | 0.443 | 0.41 | 0.45 | 0.42 | 0.44384 | 0.47 | 0.419 | 0.3463 | 0.45 | 0.43 | 0.41 | 0.416 |  |  |  | 0.42731
| 1 | 5 | 0.443 | 0.41 | 0.4518 | 0.45 | 0.44 | 0.38 | 0.416 | 0.44 | 0.43 | 0.44716 | 0.44 | 0.437 | 0.3358 | 0.41 | 0.43 | 0.43 | 0.421 |  |  |  | 57
| 1 | 6 | 0.461 | 0.41 | 0.4489 | 0.43 | 0.45 | 0.404 | 0.408 | 0.45 | 0.45 | 0.4451 | 0.46 | 0.393 | 0.3262 | 0.28 | 0.44 | 0.4 | 0.428
| Mean |  | 0.45583 | 0.40667 | 0.46525 | 0.40667 | 0.41 | 0.41683 | 0.41117 | 0.44167 | 0.425 | 0.44416 | 0.46167 | 0.41633 | 0.34565 | 0.39333 | 0.435 | 0.405 | 0.42417
| Median |  | 0.458 | 0.405 | 0.45035 | 0.4 | 0.435 | 0.42 | 0.409 | 0.44 | 0.425 | 0.44613 | 0.46 | 0.417 | 0.3443 | 0.42 | 0.435 | 0.4 | 0.423
| Std Dev. |  | 0.006824 | 0.008165 | 0.049489 | 0.049261 | 0.050596 | 0.023017 | 0.0086352 | 0.0075277 | 0.015166 | 0.0051998 | 0.013292 | 0.018457 | 0.015396 | 0.065929 | 0.0054772 | 0.013784 | 0.0062102
| Rel.Std.Dev. |  | 1.497% | 2.008% | 10.64% | 12.11% | 12.34% | 5.522% | 2.1% | 1.704% | 3.568% | 1.171% | 2.879% | 4.433% | 4.454% | 16.76% | 1.259% | 3.403% | 1.464%
| PDM3 |  | 6.675% | -4.831% | 8.878% | -4.831% | -4.051% | -2.452% | -3.778% | 3.359% | -0.5409% | 3.942% | 8.04% | -2.569% | -19.11% | -7.952% | 1.799% | -5.221% | -0.7359%
| Z-Score (Absolute) |  | 1.32 | 0.67 | 0.81 | 0.67 | 0.15 | 0.28 | 0.5 | 0.67 | 0.03 | 0.77 | 1.44 | 0.3 | 3.03 | 0.31 | 0.42 | 0.93 | 0
| Analytical results for Tl in OREAS 922 (Certified Value 0.85 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.94 | 0.88 | 0.81 | 1.01 | 0.75 | 0.9 | 0.83 | 0.9 | 0.8199 | 0.84 | 0.86 | 2 | 0.8 | 0.89 | 0.88 | 0.84 |  |  |  |  | 1
| 1 | 2 | 0.92 | 0.88 | 0.83 | 1.05 | 0.77 | 0.89 | 0.83 | 0.8 | 0.82508 | 0.85 | 0.843 | < 2 | 0.79 | 0.94 | 0.87 | 0.84 |  |  |  |  | #N/A
| 1 | 3 | 0.89 | 0.87 | 0.81 | 1 | 0.77 | 0.93 | 0.84 | 0.9 | 0.80206 | 0.88 | 0.82 | < 2 | 0.79 | 0.89 | 0.88 | 0.77 |  |  |  |  | 16
| 1 | 4 | 0.92 | 0.87 | 0.82 | 1.01 | 0.75 | 0.91 | 0.82 | 0.9 | 0.81913 | 0.85 | 0.843 | 4 | 0.74 | 0.92 | 0.88 | 0.75 |  |  |  |  | 0.84713
| 1 | 5 | 0.92 | 0.87 | 0.81 | 1.01 | 0.75 | 0.8 | 0.82 | 0.8 | 0.87817 | 0.96 | 0.836 | < 2 | 0.8 | 0.89 | 0.93 | 0.81 |  |  |  |  | 58
| 1 | 6 | 0.91 | 0.87 | 0.86 | 1 | 0.79 | 0.87 | 0.8 | 0.8 | 0.79349 | 0.88 | 0.829 | < 2 | 0.81 | 0.91 | 0.9 | 0.81
| Mean |  | 0.91667 | 0.87333 | 0.82333 | 1.0133 | 0.76333 | 0.88333 | 0.82333 | 0.85 | 0.82297 | 0.87667 | 0.8385 | 3 | 0.78833 | 0.90667 | 0.89 | 0.80333
| Median |  | 0.92 | 0.87 | 0.815 | 1.01 | 0.76 | 0.895 | 0.825 | 0.85 | 0.81951 | 0.865 | 0.8395 | 3 | 0.795 | 0.9 | 0.88 | 0.81
| Std Dev. |  | 0.01633 | 0.005164 | 0.019664 | 0.018619 | 0.01633 | 0.045461 | 0.013663 | 0.054772 | 0.029606 | 0.044121 | 0.013722 | 1.4142 | 0.024833 | 0.020656 | 0.021909 | 0.036697
| Rel.Std.Dev. |  | 1.781% | 0.5913% | 2.388% | 1.837% | 2.139% | 5.146% | 1.659% | 6.444% | 3.597% | 5.033% | 1.637% | 47.14% | 3.15% | 2.278% | 2.462% | 4.568%
| PDM3 |  | 8.209% | 3.093% | -2.809% | 19.62% | -9.892% | 4.274% | -2.809% | 0.3389% | -2.852% | 3.487% | -1.019% | 254.1% | -6.941% | 7.028% | 5.061% | -5.17%
| Z-Score (Absolute) |  | 1.04 | 0.31 | 0.53 | 2.67 | 1.55 | 0.76 | 0.53 | 0.08 | 0.54 | 0.08 | 0.28 | 36.16 | 1.12 | 0.87 | 0.59 | 0.87
| Analytical results for Tm in OREAS 922 (Certified Value 0.44 ppm) | [21 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 13: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 13: 19 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 09: 4A*MS; 13: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 01: 0.52; 02: 0.39; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.4273; 07: 0.45; 09: 0.4; 13: 0.4 |  |  |  |  | 1
| 1 | 2 | 01: 0.52; 02: 0.41; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.43866; 07: 0.46; 09: 0.5; 13: 0.4 |  |  |  |  | #N/A
| 1 | 3 | 01: 0.52; 02: 0.39; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.43336; 07: 0.46; 09: 0.4; 13: 0.5 |  |  |  |  | 16
| 1 | 4 | 01: 0.52; 02: 0.41; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.41948; 07: 0.44; 09: 0.4; 13: 0.4 |  |  |  |  | 0.44313
| 1 | 5 | 01: 0.52; 02: 0.4; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.42574; 07: 0.47; 09: 0.4; 13: 0.4 |  |  |  |  | 59
| 1 | 6 | 01: 0.52; 02: 0.41; 03: 0.5; 04: 0.4; 05: 0.46; 06: 0.42435; 07: 0.49; 09: 0.4; 13: 0.4
| Mean |  | 01: 0.52; 02: 0.40167; 03: 0.5; 04: 0.4; 05: 0.47667; 06: 0.42815; 07: 0.46167; 09: 0.41667; 13: 0.41667
| Median |  | 01: 0.52; 02: 0.405; 03: 0.5; 04: 0.4; 05: 0.48; 06: 0.42652; 07: 0.46; 09: 0.4; 13: 0.4
| Std Dev. |  | 01: 0; 02: 0.0098319; 03: 0; 04: 6.0809e-17; 05: 0.008165; 06: 0.0068388; 07: 0.017224; 09: 0.040825; 13: 0.040825
| Rel.Std.Dev. |  | 01: 0%; 02: 2.448%; 03: 0%; 04: 1.52e-14%; 05: 1.713%; 06: 1.597%; 07: 3.731%; 09: 9.798%; 13: 9.798%
| PDM3 |  | 01: 17.35%; 02: -9.356%; 03: 12.83%; 04: -9.733%; 05: 7.569%; 06: -3.38%; 07: 4.184%; 09: -5.971%; 13: -5.971%
| Z-Score (Absolute) |  | 01: 2.2; 02: 0.63; 03: 1.72; 04: 0.67; 05: 1.16; 06: 0; 07: 0.8; 09: 0.67; 13: 0.67
| Analytical results for U in OREAS 922 (Certified Value 3.31 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3.5 | 3.46 | 3.17 | 3.5 | 2.9 | 3.4 | 3.3 | 3.5 | < 20 | 3.171 | 3.37 | 3.217 | 3.1 | 3.42 | 3.43 | 3.3 |  |  |  |  | 1
| 1 | 2 | 3.47 | 3.45 | 3.2 | 3.6 | 2.9 | 3.3 | 3.3 | 3.5 | < 20 | 3.1883 | 3.37 | 3.238 | 3.1 | 3.64 | 3.39 | 3.4 |  |  |  |  | #N/A
| 1 | 3 | 3.23 | 3.42 | 3.19 | 4.7 | 3 | 3.4 | 3.3 | 3.5 | < 20 | 3.1368 | 3.43 | 3.135 | 3.1 | 3.34 | 3.37 | 3 |  |  |  |  | 16
| 1 | 4 | 3.36 | 3.41 | 3.22 | 3.5 | 2.9 | 3.3 | 3.3 | 3.7 | < 20 | 3.13 | 3.4 | 3.359 | 3.1 | 3.48 | 3.46 | 3.2 |  |  |  |  | 3.3055
| 1 | 5 | 3.42 | 3.42 | 3.18 | 3.5 | 2.9 | 3 | 3.2 | 3.5 | < 20 | 3.2353 | 3.44 | 3.291 | 3.1 | 3.3 | 3.59 | 3.3 |  |  |  |  | 60
| 1 | 6 | 3.51 | 3.43 | 3.36 | 3.5 | 3 | 3.2 | 3.3 | 3.4 | < 20 | 3.121 | 3.64 | 3.227 | 3.1 | 3.64 | 3.44 | 3.2
| Mean |  | 3.415 | 3.4317 | 3.22 | 3.7167 | 2.9333 | 3.2667 | 3.2833 | 3.5167 |  | 3.1637 | 3.4417 | 3.2445 | 3.1 | 3.47 | 3.4467 | 3.2333
| Median |  | 3.445 | 3.425 | 3.195 | 3.5 | 2.9 | 3.3 | 3.3 | 3.5 |  | 3.1539 | 3.415 | 3.2325 | 3.1 | 3.45 | 3.435 | 3.25
| Std Dev. |  | 0.10635 | 0.019408 | 0.070711 | 0.48339 | 0.05164 | 0.15055 | 0.040825 | 0.098319 |  | 0.043518 | 0.10147 | 0.075323 | 0 | 0.14574 | 0.077632 | 0.13663
| Rel.Std.Dev. |  | 3.114% | 0.5656% | 2.196% | 13.01% | 1.76% | 4.609% | 1.243% | 2.796% |  | 1.376% | 2.948% | 2.322% | 0% | 4.2% | 2.252% | 4.226%
| PDM3 |  | 3.313% | 3.817% | -2.586% | 12.44% | -11.26% | -1.174% | -0.6701% | 6.389% |  | -4.288% | 4.12% | -1.845% | -6.216% | 4.977% | 4.271% | -2.183%
| Z-Score (Absolute) |  | 0.67 | 0.76 | 0.47 | 1.21 | 1.79 | 0.09 | 0 | 1.01 | NA | 0.61 | 0.61 | 0.2 | 0.94 | 0.96 | 0.84 | 0.26
| Analytical results for V in OREAS 922 (Certified Value 91 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 87 | 101 | 95 | 70 | 77 | 92 | 82 | 90 | 93 | 95 | 108 | 85.952 | 87 | 73 | 90 | 85 | 90 |  |  |  | 1
| 1 | 2 | 90 | 104 | 124 | 91 | 87 | 91 | 88 | 95 | 91 | 99.3878 | 102 | 82.024 | 95 | 92 | 90 | 87 | 92 |  |  |  | #N/A
| 1 | 3 | 89 | 100 | 96 | 92 | 94 | 95 | 83 | 90 | 88 | 95.1498 | 105 | 86.958 | 90 | 90 | 94 | 86 | 89 |  |  |  | 16
| 1 | 4 | 91 | 99 | 97 | 97 | 92 | 95 | 85 | 95 | 93 | 90.5474 | 105 | 90.957 | 89 | 95 | 92 | 89 | 89 |  |  |  | 90.8656
| 1 | 5 | 90 | 100 | 97 | 93 | 92 | 84 | 86 | 90 | 96 | 93.6949 | 100 | 90.48 | 86 | 82 | 93 | 94 | 90 |  |  |  | 61
| 1 | 6 | 91 | 102 | 97 | 95 | 97 | 89 | 85 | 95 | 98 | 90.7628 | 104 | 85.089 | 85 | 48 | 91 | 88 | 91
| Mean |  | 89.6667 | 101 | 101 | 89.6667 | 89.8333 | 91 | 84.8333 | 92.5 | 93.1667 | 94.0905 | 104 | 86.91 | 88.6667 | 80 | 91.6667 | 88.1667 | 90.1667
| Median |  | 90 | 100.5 | 97 | 92.5 | 92 | 91.5 | 85 | 92.5 | 93 | 94.3475 | 104.5 | 86.455 | 88 | 86 | 91.5 | 87.5 | 90
| Std Dev. |  | 1.5055 | 1.7889 | 11.296 | 9.8725 | 7.0828 | 4.1473 | 2.137 | 2.7386 | 3.5449 | 3.2801 | 2.7568 | 3.3833 | 3.6148 | 17.5841 | 1.633 | 3.1885 | 1.169
| Rel.Std.Dev. |  | 1.679% | 1.771% | 11.18% | 11.01% | 7.884% | 4.557% | 2.519% | 2.961% | 3.805% | 3.486% | 2.651% | 3.893% | 4.077% | 21.98% | 1.781% | 3.616% | 1.297%
| PDM3 |  | -1.319% | 11.15% | 11.15% | -1.319% | -1.136% | 0.1479% | -6.639% | 1.799% | 2.532% | 3.549% | 14.45% | -4.353% | -2.42% | -11.96% | 0.8816% | -2.97% | -0.7692%
| Z-Score (Absolute) |  | 0.67 | 3.15 | 1.6 | 0.65 | 0.25 | 0.22 | 2.3 | 0.28 | 0.51 | 0.82 | 4.16 | 1.6 | 1.01 | 0.65 | 0 | 1.18 | 0.51
| Analytical results for W in OREAS 922 (Certified Value 3.81 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 3.8 | 6.7 | 3.3 | 5 | 0.4 | 4.1 | 3.8 | 4 | 6 | 3.7712 | 4.1 | < 10 | 3.685 | 10 | < 5 | 4 | 3.8 | 5.1 |  |  | 1
| 1 | 2 | 3.8 | 6.7 | 3.3 | < 5 | 0.5 | 4.7 | 3.4 | 4.5 | 5 | 3.487 | 3.7 | < 10 | 3.688 | 15 | < 5 | 4 | 4.1 | 4 |  |  | #N/A
| 1 | 3 | 3.7 | 6.5 | 3.3 | 5 | 1.4 | 4.2 | 3.7 | 4 | 5 | 3.4283 | 3.7 | < 10 | 3.778 | < 10 | < 5 | 4 | 3.7 | 3.4 |  |  | 16
| 1 | 4 | 3.8 | 6.5 | 3.2 | 6 | 2 | 4.4 | 3.5 | 4 | < 4 | 3.6751 | 4.7 | < 10 | 3.805 | < 10 | 9 | 4 | 4 | 3.7 |  |  | 3.8074
| 1 | 5 | 3.6 | 6.6 | 3.5 | < 5 | 2.4 | 3.8 | 3.4 | 4 | 5 | 3.7575 | 4 | < 10 | 3.658 | < 10 | < 5 | 4 | 4.2 | 3.6 |  |  | 62
| 1 | 6 | 4 | 6.7 | 3.3 | < 5 | 3 | 4 | 3.4 | 4.5 | 6 | 3.6295 | 3.9 | < 10 | 3.865 | < 10 | < 5 | 4 | 4 | 3.8
| Mean |  | 3.7833 | 6.6167 | 3.3167 | 5.3333 | 1.6167 | 4.2 | 3.5333 | 4.1667 | 5.4 | 3.6248 | 4.0167 |  | 3.7465 | 12.5 | 9 | 4 | 3.9667 | 3.9333
| Median |  | 3.8 | 6.65 | 3.3 | 5 | 1.7 | 4.15 | 3.45 | 4 | 5 | 3.6523 | 3.95 |  | 3.733 | 12.5 | 9 | 4 | 4 | 3.75
| Std Dev. |  | 0.13292 | 0.098319 | 0.098319 | 0.57735 | 1.0439 | 0.31623 | 0.17512 | 0.2582 | 0.54772 | 0.14087 | 0.37103 |  | 0.081846 | 3.5355 |  | 0 | 0.18619 | 0.60553
| Rel.Std.Dev. |  | 3.513% | 1.486% | 2.964% | 10.83% | 64.57% | 7.529% | 4.956% | 6.197% | 10.14% | 3.886% | 9.237% |  | 2.185% | 28.28% |  | 0% | 4.694% | 15.39%
| PDM3 |  | -0.6318% | 73.78% | -12.89% | 40.08% | -57.54% | 10.31% | -7.198% | 9.436% | 41.83% | -4.796% | 5.497% |  | -1.599% | 228.3% | 136.4% | 5.059% | 4.183% | 3.308%
| Z-Score (Absolute) |  | 0.36 | 5.23 | 1.35 | 2.7 | 4.63 | 0.46 | 0.85 | 0.39 | 2.83 | 0.67 | 0.17 | NA | 0.43 | 16.83 | 9.93 | 0.07 | 0 | 0.53
| Analytical results for Y in OREAS 922 (Certified Value 28.9 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 30 | 37.6 | 27.69 | 30 | 27.6 | 30 | 29.2 | 31.1 | 24 | 27.6494 | 30.46 | 20 | 32.068 | 21 | 28 | 28 | 28 | 26.9 |  |  | 1
| 1 | 2 | 30.7 | 37.8 | 27.83 | 30 | 28.5 | 29.7 | 29.9 | 31.5 | 24 | 27.88 | 30.23 | 20 | 32.753 | 22 | 28 | 28 | 29 | 28 |  |  | #N/A
| 1 | 3 | 30.1 | 37.1 | 27.89 | 30 | 28.6 | 29.7 | 28.7 | 31 | 22 | 27.4026 | 31.51 | 20 | 31.903 | 19 | 27 | 28 | 28.3 | 26.4 |  |  | 16
| 1 | 4 | 29.5 | 37.2 | 27.42 | 30 | 27.8 | 30.7 | 29 | 32.3 | 24 | 27.6031 | 31.75 | 20 | 32.54 | 19 | 28 | 28 | 28.8 | 27.4 |  |  | 28.9273
| 1 | 5 | 30.5 | 37.4 | 28.61 | 30 | 27.7 | 27.2 | 28.2 | 30.2 | 25 | 27.7995 | 32.26 | 19 | 32.529 | 17 | 29 | 28 | 30.7 | 27.3 |  |  | 63
| 1 | 6 | 30.3 | 37.7 | 28.18 | 30 | 27.6 | 28.4 | 28.9 | 31.5 | 26 | 27.8926 | 32.9 | 20 | 31.806 | 30 | 28 | 28 | 29.1 | 27.3
| Mean |  | 30.1833 | 37.4667 | 27.9367 | 30 | 27.9667 | 29.2833 | 28.9833 | 31.2667 | 24.1667 | 27.7045 | 31.5183 | 19.8333 | 32.2665 | 21.3333 | 28 | 28 | 28.9833 | 27.2167
| Median |  | 30.2 | 37.5 | 27.86 | 30 | 27.75 | 29.7 | 28.95 | 31.3 | 24 | 27.7245 | 31.63 | 20 | 32.2985 | 20 | 28 | 28 | 28.9 | 27.3
| Std Dev. |  | 0.42151 | 0.28048 | 0.41326 | 0 | 0.45898 | 1.264 | 0.56362 | 0.69474 | 1.3292 | 0.18955 | 1.0289 | 0.40825 | 0.39089 | 4.5898 | 0.63246 | 0 | 0.9411 | 0.53448
| Rel.Std.Dev. |  | 1.396% | 0.7486% | 1.479% | 0% | 1.641% | 4.316% | 1.945% | 2.222% | 5.5% | 0.6842% | 3.264% | 2.058% | 1.211% | 21.51% | 2.259% | 0% | 3.247% | 1.964%
| PDM3 |  | 4.342% | 29.52% | -3.425% | 3.708% | -3.321% | 1.231% | 0.1937% | 8.087% | -16.46% | -4.227% | 8.957% | -31.44% | 11.54% | -26.25% | -3.206% | -3.206% | 0.1937% | -5.914%
| Z-Score (Absolute) |  | 0.82 | 4.35 | 0.27 | 0.73 | 0.25 | 0.38 | 0.24 | 1.34 | 1.89 | 0.38 | 1.47 | 4.11 | 1.83 | 4.31 | 0.24 | 0.24 | 0.24 | 0.62
| Analytical results for Yb in OREAS 922 (Certified Value 2.8 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 3.9 | 2.65 | 2.8 | 2.7 | 2.9 | 2.9637 | 2.87 | 2.673 | 2.6 | 2.9 | 2.8 | [9 empty] | 1
| 1 | 2 | 3.8 | 2.66 | 2.9 | 2.7 | 2.95 | 2.8608 | 2.85 | 2.764 | 2.7 | 2.9 | 2.8 | [9 empty] | #N/A
| 1 | 3 | 3.8 | 2.62 | 2.8 | 2.7 | 2.9 | 2.9954 | 3.02 | 2.66 | 2.6 | 2.8 | 2.8 | [9 empty] | 16
| 1 | 4 | 3.8 | 2.69 | 2.8 | 2.6 | 2.95 | 2.9062 | 3.04 | 2.753 | 2.6 | 2.8 | 2.9 | [9 empty] | 2.7989
| 1 | 5 | 3.8 | 2.65 | 2.9 | 2.6 | 2.85 | 2.9541 | 3.17 | 2.762 | 2.5 | 2.7 | 3 | [9 empty] | 64
| 1 | 6 | 3.8 | 2.71 | 2.8 | 2.7 | 2.8 | 2.9604 | 3.07 | 2.711 | 2.6 | 3 | 2.8
| Mean |  | 3.8167 | 2.6633 | 2.8333 | 2.6667 | 2.8917 | 2.9401 | 3.0033 | 2.7205 | 2.6 | 2.85 | 2.85
| Median |  | 3.8 | 2.655 | 2.8 | 2.7 | 2.9 | 2.9573 | 3.03 | 2.732 | 2.6 | 2.85 | 2.8
| Std Dev. |  | 0.040825 | 0.032042 | 0.05164 | 0.05164 | 0.058452 | 0.048285 | 0.12258 | 0.046211 | 0.063246 | 0.10488 | 0.083666
| Rel.Std.Dev. |  | 1.07% | 1.203% | 1.823% | 1.936% | 2.021% | 1.642% | 4.082% | 1.699% | 2.433% | 3.68% | 2.936%
| PDM3 |  | 36.36% | -4.843% | 1.23% | -4.724% | 3.315% | 5.046% | 7.304% | -2.801% | -7.106% | 1.826% | 1.826%
| Z-Score (Absolute) |  | 5.88 | 1.02 | 0 | 1 | 0.35 | 0.64 | 1.02 | 0.67 | 1.39 | 0.1 | 0.08
| Analytical results for Zn in OREAS 922 (Certified Value 267 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 288 | 261 | 266 | 273 | 263 | 261 | 259 | 279.4226 | 274 | 256 | 230 | 262.6782 | 280 | 245.122 | 230 | 267 | 260 | 259 | 275 |  | 1
| 1 | 2 | 289 | 282 | 266 | 273 | 284 | 263 | 280 | 266.3496 | 268 | 257 | 236 | 259.751 | 280 | 229.699 | 250 | 261 | 267 | 258 | 278 |  | #N/A
| 1 | 3 | 283 | 264 | 267 | 264 | 280 | 283 | 267 | 266.2308 | 266 | 251 | 223 | 252.6816 | 280 | 228.977 | 236 | 260 | 270 | 253 | 266 |  | 16
| 1 | 4 | 291 | 273 | 267 | 262 | 269 | 274 | 273 | 260.6458 | 266 | 260 | 238 | 257.0432 | 280 | 233.526 | 235 | 258 | 260 | 268 | 263 |  | 267.2858
| 1 | 5 | 281 | 270 | 268 | 271 | 275 | 249 | 270 | 262.3491 | 274 | 252 | 237 | 262.7241 | 280 | 239.652 | 223 | 262 | 269 | 271 | 268 |  | 65
| 1 | 6 | 292 | 281 | 272 | 268 | 275 | 259 | 275 | 270.2674 | 262 | 265 | 241 | 261.5774 | 280 | 233.414 | 239 | 250 | 270 | 268 | 275
| Mean |  | 287.3333 | 271.8333 | 267.6667 | 268.5 | 274.3333 | 264.8333 | 270.6667 | 267.5442 | 268.3333 | 256.8333 | 234.1667 | 259.4092 | 280 | 235.065 | 235.5 | 259.6667 | 266 | 262.8333 | 270.8333
| Median |  | 288.5 | 271.5 | 267 | 269.5 | 275 | 262 | 271.5 | 266.2902 | 267 | 256.5 | 236.5 | 260.6642 | 280 | 233.47 | 235.5 | 260.5 | 268 | 263.5 | 271.5
| Std Dev. |  | 4.4121 | 8.612 | 2.2509 | 4.6797 | 7.5277 | 11.9736 | 7.2296 | 6.7267 | 4.8028 | 5.1929 | 6.5549 | 3.9332 | 0 | 6.2151 | 9.0499 | 5.6095 | 4.7749 | 7.1391 | 5.9805
| Rel.Std.Dev. |  | 1.536% | 3.168% | 0.8409% | 1.743% | 2.744% | 4.521% | 2.671% | 2.514% | 1.79% | 2.022% | 2.799% | 1.516% | 0% | 2.644% | 3.843% | 2.16% | 1.795% | 2.716% | 2.208%
| PDM3 |  | 7.5% | 1.701% | 0.1425% | 0.4543% | 2.637% | -0.9175% | 1.265% | 0.09669% | 0.3919% | -3.911% | -12.39% | -2.947% | 4.757% | -12.05% | -11.89% | -2.851% | -0.4811% | -1.666% | 1.327%
| Z-Score (Absolute) |  | 2.83 | 0.61 | 0.02 | 0.14 | 0.97 | 0.39 | 0.45 | 0 | 0.11 | 1.53 | 4.46 | 1.16 | 1.78 | 4.65 | 4.59 | 1.13 | 0.22 | 0.67 | 0.47
| Analytical results for Zr in OREAS 922 (Certified Value 127 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 117 | 118 | 141.6 | 120 | 125 | 132.5 | 126 | 155 | 99 | 126.4298 | 132.8 | 124 | 127.59 | 120 | 124 | 122 | 121 |  |  |  | 1
| 1 | 2 | 110 | 118 | 133.5 | 128 | 129 | 131 | 126.5 | 151 | 97 | 125.1203 | 130.2 | 123 | 127.65 | 139 | 127 | 121 | 126 |  |  |  | 9
| 1 | 3 | 113 | 115 | 125.8 | 120 | 142 | 135 | 126 | 149 | 97 | 124.4462 | 130.8 | 123 | 128.76 | 139 | 122 | 121 | 147 |  |  |  | 16
| 1 | 4 | 112 | 112 | 129.5 | 159 | 143 | 139 | 128 | 152 | 99 | 126.6797 | 130.8 | 127 | 125.98 | 154 | 121 | 129 | 120 |  |  |  | 127.3306
| 1 | 5 | 112 | 118 | 144.9 | 147 | 139 | 120.5 | 127.5 | 146 | 103 | 127.8033 | 136.8 | 120 | 129.92 | 118 | 121 | 130 | 124.5 |  |  |  | 66
| 1 | 6 | 112 | 120 | 139.1 | 138 | 144 | 128 | 126 | 162 | 107 | 125.9049 | 139.6 | 124 | 128.87 | 71 | 122 | 123 | 124
| Mean |  | 112.6667 | 116.8333 | 135.7333 | 135.3333 | 137 | 131 | 126.6667 | 152.5 | 100.3333 | 126.064 | 133.5 | 123.5 | 128.1283 | 123.5 | 122.8333 | 124.3333 | 127.0833
| Median |  | 112 | 118 | 136.3 | 133 | 140.5 | 131.75 | 126.25 | 151.5 | 99 | 126.1673 | 131.8 | 123.5 | 128.205 | 129.5 | 122 | 122.5 | 124.25
| Std Dev. |  | 2.3381 | 2.8577 | 7.3796 | 15.6674 | 8.025 | 6.3482 | 0.8756 | 5.5408 | 3.9328 | 1.1892 | 3.8486 | 2.2583 | 1.3633 | 29.0155 | 2.3166 | 4.0825 | 10.0121
| Rel.Std.Dev. |  | 2.075% | 2.446% | 5.437% | 11.58% | 5.858% | 4.846% | 0.6913% | 3.633% | 3.92% | 0.9434% | 2.883% | 1.829% | 1.064% | 23.49% | 1.886% | 3.283% | 7.878%
| PDM3 |  | -11.52% | -8.244% | 6.599% | 6.285% | 7.594% | 2.882% | -0.5214% | 19.77% | -21.2% | -0.9947% | 4.845% | -3.008% | 0.6265% | -3.008% | -3.532% | -2.354% | -0.1942%
| Z-Score (Absolute) |  | 1.47 | 0.87 | 0.89 | 0.86 | 1.02 | 0.43 | 0 | 2.55 | 2.6 | 0.06 | 0.67 | 0.31 | 0.14 | 0.72 | 0.46 | 0.23 | 0.35
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 922 (Certified Value 0.851 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  | 3
3
| 1 | 1 | 0.88 | 0.84 | 0.9 | 1.18 | 0.84 | 0.91 | 0.95 | 0.95 | 0.8 | 0.86 | 0.68 | 0.91 | 0.82 | 5 | 0.8 | 1.6 | < 1 | 0.84 | 1.05 |  | 1
| 1 | 2 | 0.8 | 0.82 | 0.8 | 0.955 | 0.94 | 1.15 | 0.95 | 0.95 | 0.7 | 0.79614 | 0.47 | 0.83 | 0.969 | 5 | 0.7 | 1.1 | < 1 | 0.85 | 0.85 |  | 8
| 1 | 3 | 0.81 | 0.93 | 0.9 | 0.937 | 1.08 | 0.83 | 0.9 | 0.78 | 0.8 | 0.7628 | 0.8 | 0.86 | 0.79 | 7 | 0.7 | 1.3 | < 1 | 2.03 | 0.87 |  | 16
| 1 | 4 | 0.89 | 0.78 | 0.7 | 0.814 | 0.92 | 0.86 | 1 | 0.85 | 0.8 | 0.84 | 0.69 | 1.14 | 0.81 | 7 | 0.8 | 1.2 | < 1 | 0.84 | 0.98 |  | 0.85127
| 1 | 5 | 0.82 | 0.82 | 1 | 0.894 | 0.94 | 0.83 | 0.9 | 1.31 | 0.8 | 0.85122 | 0.52 | 0.81 | 0.84 | 6 | 0.7 | 1.6 | < 1 | 0.81 | 0.9 |  | 68
| 1 | 6 | 0.84 | 0.8 | 0.8 | 0.995 | 0.83 | 0.86 | 0.95 | 0.81 | 0.7 | 0.8178 | 0.64 | 0.79 | 0.872 | 7 | 0.8 | 1.9 | < 1 | 0.84 | 0.84
| Mean |  | 0.84 | 0.83167 | 0.85 | 0.9625 | 0.925 | 0.90667 | 0.94167 | 0.94167 | 0.76667 | 0.82133 | 0.63333 | 0.89 | 0.85017 | 6.1667 | 0.75 | 1.45 |  | 1.035 | 0.915
| Median |  | 0.83 | 0.82 | 0.85 | 0.946 | 0.93 | 0.86 | 0.95 | 0.9 | 0.8 | 0.8289 | 0.66 | 0.845 | 0.83 | 6.5 | 0.75 | 1.45 |  | 0.84 | 0.885
| Std Dev. |  | 0.037417 | 0.052313 | 0.10488 | 0.12314 | 0.090277 | 0.12275 | 0.037639 | 0.19374 | 0.05164 | 0.036905 | 0.12061 | 0.12946 | 0.064568 | 0.98319 | 0.054772 | 0.30166 |  | 0.48764 | 0.083126
| Rel.Std.Dev. |  | 4.454% | 6.29% | 12.34% | 12.79% | 9.76% | 13.54% | 3.997% | 20.57% | 6.736% | 4.493% | 19.04% | 14.55% | 7.595% | 15.94% | 7.303% | 20.8% |  | 47.11% | 9.085%
| PDM3 |  | -1.324% | -2.303% | -0.1493% | 13.07% | 8.661% | 6.507% | 10.62% | 10.62% | -9.939% | -3.518% | -25.6% | 4.55% | -0.1297% | 624.4% | -11.9% | 70.33% |  | 21.58% | 7.486%
| Z-Score (Absolute) |  | 0.07 | 0.43 | 0.07 | 0.97 | 1.05 | 0.17 | 1.27 | 0.3 | 1.03 | 0.31 | 2.77 | 0.07 | 0.24 | 69.68 | 1.24 | 7.92 | NA | 0.12 | 0.92
| Analytical results for Al in OREAS 922 (Certified Value 2.72 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 2
3
| 1 | 1 | 2.62 | 2.5123 | 2.96 | 2.75 | 2.73 | 2.77 | 2.66 | 2.63 | 2.9462 | 2.72 | 2.939 | 2.42 | 2.74 | 2.594 | 2.56 | 2.75 |  |  |  |  | 1
| 1 | 2 | 2.59 | 2.5739 | 2.91 | 2.75 | 2.8 | 2.78 | 2.63 | 2.65 | 2.9778 | 2.7 | 2.847 | 2.44 | 2.79 | 2.652 | 2.57 | 2.75 |  |  |  |  | #N/A
| 1 | 3 | 2.61 | 2.5722 | 2.9 | 2.89 | 2.83 | 2.79 | 2.66 | 2.67 | 2.9874 | 2.67 | 2.909 | 2.45 | 2.7 | 2.608 | 2.59 | 2.73 |  |  |  |  | 16
| 1 | 4 | 2.54 | 2.5235 | 2.9 | 2.86 | 2.92 | 2.74 | 2.65 | 2.73 | 2.8359 | 2.66 | 2.863 | 2.52 | 2.78 | 2.671 | 2.54 | 2.74 |  |  |  |  | 2.715
| 1 | 5 | 2.46 | 2.5521 | 3.09 | 2.86 | 2.79 | 2.74 | 2.65 | 2.68 | 2.9896 | 2.61 | 2.795 | 2.53 | 2.79 | 2.639 | 2.56 | 2.73 |  |  |  |  | 69
| 1 | 6 | 2.42 | 2.4422 | 3.02 | 2.75 | 2.77 | 2.64 | 2.62 | 2.7 | 2.9807 | 2.63 | 2.962 | 2.52 | 2.77 | 2.643 | 2.56 | 2.73
| Mean |  | 2.54 | 2.5294 | 2.9633 | 2.81 | 2.8067 | 2.7433 | 2.645 | 2.6767 | 2.9529 | 2.665 | 2.8858 | 2.48 | 2.7617 | 2.6345 | 2.5633 | 2.7383
| Median |  | 2.565 | 2.5378 | 2.935 | 2.805 | 2.795 | 2.755 | 2.65 | 2.675 | 2.9792 | 2.665 | 2.886 | 2.485 | 2.775 | 2.641 | 2.56 | 2.735
| Std Dev. |  | 0.083187 | 0.049515 | 0.077632 | 0.066633 | 0.064704 | 0.05465 | 0.016432 | 0.03559 | 0.059442 | 0.041352 | 0.062362 | 0.04858 | 0.035449 | 0.028543 | 0.01633 | 0.0098319
| Rel.Std.Dev. |  | 3.275% | 1.958% | 2.62% | 2.371% | 2.305% | 1.992% | 0.6212% | 1.33% | 2.013% | 1.552% | 2.161% | 1.959% | 1.284% | 1.083% | 0.6371% | 0.359%
| PDM3 |  | -6.446% | -6.837% | 9.147% | 3.499% | 3.376% | 1.044% | -2.578% | -1.412% | 8.763% | -1.842% | 6.292% | -8.656% | 1.719% | -2.965% | -5.586% | 0.8593%
| Z-Score (Absolute) |  | 1.12 | 1.19 | 1.71 | 0.69 | 0.66 | 0.38 | 0.42 | 0.21 | 1.8 | 0.28 | 1.19 | 1.52 | 0.36 | 0.49 | 0.96 | 0.21
| Analytical results for As in OREAS 922 (Certified Value 6.12 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  | 2
3
| 1 | 1 | 12 | 4 | 7 | 6.3 | 6.3 | 6.8 | 6 | 6 | 7 | 6.5079 | 5 | 8 | 5.828 | 7 | 5.7 | 6 | < 10 | 6 | 6 |  | 1
| 1 | 2 | 12 | 5 | 5 | 6.1 | 6.6 | 6.5 | 6 | 6 | 8 | 6.4568 | 6 | 8 | 6.082 | 4 | 6.1 | 6 | < 10 | 6 | 6.6 |  | #N/A
| 1 | 3 | 12 | 4 | 7 | 6.4 | 6.8 | 6.9 | 6.2 | 6 | 9 | 6.4136 | 6 | 7 | 5.423 | 6 | 5.7 | 6 | < 10 | 6 | 6.5 |  | 16
| 1 | 4 | 12 | 5 | 6 | 6.1 | 7 | 6.3 | 6.2 | 6 | 9 | 6.3196 | 5 | 8 | 5.772 | 7 | 5.9 | 5 | < 10 | 6 | 5.8 |  | 6.1175
| 1 | 5 | 12 | 5 | 6 | 6.4 | 6.6 | 6 | 6.2 | 6 | 9 | 6.7287 | 6 | 8 | 5.47 | 5 | 5.5 | 6 | < 10 | 6 | 6 |  | 70
| 1 | 6 | 12 | 4 | 4 | 6.3 | 6.3 | 6.2 | 6.4 | 5 | 8 | 6.614 | 5 | 8 | 5.754 | 7 | 5.5 | 6 | < 10 | 6 | 5.7
| Mean |  | 12 | 4.5 | 5.8333 | 6.2667 | 6.6 | 6.45 | 6.1667 | 5.8333 | 8.3333 | 6.5068 | 5.5 | 7.8333 | 5.7215 | 6 | 5.7333 | 5.8333 |  | 6 | 6.1
| Median |  | 12 | 4.5 | 6 | 6.3 | 6.6 | 6.4 | 6.2 | 6 | 8.5 | 6.4824 | 5.5 | 8 | 5.763 | 6.5 | 5.7 | 6 |  | 6 | 6
| Std Dev. |  | 0 | 0.54772 | 1.169 | 0.13663 | 0.27568 | 0.35071 | 0.15055 | 0.40825 | 0.8165 | 0.14627 | 0.54772 | 0.40825 | 0.24383 | 1.2649 | 0.23381 | 0.40825 |  | 0 | 0.36878
| Rel.Std.Dev. |  | 0% | 12.17% | 20.04% | 2.18% | 4.177% | 5.437% | 2.441% | 6.999% | 9.798% | 2.248% | 9.959% | 5.212% | 4.262% | 21.08% | 4.078% | 6.999% |  | 0% | 6.046%
| PDM3 |  | 96.16% | -26.44% | -4.645% | 2.438% | 7.887% | 5.435% | 0.8038% | -4.645% | 36.22% | 6.363% | -10.09% | 28.05% | -6.473% | -1.921% | -6.28% | -4.645% |  | -1.921% | -0.286%
| Z-Score (Absolute) |  | 13.29 | 3.85 | 0.04 | 0.19 | 0.95 | 0.61 | 0.04 | 0.42 | 4.91 | 0.74 | 1.56 | 4.15 | 1.06 | 0.5 | 1.03 | 0.42 | NA | 0.42 | 0.19
| Analytical results for Au in OREAS 922 (Certified Value < 5 ppb) | [21 empty] | Certified
Au | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 06; 06: 07; 07: 09; 09: 14; 10: 18; 11: 22 |  | ppb
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 09: AR*MS; 10: AR*MS; 11: AR*MS |  | 0
0
| 1 | 1 | 02: < 1; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 2; 09: < 1; 10: < 5; 11: < 200 |  | 1
| 1 | 2 | 02: 5; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 2; 09: < 1; 10: < 5; 11: < 200 |  | #N/A
| 1 | 3 | 02: 2; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 2; 09: < 1; 10: < 5; 11: < 200 |  | 16
| 1 | 4 | 02: 11; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 3; 09: < 1; 10: < 5; 11: < 200 |  | < 5
| 1 | 5 | 02: 5; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 2; 09: < 1; 10: < 5; 11: < 200 |  | 71
| 1 | 6 | 02: < 1; 03: < 5; 04: < 5; 05: < 200; 06: < 200; 07: 2; 09: < 1; 10: < 5; 11: < 200
| Mean |  | 02: 5.75; 07: 2.1667
| Median |  | 02: 5; 07: 2
| Std Dev. |  | 02: 3.7749; 07: 0.40825
| Rel.Std.Dev. |  | 02: 65.65%; 07: 18.84%
PDM3
| Z-Score (Absolute) |  | 02: 0.67; 03: NA; 04: NA; 05: NA; 06: NA; 07: 0.67; 09: NA; 10: NA; 11: NA
| Analytical results for B in OREAS 922 (Indicative Value 35.6 ppm) | [21 empty] | Indicative
B | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 04 | 05 | 06 | 07 | 11 | 15 | 16 | 17 | 18 | 21 | 22 | [9 empty] | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 1
1
| 1 | 1 | < 10 | < 10 | < 10 | < 10 | < 20 | 32 | 43.528 | 47 | < 10 | 20 | < 10 | [9 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 10 | < 20 | 31 | 42.041 | 49 | < 10 | 20 | < 10 | [9 empty] | 1
| 1 | 3 | < 10 | < 10 | < 10 | < 10 | < 20 | 30 | 42.568 | 49 | < 10 | 20 | < 10 | [9 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 10 | < 20 | 31 | 42.838 | 53 | < 10 | 30 | < 10 | [9 empty] | 35.6256
| 1 | 5 | < 10 | < 10 | < 10 | < 10 | < 20 | 28 | 41.934 | 51 | < 10 | 20 | < 10 | [9 empty] | 9
| 1 | 6 | < 10 | < 10 | < 10 | < 10 | < 20 | 28 | 44.106 | 49 | < 10 | 20 | < 10
| Mean |  | 15: 30; 16: 42.8358; 17: 49.6667; 21: 21.6667
| Median |  | 15: 30.5; 16: 42.703; 17: 49; 21: 20
| Std Dev. |  | 15: 1.6733; 16: 0.84961; 17: 2.0656; 21: 4.0825
| Rel.Std.Dev. |  | 15: 5.578%; 16: 1.983%; 17: 4.159%; 21: 18.84%
| PDM3 |  | 15: -15.79%; 16: 20.24%; 17: 39.41%; 21: -39.18%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.44 | 0.44 | 0.91 | NA | 1.13 | NA
| Analytical results for Ba in OREAS 922 (Certified Value 70 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 0
0
| 1 | 1 | 76 | 67 | 88 | 92 | 70 | 70 | 66 | 86 | 85.3242 | 42 | 129 | 66.156 | 69 | 64 | 65 | 63 | 60 |  |  |  | 1
| 1 | 2 | 63 | 71 | 82 | 90 | 70 | 70 | 63 | 87 | 86.9783 | 41 | 123 | 65.144 | 68 | 64 | 64 | 63 | 60 |  |  |  | #N/A
| 1 | 3 | 64 | 68 | 88 | 91 | 70 | 70 | 64 | 88 | 84.528 | 46 | 123 | 64.755 | 70 | 62 | 65 | 64 | 60 |  |  |  | 16
| 1 | 4 | 62 | 66 | 90 | 91 | 70 | 70 | 65 | 91 | 83.3303 | 44 | 123 | 65.608 | 70 | 65 | 64 | 65 | 60 |  |  |  | 69.9699
| 1 | 5 | 59 | 66 | 93 | 91 | 70 | 70 | 66 | 89 | 86.0738 | 45 | 118 | 65.084 | 69 | 65 | 65 | 62 | 60 |  |  |  | 72
| 1 | 6 | 60 | 62 | 96 | 92 | 60 | 60 | 69 | 89 | 89.1303 | 45 | 117 | 65.397 | 70 | 65 | 65 | 64 | 60
| Mean |  | 64 | 66.6667 | 89.5 | 91.1667 | 68.3333 | 68.3333 | 65.5 | 88.3333 | 85.8941 | 43.8333 | 122.1667 | 65.3573 | 69.3333 | 64.1667 | 64.6667 | 63.5 | 60
| Median |  | 62.5 | 66.5 | 89 | 91 | 70 | 70 | 65.5 | 88.5 | 85.699 | 44.5 | 123 | 65.2705 | 69.5 | 64.5 | 65 | 63.5 | 60
| Std Dev. |  | 6.1644 | 2.9439 | 4.8062 | 0.75277 | 4.0825 | 4.0825 | 2.0736 | 1.7512 | 2.0221 | 1.9408 | 4.3089 | 0.48689 | 0.8165 | 1.169 | 0.5164 | 1.0488 | 0
| Rel.Std.Dev. |  | 9.632% | 4.416% | 5.37% | 0.8257% | 5.974% | 5.974% | 3.166% | 1.982% | 2.354% | 4.428% | 3.527% | 0.745% | 1.178% | 1.822% | 0.7986% | 1.652% | 0%
| PDM3 |  | -8.532% | -4.721% | 27.91% | 30.29% | -2.339% | -2.339% | -6.388% | 26.24% | 22.76% | -37.35% | 74.6% | -6.592% | -0.9097% | -8.294% | -7.579% | -9.247% | -14.25%
| Z-Score (Absolute) |  | 1.02 | 0 | 4.62 | 4.96 | 0.67 | 0.67 | 0.24 | 4.38 | 3.89 | 4.62 | 11.23 | 0.26 | 0.54 | 0.51 | 0.4 | 0.64 | 1.35
| Analytical results for Be in OREAS 922 (Certified Value 0.65 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
2
| 1 | 1 | 0.7 | 0.65 | 0.7 | 0.7 | 0.67 | 0.58 | 0.81633 | 0.53 | < 5 | 0.705 | < 2 | 0.6 | 0.7 | 0.6 | 0.68 | [5 empty] | 1
| 1 | 2 | 0.7 | 0.65 | 0.7 | 0.7 | 0.7 | 0.54 | 0.76134 | 0.39 | < 5 | 0.732 | 2 | 0.6 | 0.6 | 0.6 | 0.65 | [5 empty] | #N/A
| 1 | 3 | 0.7 | 0.66 | 0.7 | 0.7 | 0.73 | 0.54 | 0.76923 | 0.53 | < 5 | 0.72 | 2 | 0.6 | 0.6 | 0.6 | 0.65 | [5 empty] | 16
| 1 | 4 | 0.7 | 0.65 | 0.7 | 0.7 | 0.7 | 0.54 | 0.81492 | 0.47 | < 5 | 0.728 | < 2 | 0.6 | 0.6 | 0.6 | 0.69 | [5 empty] | 0.65166
| 1 | 5 | 0.7 | 0.63 | 0.7 | 0.7 | 0.64 | 0.54 | 0.80915 | 0.5 | < 5 | 0.719 | 2 | 0.6 | 0.6 | 0.6 | 0.67 | [5 empty] | 73
| 1 | 6 | 0.7 | 0.64 | 0.9 | 0.7 | 0.62 | 0.53 | 0.79128 | 0.58 | < 5 | 0.755 | 2 | 0.6 | 0.6 | 0.6 | 0.58
| Mean |  | 0.7 | 0.64667 | 0.73333 | 0.7 | 0.67667 | 0.545 | 0.79371 | 0.5 |  | 0.7265 | 2 | 0.6 | 0.61667 | 0.6 | 0.65333
| Median |  | 0.7 | 0.65 | 0.7 | 0.7 | 0.685 | 0.54 | 0.80021 | 0.515 |  | 0.724 | 2 | 0.6 | 0.6 | 0.6 | 0.66
| Std Dev. |  | 1.2162e-16 | 0.010328 | 0.08165 | 1.2162e-16 | 0.041312 | 0.017607 | 0.023889 | 0.065115 |  | 0.016766 | 0 | 0 | 0.040825 | 0 | 0.039328
| Rel.Std.Dev. |  | 1.737e-14% | 1.597% | 11.13% | 1.737e-14% | 6.105% | 3.231% | 3.01% | 13.02% |  | 2.308% | 0% | 0% | 6.62% | 0% | 6.02%
| PDM3 |  | 7.418% | -0.7658% | 12.53% | 7.418% | 3.838% | -16.37% | 21.8% | -23.27% |  | 11.49% | 206.9% | -7.927% | -5.369% | -7.927% | 0.2572%
| Z-Score (Absolute) |  | 0.29 | 0.27 | 0.29 | 0.29 | 0.05 | 1.43 | 1.29 | 1.6 | NA | 0.58 | 14.15 | 0.77 | 0.77 | 0.77 | 0.05
| Analytical results for Bi in OREAS 922 (Certified Value 10.3 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
1
| 1 | 1 | 9.16 | 9.63 | 11.4 | 13.8 | 10.2 | 8.86 | 10.8 | 10.88 | 6 | 10.3477 | 8.54 | 10 | 10.016 | 20 | 8.62 | 12.3 | 8.7 | 10.55 |  |  | 1
| 1 | 2 | 9.32 | 10.25 | 10.2 | 12.3 | 12.8 | 8.61 | 10.2 | 10.11 | 11 | 10.5998 | 8.37 | 10 | 10.474 | 21 | 10 | 10.2 | 11.5 | 10.1 |  |  | #N/A
| 1 | 3 | 9.42 | 11.09 | 12.6 | 11.7 | 11.4 | 8.76 | 10.3 | 9.77 | 9 | 10.08 | 9.24 | 10 | 13.032 | 11 | 9.48 | 11.5 | 11.3 | 12.05 |  |  | 16
| 1 | 4 | 9.15 | 10.58 | 9.89 | 12.5 | 10.05 | 10.6 | 11.3 | 10.43 | 9 | 10.1156 | 9.14 | 8 | 10.042 | 18 | 9.08 | 12.1 | 8.47 | 10.5 |  |  | 10.3066
| 1 | 5 | 9.26 | 9.69 | 12.3 | 11.1 | 11.95 | 10.55 | 11.2 | 10.61 | 7 | 10.2329 | 9.05 | 7 | 9.337 | 18 | 9.54 | 11.3 | 9.96 | 11.3 |  |  | 74
| 1 | 6 | 9.25 | 10.64 | 10.7 | 10.9 | 10.1 | 9.38 | 10.7 | 12.34 | 10 | 10.2721 | 8.84 | 6 | 13.689 | 27 | 10.4 | 12 | 11.4 | 11.55
| Mean |  | 9.26 | 10.3133 | 11.1817 | 12.05 | 11.0833 | 9.46 | 10.75 | 10.69 | 8.6667 | 10.2747 | 8.8633 | 8.5 | 11.0983 | 19.1667 | 9.52 | 11.5667 | 10.2217 | 11.0083
| Median |  | 9.255 | 10.415 | 11.05 | 12 | 10.8 | 9.12 | 10.75 | 10.52 | 9 | 10.2525 | 8.945 | 9 | 10.258 | 19 | 9.51 | 11.75 | 10.63 | 10.925
| Std Dev. |  | 0.10139 | 0.57281 | 1.1112 | 1.0654 | 1.1501 | 0.90184 | 0.45056 | 0.89615 | 1.8619 | 0.18761 | 0.34691 | 1.7607 | 1.8017 | 5.1929 | 0.6341 | 0.76855 | 1.3884 | 0.74257
| Rel.Std.Dev. |  | 1.095% | 5.554% | 9.938% | 8.841% | 10.38% | 9.533% | 4.191% | 8.383% | 21.48% | 1.826% | 3.914% | 20.71% | 16.23% | 27.09% | 6.661% | 6.645% | 13.58% | 6.746%
| PDM3 |  | -10.15% | 0.06581% | 8.491% | 16.92% | 7.537% | -8.214% | 4.303% | 3.72% | -15.91% | -0.3091% | -14% | -17.53% | 7.682% | 85.97% | -7.632% | 12.23% | -0.8236% | 6.809%
| Z-Score (Absolute) |  | 0.87 | 0.02 | 0.69 | 1.39 | 0.61 | 0.71 | 0.34 | 0.02 | 0.92 | 0.05 | 1.2 | 1.08 | 0.62 | 7.17 | 0.66 | 1 | 0.09 | 0.55
| Analytical results for Ca in OREAS 922 (Certified Value 0.324 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 3
3
| 1 | 1 | 0.33 | 0.33 | 0.45 | 0.4 | 0.33 | 0.32 | 0.32 | 0.28 | 0.40136 | 0.35 | 0.349 | 0.35 | 0.33 | 0.311 | 0.31 | 0.32 |  |  |  |  | 1
| 1 | 2 | 0.32 | 0.34 | 0.45 | 0.4 | 0.35 | 0.31 | 0.32 | 0.28 | 0.40209 | 0.34 | 0.334 | 0.34 | 0.33 | 0.32 | 0.31 | 0.32 |  |  |  |  | #N/A
| 1 | 3 | 0.33 | 0.34 | 0.43 | 0.42 | 0.34 | 0.32 | 0.32 | 0.28 | 0.41217 | 0.34 | 0.345 | 0.35 | 0.32 | 0.317 | 0.31 | 0.32 |  |  |  |  | 16
| 1 | 4 | 0.32 | 0.33 | 0.42 | 0.42 | 0.34 | 0.31 | 0.32 | 0.29 | 0.387 | 0.34 | 0.345 | 0.36 | 0.33 | 0.332 | 0.31 | 0.31 |  |  |  |  | 0.3238
| 1 | 5 | 0.31 | 0.34 | 0.44 | 0.42 | 0.33 | 0.31 | 0.31 | 0.29 | 0.40437 | 0.33 | 0.335 | 0.35 | 0.33 | 0.317 | 0.31 | 0.32 |  |  |  |  | 75
| 1 | 6 | 0.3 | 0.32 | 0.43 | 0.4 | 0.32 | 0.3 | 0.31 | 0.29 | 0.3924 | 0.33 | 0.351 | 0.36 | 0.33 | 0.316 | 0.31 | 0.32
| Mean |  | 0.31833 | 0.33333 | 0.43667 | 0.41 | 0.335 | 0.31167 | 0.31667 | 0.285 | 0.3999 | 0.33833 | 0.34317 | 0.35167 | 0.32833 | 0.31883 | 0.31 | 0.31833
| Median |  | 0.32 | 0.335 | 0.435 | 0.41 | 0.335 | 0.31 | 0.32 | 0.285 | 0.40173 | 0.34 | 0.345 | 0.35 | 0.33 | 0.317 | 0.31 | 0.32
| Std Dev. |  | 0.01169 | 0.008165 | 0.012111 | 0.010954 | 0.010488 | 0.0075277 | 0.005164 | 0.0054772 | 0.0089472 | 0.0075277 | 0.007111 | 0.0075277 | 0.0040825 | 0.0070828 | 0 | 0.0040825
| Rel.Std.Dev. |  | 3.672% | 2.449% | 2.773% | 2.672% | 3.131% | 2.415% | 1.631% | 1.922% | 2.237% | 2.225% | 2.072% | 2.141% | 1.243% | 2.221% | 0% | 1.282%
| PDM3 |  | -1.688% | 2.945% | 34.86% | 26.62% | 3.46% | -3.746% | -2.202% | -11.98% | 23.5% | 4.489% | 5.982% | 8.607% | 1.401% | -1.533% | -4.261% | -1.688%
| Z-Score (Absolute) |  | 0.59 | 0.07 | 4.65 | 3.47 | 0.15 | 0.89 | 0.66 | 2.07 | 3.02 | 0.3 | 0.51 | 0.89 | 0.07 | 0.68 | 0.96 | 0.52
| Analytical results for Cd in OREAS 922 (Certified Value 0.28 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  |  |  | 2
3
| 1 | 1 | 0.27 | 0.28 | 0.27 | 0.29 | 0.28 | 0.29 | 0.3 | 0.8 | 0.26997 | 0.22 | < 0.5 | 0.306 | 0.26 | 0.3 | < 1 | 0.24 | 0.26 |  |  |  | 1
| 1 | 2 | 0.28 | 0.28 | 0.26 | 0.27 | 0.29 | 0.29 | 0.3 | 0.9 | 0.28444 | 0.25 | < 0.5 | 0.326 | 0.24 | 0.3 | < 1 | 0.26 | 0.25 |  |  |  | #N/A
| 1 | 3 | 0.28 | 0.3 | 0.26 | 0.24 | 0.27 | 0.3 | 0.3 | 0.7 | 0.28409 | 0.25 | < 0.5 | 0.311 | 0.25 | 0.3 | < 1 | 0.27 | 0.29 |  |  |  | 16
| 1 | 4 | 0.27 | 0.28 | 0.26 | 0.27 | 0.29 | 0.29 | 0.3 | 0.9 | 0.26994 | 0.23 | < 0.5 | 0.312 | 0.28 | 0.3 | < 1 | 0.28 | 0.27 |  |  |  | 0.27605
| 1 | 5 | 0.27 | 0.26 | 0.25 | 0.28 | 0.28 | 0.27 | 0.3 | 0.9 | 0.28375 | 0.24 | < 0.5 | 0.301 | 0.25 | 0.3 | < 1 | 0.26 | 0.27 |  |  |  | 76
| 1 | 6 | 0.27 | 0.27 | 0.28 | 0.29 | 0.27 | 0.27 | 0.3 | 0.7 | 0.2933 | 0.21 | < 0.5 | 0.307 | 0.24 | 0.3 | < 1 | 0.27 | 0.28
| Mean |  | 0.27333 | 0.27833 | 0.26333 | 0.27333 | 0.28 | 0.285 | 0.3 | 0.81667 | 0.28092 | 0.23333 |  | 0.3105 | 0.25333 | 0.3 |  | 0.26333 | 0.27
| Median |  | 0.27 | 0.28 | 0.26 | 0.275 | 0.28 | 0.29 | 0.3 | 0.85 | 0.28392 | 0.235 |  | 0.309 | 0.25 | 0.3 |  | 0.265 | 0.27
| Std Dev. |  | 0.005164 | 0.013292 | 0.010328 | 0.018619 | 0.0089443 | 0.012247 | 0 | 0.098319 | 0.0092122 | 0.01633 |  | 0.0085499 | 0.015055 | 0 |  | 0.013663 | 0.014142
| Rel.Std.Dev. |  | 1.889% | 4.775% | 3.922% | 6.812% | 3.194% | 4.297% | 0% | 12.04% | 3.279% | 6.999% |  | 2.754% | 5.943% | 0% |  | 5.188% | 5.238%
| PDM3 |  | -0.9854% | 0.8258% | -4.608% | -0.9854% | 1.43% | 3.241% | 8.675% | 195.8% | 1.762% | -15.48% |  | 12.48% | -8.23% | 8.675% |  | -4.608% | -2.193%
| Z-Score (Absolute) |  | 0.22 | 0 | 0.67 | 0.22 | 0.07 | 0.3 | 0.97 | 24.2 | 0.12 | 2.02 | NA | 1.45 | 1.12 | 0.97 | NA | 0.67 | 0.37
| Analytical results for Ce in OREAS 922 (Certified Value 63 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 0
0
| 1 | 1 | 56.3 | 63.19 | 77.6 | 74.3 | 65 | 65.1 | 55.9 | 80.1024 | 47 | 66.618 | 57.8 | 65.5 | 61.8 | 61.9 | [6 empty] | 1
| 1 | 2 | 56.2 | 65.06 | 77.7 | 72.8 | 67.6 | 64 | 54.7 | 82.4868 | 45 | 66.773 | 58.6 | 67.4 | 61.9 | 62.2 | [6 empty] | #N/A
| 1 | 3 | 55.9 | 65.63 | 78 | 72.9 | 66.3 | 66 | 55.6 | 80.062 | 53 | 68.638 | 58.6 | 65.3 | 62.5 | 62.6 | [6 empty] | 16
| 1 | 4 | 56 | 64.6 | 77.1 | 71.7 | 66.2 | 63.1 | 57.1 | 80.9575 | 51 | 67.767 | 58.5 | 66.1 | 62.6 | 60.7 | [6 empty] | 63.2671
| 1 | 5 | 55.6 | 66.34 | 80.2 | 73.7 | 64.9 | 60.4 | 56.5 | 81.4116 | 51 | 67.721 | 59.1 | 66.1 | 62.5 | 62.7 | [6 empty] | 77
| 1 | 6 | 55.2 | 62.34 | 82.2 | 75.5 | 62.8 | 60 | 56.7 | 82.8758 | 49 | 68.035 | 58.2 | 67.3 | 61.8 | 61.7
| Mean |  | 55.8667 | 64.5267 | 78.8 | 73.4833 | 65.4667 | 63.1 | 56.0833 | 81.316 | 49.3333 | 67.592 | 58.4667 | 66.2833 | 62.1833 | 61.9667
| Median |  | 55.95 | 64.83 | 77.85 | 73.3 | 65.6 | 63.55 | 56.2 | 81.1846 | 50 | 67.744 | 58.55 | 66.1 | 62.2 | 62.05
| Std Dev. |  | 0.40825 | 1.5075 | 1.9849 | 1.3243 | 1.6391 | 2.4544 | 0.86814 | 1.1821 | 2.9439 | 0.76903 | 0.43665 | 0.88638 | 0.38687 | 0.73121
| Rel.Std.Dev. |  | 0.7308% | 2.336% | 2.519% | 1.802% | 2.504% | 3.89% | 1.548% | 1.454% | 5.967% | 1.138% | 0.7468% | 1.337% | 0.6221% | 1.18%
| PDM3 |  | -11.7% | 1.991% | 24.55% | 16.15% | 3.477% | -0.2641% | -11.35% | 28.53% | -22.02% | 6.836% | -7.588% | 4.767% | -1.713% | -2.055%
| Z-Score (Absolute) |  | 1.17 | 0.11 | 2.11 | 1.43 | 0.24 | 0.11 | 1.14 | 2.59 | 2.14 | 0.56 | 0.79 | 0.37 | 0.24 | 0.27
| Analytical results for Co in OREAS 922 (Certified Value 19.4 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | 1
2
| 1 | 1 | 17.6 | 19.3 | 20.6 | 18 | 18.8 | 19.1 | 21.2671 | 21 | 19.3 | 17 | 19.6409 | 17 | 16 | 19.53 | 22 | 19 | 19.1 | 18.2 | 19.8 | 18.9 | 1
| 1 | 2 | 17.6 | 19.9 | 20.2 | 18 | 20 | 18.6 | 22.0519 | 20 | 18.6 | 18 | 19.3131 | 17 | 16 | 19.87 | 21 | 19 | 19 | 18.3 | 20.3 | 19.4 | #N/A
| 1 | 3 | 17.7 | 19.8 | 20.8 | 19 | 19.6 | 18.9 | 20.8084 | 20 | 19.4 | 17 | 19.7708 | 19 | 16 | 19.87 | 20 | 18 | 19.2 | 18.9 | 20.2 | 19.5 | 16
| 1 | 4 | 17.5 | 19.3 | 20.1 | 19 | 20 | 18.3 | 22.5704 | 20 | 23.9 | 18 | 19.3491 | 19 | 16 | 20.16 | 21 | 19 | 19 | 19.2 | 20.1 | 18.7 | 19.3507
| 1 | 5 | 17.5 | 18.6 | 21 | 19 | 19.4 | 17.2 | 21.0127 | 20 | 19.2 | 18 | 19.7637 | 19 | 16 | 19.78 | 21 | 19 | 19.7 | 18.6 | 20 | 19.5 | 78
| 1 | 6 | 17.5 | 19 | 21.4 | 19 | 18.8 | 18.3 | 21.6109 | 21 | 19.3 | 18 | 19.7311 | 19 | 16 | 19.72 | 21 | 19 | 19.8 | 18.4 | 20.3 | 18.7
| Mean |  | 17.5667 | 19.3167 | 20.6833 | 18.6667 | 19.4333 | 18.4 | 21.5536 | 20.3333 | 19.95 | 17.6667 | 19.5948 | 18.3333 | 16 | 19.8217 | 21 | 18.8333 | 19.3 | 18.6 | 20.1167 | 19.1167
| Median |  | 17.55 | 19.3 | 20.7 | 19 | 19.5 | 18.45 | 21.439 | 20 | 19.3 | 18 | 19.686 | 19 | 16 | 19.825 | 21 | 19 | 19.15 | 18.5 | 20.15 | 19.15
| Std Dev. |  | 0.08165 | 0.48751 | 0.4916 | 0.5164 | 0.54283 | 0.66933 | 0.66566 | 0.5164 | 1.9563 | 0.5164 | 0.20974 | 1.0328 | 0 | 0.20798 | 0.63246 | 0.40825 | 0.35777 | 0.38471 | 0.19408 | 0.392
| Rel.Std.Dev. |  | 0.4648% | 2.524% | 2.377% | 2.766% | 2.793% | 3.638% | 3.088% | 2.54% | 9.806% | 2.923% | 1.07% | 5.633% | 0% | 1.049% | 3.012% | 2.168% | 1.854% | 2.068% | 0.9648% | 2.051%
| PDM3 |  | -9.219% | -0.1759% | 6.887% | -3.535% | 0.427% | -4.913% | 11.38% | 5.078% | 3.097% | -8.703% | 1.261% | -5.258% | -17.32% | 2.434% | 8.523% | -2.674% | -0.262% | -3.879% | 3.958% | -1.209%
| Z-Score (Absolute) |  | 1.54 | 0.08 | 1.34 | 0.52 | 0.19 | 0.77 | 2.15 | 1.02 | 0.06 | 1.44 | 0.34 | 0.83 | 2.98 | 0.55 | 1.63 | 0.21 | 0.06 | 0.58 | 0.82 | 0.1
| Analytical results for Cr in OREAS 922 (Certified Value 40.7 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  | 1
2
| 1 | 1 | 40 | 43 | 46.5 | 44 | 42 | 42 | < 50 | 39 | 46.7359 | 39 | 54 | 42.038 | 39 | 38.4 | 42 | 39.9 | 40 | 41 |  |  | 1
| 1 | 2 | 40 | 41 | 47.3 | 45 | 43 | 41 | < 50 | 39 | 43.3018 | 38 | 51 | 41.554 | 40 | 40.1 | 42 | 40.1 | 40 | 41 |  |  | #N/A
| 1 | 3 | 41 | 43 | 46.6 | 46 | 42 | 43 | < 50 | 39 | 42.8981 | 41 | 52 | 39.798 | 39 | 39.4 | 42 | 40.4 | 41 | 42 |  |  | 16
| 1 | 4 | 38 | 41 | 46 | 47 | 44 | 41 | < 50 | 41 | 42.2729 | 41 | 52 | 42.013 | 41 | 39.4 | 43 | 39.9 | 42 | 40 |  |  | 40.7392
| 1 | 5 | 38 | 40 | 48.6 | 47 | 42 | 41 | < 50 | 40 | 43.714 | 40 | 51 | 40.644 | 40 | 39.3 | 41 | 40 | 39 | 40 |  |  | 79
| 1 | 6 | 37 | 41 | 48.9 | 46 | 41 | 40 | < 50 | 40 | 44.195 | 39 | 51 | 41.085 | 41 | 38.3 | 41 | 41.1 | 40 | 41
| Mean |  | 39 | 41.5 | 47.3167 | 45.8333 | 42.3333 | 41.3333 |  | 39.6667 | 43.8529 | 39.6667 | 51.8333 | 41.1887 | 40 | 39.15 | 41.8333 | 40.2333 | 40.3333 | 40.8333
| Median |  | 39 | 41 | 46.95 | 46 | 42 | 41 |  | 39.5 | 43.5079 | 39.5 | 51.5 | 41.3195 | 40 | 39.35 | 42 | 40.05 | 40 | 41
| Std Dev. |  | 1.5492 | 1.2247 | 1.189 | 1.169 | 1.0328 | 1.0328 |  | 0.8165 | 1.5595 | 1.2111 | 1.169 | 0.86899 | 0.89443 | 0.68337 | 0.75277 | 0.46332 | 1.0328 | 0.75277
| Rel.Std.Dev. |  | 3.972% | 2.951% | 2.513% | 2.551% | 2.44% | 2.499% |  | 2.058% | 3.556% | 3.053% | 2.255% | 2.11% | 2.236% | 1.746% | 1.799% | 1.152% | 2.561% | 1.844%
| PDM3 |  | -4.269% | 1.868% | 16.15% | 12.5% | 3.913% | 1.458% |  | -2.633% | 7.643% | -2.633% | 27.23% | 1.103% | -1.814% | -3.901% | 2.686% | -1.242% | -0.9962% | 0.2311%
| Z-Score (Absolute) |  | 1.24 | 0.18 | 3.48 | 2.63 | 0.65 | 0.08 | NA | 0.86 | 1.18 | 0.86 | 5.79 | 0 | 0.67 | 1.16 | 0.37 | 0.54 | 0.49 | 0.2
| Analytical results for Cs in OREAS 922 (Certified Value 1.76 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
2
| 1 | 1 | 1.54 | 2.07 | 1.56 | 2.16 | 1.53 | 1.72 | 1.7 | 2.4935 | 1.36 | 1.83 | 0.98 | 1.9 | 1.85 | 1.81 | [6 empty] | 1
| 1 | 2 | 1.57 | 2.17 | 1.5 | 2.1 | 1.62 | 1.67 | 1.7 | 2.4971 | 1.3 | 1.874 | 1.03 | 2 | 1.85 | 1.84 | [6 empty] | #N/A
| 1 | 3 | 1.55 | 2.09 | 1.55 | 2.12 | 1.6 | 1.72 | 1.7 | 2.4939 | 1.35 | 1.882 | 1.04 | 2 | 1.86 | 1.86 | [6 empty] | 16
| 1 | 4 | 1.55 | 2.04 | 1.68 | 2.16 | 1.61 | 1.67 | 1.7 | 2.3522 | 1.33 | 1.845 | 1.04 | 1.9 | 1.9 | 1.65 | [6 empty] | 1.763
| 1 | 5 | 1.56 | 2.15 | 1.68 | 2.13 | 1.53 | 1.57 | 1.7 | 2.5075 | 1.3 | 1.87 | 1.03 | 2 | 1.87 | 1.84 | [6 empty] | 80
| 1 | 6 | 1.54 | 2.01 | 1.64 | 2.11 | 1.5 | 1.59 | 1.75 | 2.5594 | 1.33 | 1.897 | 1.01 | 2 | 1.92 | 1.74
| Mean |  | 1.5517 | 2.0883 | 1.6017 | 2.13 | 1.565 | 1.6567 | 1.7083 | 2.4839 | 1.3283 | 1.8663 | 1.0217 | 1.9667 | 1.875 | 1.79
| Median |  | 1.55 | 2.08 | 1.6 | 2.125 | 1.565 | 1.67 | 1.7 | 2.4955 | 1.33 | 1.872 | 1.03 | 2 | 1.865 | 1.825
| Std Dev. |  | 0.01169 | 0.062102 | 0.075476 | 0.025298 | 0.050892 | 0.06377 | 0.020412 | 0.069245 | 0.024833 | 0.024631 | 0.023166 | 0.05164 | 0.02881 | 0.080498
| Rel.Std.Dev. |  | 0.7534% | 2.974% | 4.712% | 1.188% | 3.252% | 3.849% | 1.195% | 2.788% | 1.869% | 1.32% | 2.267% | 2.626% | 1.537% | 4.497%
| PDM3 |  | -11.99% | 18.45% | -9.151% | 20.82% | -11.23% | -6.031% | -3.101% | 40.89% | -24.65% | 5.861% | -42.05% | 11.55% | 6.353% | 1.531%
| Z-Score (Absolute) |  | 0.71 | 1.09 | 0.54 | 1.23 | 0.67 | 0.36 | 0.18 | 2.51 | 1.46 | 0.35 | 2.46 | 0.68 | 0.38 | 0.18
| Analytical results for Cu in OREAS 922 (Certified Value 2176 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  | 0
0
| 1 | 1 | 2239 | 2071 | 2210 | 2010 | 2240 | 2340 | 2190 | 2293 | 2014 | 2152.088 | 2043 | 2231 | 2187.06 | 2110 | 2230 | 2267 | 2200 | 2170 | 2130 |  | 1
| 1 | 2 | 2159 | 2100 | 2310 | 1950 | 2330 | 2320 | 2130 | 2226 | 2040 | 2137.9597 | 2023 | 2218 | 2090.16 | 2206 | 2250 | 2229 | 2160 | 2190 | 2150 |  | #N/A
| 1 | 3 | 2199 | 2147 | 2140 | 1970 | 2290 | 2330 | 2150 | 2239 | 2023 | 2136.4984 | 2175 | 2230 | 2108.52 | 2203 | 2200 | 2220 | 2190 | 2200 | 2120 |  | 16
| 1 | 4 | 2110 | 2117 | 2090 | 2120 | 2330 | 2290 | 2200 | 2184 | 2079 | 2085.4548 | 2142 | 2266 | 2131.54 | 2209 | 2270 | 2157 | 2190 | 2170 | 2110 |  | 2175.5245
| 1 | 5 | 2090 | 2120 | 2160 | 2210 | 2280 | 2240 | 2210 | 2328 | 2068 | 2172.1671 | 2071 | 2290 | 2061.83 | 2182 | 2290 | 2179 | 2180 | 2140 | 2110 |  | 81
| 1 | 6 | 2081 | 2109 | 2270 | 2200 | 2230 | 2190 | 2170 | 2281 | 2085 | 2182.5364 | 2069 | 2239 | 2198.98 | 2308 | 2240 | 2196 | 2180 | 2190 | 2140
| Mean |  | 2146.3333 | 2110.6667 | 2196.6667 | 2076.6667 | 2283.3333 | 2285 | 2175 | 2258.5 | 2051.5 | 2144.4507 | 2087.1667 | 2245.6667 | 2129.6817 | 2203 | 2246.6667 | 2208 | 2183.3333 | 2176.6667 | 2126.6667
| Median |  | 2134.5 | 2113 | 2185 | 2065 | 2285 | 2305 | 2180 | 2260 | 2054 | 2145.0239 | 2070 | 2235 | 2120.03 | 2204.5 | 2245 | 2208 | 2185 | 2180 | 2125
| Std Dev. |  | 63.692 | 25.0493 | 82.8654 | 115.5278 | 42.7395 | 58.9067 | 30.8221 | 51.9875 | 30.005 | 34.2431 | 58.959 | 27.016 | 54.2432 | 63.498 | 31.4113 | 39.1203 | 13.6626 | 21.6025 | 16.3299
| Rel.Std.Dev. |  | 2.967% | 1.187% | 3.772% | 5.563% | 1.872% | 2.578% | 1.417% | 2.302% | 1.463% | 1.597% | 2.825% | 1.203% | 2.547% | 2.882% | 1.398% | 1.772% | 0.6258% | 0.9925% | 0.7679%
| PDM3 |  | -1.342% | -2.981% | 0.9718% | -4.544% | 4.956% | 5.032% | -0.02411% | 3.814% | -5.701% | -1.428% | -4.061% | 3.224% | -2.107% | 1.263% | 3.27% | 1.493% | 0.3589% | 0.0525% | -2.246%
| Z-Score (Absolute) |  | 0.41 | 0.89 | 0.27 | 1.35 | 1.44 | 1.46 | 0.02 | 1.1 | 1.69 | 0.43 | 1.21 | 0.93 | 0.63 | 0.36 | 0.94 | 0.42 | 0.09 | 0 | 0.67
| Analytical results for Dy in OREAS 922 (Certified Value < 5 ppm) | [21 empty] | Certified
Dy | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  | 2
2
| 1 | 1 | 02: 3.26; 03: 4.3; 04: 4.7; 05: 2.56; 06: 4.0845; 07: 1.69; 09: 2.8; 10: 3 |  | 1
| 1 | 2 | 02: 3.33; 03: 4.3; 04: 4.5; 05: 2.42; 06: 4.24; 07: 1.51; 09: 2.9; 10: 3 |  | #N/A
| 1 | 3 | 02: 3.38; 03: 4.3; 04: 4.5; 05: 2.45; 06: 4.1535; 07: 1.8; 09: 2.9; 10: 3 |  | 16
| 1 | 4 | 02: 3.28; 03: 4.3; 04: 4.5; 05: 2.58; 06: 4.0093; 07: 1.65; 09: 2.9; 10: 3 |  | < 5
| 1 | 5 | 02: 3.41; 03: 4.5; 04: 4.5; 05: 2.62; 06: 4.1528; 07: 1.75; 09: 2.9; 10: 2.9 |  | 82
| 1 | 6 | 02: 2.96; 03: 4.6; 04: 4.5; 05: 2.65; 06: 4.1208; 07: 1.72; 09: 2.9; 10: 3.1
| Mean |  | 02: 3.27; 03: 4.3833; 04: 4.5333; 05: 2.5467; 06: 4.1268; 07: 1.6867; 09: 2.8833; 10: 3
| Median |  | 02: 3.305; 03: 4.3; 04: 4.5; 05: 2.57; 06: 4.1368; 07: 1.705; 09: 2.9; 10: 3
| Std Dev. |  | 02: 0.16223; 03: 0.13292; 04: 0.08165; 05: 0.092448; 06: 0.077228; 07: 0.10053; 09: 0.040825; 10: 0.063246
| Rel.Std.Dev. |  | 02: 4.961%; 03: 3.032%; 04: 1.801%; 05: 3.63%; 06: 1.871%; 07: 5.96%; 09: 1.416%; 10: 2.108%
PDM3
| Z-Score (Absolute) |  | 02: 0.14; 03: 1; 04: 1.17; 05: 0.53; 06: 0.82; 07: 1.26; 09: 0.24; 10: 0.14
| Analytical results for Er in OREAS 922 (Certified Value < 3 ppm) | [21 empty] | Certified
Er | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  | 2
2
| 1 | 1 | 02: 1.7; 03: 2.3; 04: 2.3; 05: 1.18; 06: 2.3633; 07: 0.77; 09: 1.4; 10: 1.5 |  | 1
| 1 | 2 | 02: 1.82; 03: 2.3; 04: 2.3; 05: 1.17; 06: 2.2604; 07: 0.72; 09: 1.4; 10: 1.4 |  | #N/A
| 1 | 3 | 02: 1.79; 03: 2.3; 04: 2.3; 05: 1.19; 06: 2.1928; 07: 0.86; 09: 1.4; 10: 1.4 |  | 16
| 1 | 4 | 02: 1.75; 03: 2.3; 04: 2.3; 05: 1.26; 06: 2.1055; 07: 0.82; 09: 1.4; 10: 1.5 |  | < 3
| 1 | 5 | 02: 1.77; 03: 2.4; 04: 2.3; 05: 1.26; 06: 2.2169; 07: 0.87; 09: 1.4; 10: 1.4 |  | 83
| 1 | 6 | 02: 1.58; 03: 2.4; 04: 2.4; 05: 1.25; 06: 2.3551; 07: 0.83; 09: 1.5; 10: 1.5
| Mean |  | 02: 1.735; 03: 2.3333; 04: 2.3167; 05: 1.2183; 06: 2.249; 07: 0.81167; 09: 1.4167; 10: 1.45
| Median |  | 02: 1.76; 03: 2.3; 04: 2.3; 05: 1.22; 06: 2.2387; 07: 0.825; 09: 1.4; 10: 1.45
| Std Dev. |  | 02: 0.085965; 03: 0.05164; 04: 0.040825; 05: 0.042622; 06: 0.099237; 07: 0.057067; 09: 0.040825; 10: 0.054772
| Rel.Std.Dev. |  | 02: 4.955%; 03: 2.213%; 04: 1.762%; 05: 3.498%; 06: 4.412%; 07: 7.031%; 09: 2.882%; 10: 3.777%
PDM3
| Z-Score (Absolute) |  | 02: 0.21; 03: 0.95; 04: 0.93; 05: 0.51; 06: 0.84; 07: 1.04; 09: 0.27; 10: 0.21
| Analytical results for Eu in OREAS 922 (Certified Value < 1.5 ppm) | [21 empty] | Certified
Eu | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  | 2
2
| 1 | 1 | 02: 0.86; 03: 1.2; 04: 1.3; 05: 0.73; 06: 1.1524; 07: 0.54; 09: 0.8; 10: 0.9 |  | 1
| 1 | 2 | 02: 0.92; 03: 1.1; 04: 1.2; 05: 0.718; 06: 1.1727; 07: 0.5; 09: 0.8; 10: 0.9 |  | #N/A
| 1 | 3 | 02: 0.92; 03: 1.2; 04: 1.2; 05: 0.718; 06: 1.1799; 07: 0.58; 09: 0.8; 10: 0.8 |  | 16
| 1 | 4 | 02: 0.89; 03: 1.1; 04: 1.2; 05: 0.73; 06: 1.1121; 07: 0.54; 09: 0.8; 10: 0.8 |  | < 1.5
| 1 | 5 | 02: 0.91; 03: 1.2; 04: 1.2; 05: 0.73; 06: 1.1553; 07: 0.57; 09: 0.8; 10: 0.9 |  | 84
| 1 | 6 | 02: 0.83; 03: 1.2; 04: 1.3; 05: 0.732; 06: 1.2143; 07: 0.55; 09: 0.8; 10: 0.9
| Mean |  | 02: 0.88833; 03: 1.1667; 04: 1.2333; 05: 0.72633; 06: 1.1644; 07: 0.54667; 09: 0.8; 10: 0.86667
| Median |  | 02: 0.9; 03: 1.2; 04: 1.2; 05: 0.73; 06: 1.164; 07: 0.545; 09: 0.8; 10: 0.9
| Std Dev. |  | 02: 0.03656; 03: 0.05164; 04: 0.05164; 05: 0.0065013; 06: 0.033944; 07: 0.028048; 09: 1.2162e-16; 10: 0.05164
| Rel.Std.Dev. |  | 02: 4.116%; 03: 4.426%; 04: 4.187%; 05: 0.8951%; 06: 2.915%; 07: 5.131%; 09: 1.52e-14%; 10: 5.958%
PDM3
| Z-Score (Absolute) |  | 02: 0.03; 03: 0.89; 04: 1.1; 05: 0.47; 06: 0.88; 07: 1.02; 09: 0.24; 10: 0.03
| Analytical results for Fe in OREAS 922 (Certified Value 5.05 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  | 2
3
| 1 | 1 | 5.33 | 5.06 | 5.18 | 5.07 | 4.9 | 4.96 | 5.53 | 5.64 | 4.93 | 5.2509 | 5.16 | 4.952 | 4.63 | 4.87 | 5.295 | 4.97 | 4.91 | 4.97 |  |  | 1
| 1 | 2 | 5.27 | 5.11 | 5.3 | 5.08 | 5.08 | 4.98 | 5.41 | 5.51 | 4.89 | 5.2342 | 5.17 | 4.798 | 4.69 | 4.96 | 5.246 | 4.95 | 4.91 | 5.01 |  |  | #N/A
| 1 | 3 | 5.32 | 5.15 | 4.91 | 5.29 | 5.11 | 5.02 | 5.46 | 5.58 | 4.98 | 5.212 | 5.03 | 4.834 | 4.66 | 5.03 | 5.162 | 4.98 | 4.94 | 4.94 |  |  | 16
| 1 | 4 | 5.18 | 5.06 | 4.8 | 5.28 | 5.18 | 4.91 | 5.46 | 5.38 | 5.02 | 5.0733 | 5.04 | 4.784 | 4.85 | 4.97 | 5.247 | 4.96 | 4.88 | 4.85 |  |  | 5.0477
| 1 | 5 | 5.04 | 5 | 5.07 | 5.18 | 5.09 | 4.91 | 5.47 | 5.57 | 4.91 | 5.2756 | 4.98 | 4.667 | 4.77 | 5 | 5.185 | 5 | 4.87 | 4.93 |  |  | 85
| 1 | 6 | 5.23 | 5.02 | 5.17 | 5.05 | 4.93 | 4.8 | 5.46 | 5.65 | 5.04 | 5.3352 | 5.03 | 5.034 | 4.85 | 4.86 | 5.237 | 4.98 | 4.9 | 4.96
| Mean |  | 5.2283 | 5.0667 | 5.0717 | 5.1583 | 5.0483 | 4.93 | 5.465 | 5.555 | 4.9617 | 5.2302 | 5.0683 | 4.8448 | 4.7417 | 4.9483 | 5.2287 | 4.9733 | 4.9017 | 4.9433
| Median |  | 5.25 | 5.06 | 5.12 | 5.13 | 5.085 | 4.935 | 5.46 | 5.575 | 4.955 | 5.2425 | 5.035 | 4.816 | 4.73 | 4.965 | 5.2415 | 4.975 | 4.905 | 4.95
| Std Dev. |  | 0.10797 | 0.055737 | 0.18627 | 0.10797 | 0.10944 | 0.07642 | 0.038341 | 0.09975 | 0.061128 | 0.087712 | 0.077825 | 0.13031 | 0.096003 | 0.069113 | 0.047869 | 0.017512 | 0.024833 | 0.053541
| Rel.Std.Dev. |  | 2.065% | 1.1% | 3.673% | 2.093% | 2.168% | 1.55% | 0.7016% | 1.796% | 1.232% | 1.677% | 1.536% | 2.69% | 2.025% | 1.397% | 0.9155% | 0.3521% | 0.5066% | 1.083%
| PDM3 |  | 3.579% | 0.3764% | 0.4754% | 2.192% | 0.01318% | -2.331% | 8.268% | 10.05% | -1.704% | 3.616% | 0.4094% | -4.018% | -6.062% | -1.968% | 3.586% | -1.473% | -2.892% | -2.067%
| Z-Score (Absolute) |  | 0.95 | 0.05 | 0.08 | 0.56 | 0.05 | 0.71 | 2.27 | 2.78 | 0.53 | 0.96 | 0.06 | 1.19 | 1.76 | 0.61 | 0.96 | 0.47 | 0.87 | 0.64
| Analytical results for Ga in OREAS 922 (Certified Value 7.62 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
3
| 1 | 1 | 7.5 | 7.47 | 9.21 | 7.93 | 7.66 | 8.21 | 7.8 | < 5 | 8.3805 | 5.53 | 10 | 8.92 | 6.85 | 7.5 | 7.6 | 7.45 |  |  |  |  | 1
| 1 | 2 | 7.5 | 7.82 | 8.89 | 7.39 | 8.01 | 8.03 | 8 | 5 | 8.3537 | 5.56 | 11 | 8.94 | 6.84 | 7.4 | 7.8 | 7.63 |  |  |  |  | #N/A
| 1 | 3 | 7.5 | 7.7 | 9.12 | 7.32 | 7.86 | 8.18 | 8.2 | 10 | 8.0688 | 6.12 | 10 | 9.18 | 7 | 7.5 | 7.6 | 7.97 |  |  |  |  | 16
| 1 | 4 | 7.5 | 7.31 | 8.63 | 7.66 | 7.94 | 7.78 | 8.2 | 7 | 8.2729 | 5.8 | 10 | 9.12 | 6.93 | 7.6 | 7.7 | 7.91 |  |  |  |  | 7.623
| 1 | 5 | 7.4 | 7.3 | 9.5 | 8.06 | 7.66 | 7.43 | 7.8 | 7 | 8.619 | 5.56 | 11 | 8.94 | 6.87 | 7.5 | 7.6 | 7.65 |  |  |  |  | 86
| 1 | 6 | 7.4 | 7.24 | 9.66 | 8.52 | 7.6 | 7.47 | 7.8 | < 5 | 8.5789 | 5.76 | 10 | 9.14 | 6.74 | 7.3 | 7.7 | 7.79
| Mean |  | 7.4667 | 7.4733 | 9.1683 | 7.8133 | 7.7883 | 7.85 | 7.9667 | 7.25 | 8.379 | 5.7217 | 10.3333 | 9.04 | 6.8717 | 7.4667 | 7.6667 | 7.7333
| Median |  | 7.5 | 7.39 | 9.165 | 7.795 | 7.76 | 7.905 | 7.9 | 7 | 8.3671 | 5.66 | 10 | 9.03 | 6.86 | 7.5 | 7.65 | 7.72
| Std Dev. |  | 0.05164 | 0.2378 | 0.38028 | 0.45165 | 0.17069 | 0.34543 | 0.19664 | 2.0616 | 0.20281 | 0.22596 | 0.5164 | 0.11866 | 0.087958 | 0.10328 | 0.08165 | 0.19408
| Rel.Std.Dev. |  | 0.6916% | 3.182% | 4.148% | 5.78% | 2.192% | 4.4% | 2.468% | 28.44% | 2.42% | 3.949% | 4.997% | 1.313% | 1.28% | 1.383% | 1.065% | 2.51%
| PDM3 |  | -2.05% | -1.963% | 20.27% | 2.497% | 2.169% | 2.978% | 4.509% | -4.893% | 9.918% | -24.94% | 35.56% | 18.59% | -9.856% | -2.05% | 0.5732% | 1.448%
| Z-Score (Absolute) |  | 0.67 | 0.66 | 3.23 | 0.12 | 0.06 | 0.2 | 0.47 | 1.74 | 1.42 | 4.67 | 5.9 | 2.93 | 2.04 | 0.67 | 0.22 | 0.06
| Analytical results for Gd in OREAS 922 (Certified Value 4.44 ppm) | [21 empty] | Certified
Gd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  |  | 2
2
| 1 | 1 | 02: 4.28; 03: 5.1; 04: 5.6; 05: 3.3; 06: 5.0738; 07: 2.37; 09: 3.6; 10: 3.9 |  |  | 1
| 1 | 2 | 02: 4.38; 03: 5.1; 04: 5.4; 05: 3.24; 06: 5.2207; 07: 2.19; 09: 3.7; 10: 4.1 |  |  | #N/A
| 1 | 3 | 02: 4.43; 03: 5.2; 04: 5.4; 05: 3.25; 06: 5.1064; 07: 2.58; 09: 3.7; 10: 3.8 |  |  | 16
| 1 | 4 | 02: 4.36; 03: 5.1; 04: 5.3; 05: 3.39; 06: 5.2487; 07: 2.47; 09: 3.7; 10: 3.7 |  |  | 4.4363
| 1 | 5 | 02: 4.41; 03: 5.3; 04: 5.2; 05: 3.35; 06: 5.2796; 07: 2.44; 09: 3.8; 10: 3.8 |  |  | 87
| 1 | 6 | 02: 4.04; 03: 5.5; 04: 5.3; 05: 3.38; 06: 5.3526; 07: 2.53; 09: 3.7; 10: 3.9
| Mean |  | 02: 4.3167; 03: 5.2167; 04: 5.3667; 05: 3.3183; 06: 5.2136; 07: 2.43; 09: 3.7; 10: 3.8667
| Median |  | 02: 4.37; 03: 5.15; 04: 5.35; 05: 3.325; 06: 5.2347; 07: 2.455; 09: 3.7; 10: 3.85
| Std Dev. |  | 02: 0.14514; 03: 0.16021; 04: 0.13663; 05: 0.064936; 06: 0.10581; 07: 0.13813; 09: 0.063246; 10: 0.13663
| Rel.Std.Dev. |  | 02: 3.362%; 03: 3.071%; 04: 2.546%; 05: 1.957%; 06: 2.029%; 07: 5.684%; 09: 1.709%; 10: 3.533%
| PDM3 |  | 02: -2.696%; 03: 17.59%; 04: 20.97%; 05: -25.2%; 06: 17.52%; 07: -45.22%; 09: -16.6%; 10: -12.84%
| Z-Score (Absolute) |  | 02: 0.18; 03: 0.78; 04: 0.89; 05: 0.57; 06: 0.78; 07: 1.2; 09: 0.3; 10: 0.18
| Analytical results for Ge in OREAS 922 (Certified Value 0.1 ppm) | [21 empty] | Certified
Ge | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab |  |  | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 06; 06: 07; 07: 16; 09: 18; 10: 19; 11: 21; 13: 22 |  |  | ppm
|  |  | 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 09: AR*MS; 10: AR*OES/MS; 11: AR*OES/MS; 13: AR*MS |  |  | 2
2
| 1 | 1 | 02: < 0.1; 03: 0.2; 04: 0.1; 05: 0.09; 06: 0.13; 07: 0.208; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.09 |  |  | 1
| 1 | 2 | 02: < 0.1; 03: 0.1; 04: 0.1; 05: 0.1; 06: 0.12; 07: 0.206; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.1 |  |  | 10
| 1 | 3 | 02: < 0.1; 03: 0.2; 04: 0.1; 05: 0.08; 06: 0.13; 07: 0.201; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.11 |  |  | 16
| 1 | 4 | 02: < 0.1; 03: 0.1; 04: 0.1; 05: 0.09; 06: 0.12; 07: 0.201; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.16 |  |  | 0.102
| 1 | 5 | 02: < 0.1; 03: 0.2; 04: 0.1; 05: 0.08; 06: 0.11; 07: 0.193; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.11 |  |  | 88
| 1 | 6 | 02: < 0.1; 03: 0.2; 04: 0.1; 05: 0.09; 06: 0.12; 07: 0.204; 09: 0.1; 10: < 0.5; 11: 0.1; 13: 0.1
| Mean |  | 03: 0.16667; 04: 0.1; 05: 0.088333; 06: 0.12167; 07: 0.20217; 09: 0.1; 11: 0.1; 13: 0.11167
| Median |  | 03: 0.2; 04: 0.1; 05: 0.09; 06: 0.12; 07: 0.2025; 09: 0.1; 11: 0.1; 13: 0.105
| Std Dev. |  | 03: 0.05164; 04: 1.5202e-17; 05: 0.0075277; 06: 0.0075277; 07: 0.0052694; 09: 1.5202e-17; 11: 1.5202e-17; 13: 0.024833
| Rel.Std.Dev. |  | 03: 30.98%; 04: 1.52e-14%; 05: 8.522%; 06: 6.187%; 07: 2.606%; 09: 1.52e-14%; 11: 1.52e-14%; 13: 22.24%
| PDM3 |  | 03: 63.4%; 04: -1.961%; 05: -13.4%; 06: 19.28%; 07: 98.2%; 09: -1.961%; 11: -1.961%; 13: 9.477%
| Z-Score (Absolute) |  | 02: NA; 03: 6.48; 04: 0.1; 05: 1.25; 06: 2.04; 07: 9.98; 09: 0.1; 10: NA; 11: 0.1; 13: 0.1
| Analytical results for Hf in OREAS 922 (Certified Value 0.61 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
3
| 1 | 1 | 0.42 | 0.65 | < 0.1 | 0.4 | 0.61 | 0.6 | 0.92564 | 0.17 | 0.676 | < 0.1 | 0.6 | 0.6 | 0.64 | [7 empty] | 1
| 1 | 2 | 0.41 | 0.62 | < 0.1 | 0.5 | 0.67 | 0.59 | 0.94348 | 0.15 | 0.699 | < 0.1 | 0.7 | 0.6 | 0.62 | [7 empty] | #N/A
| 1 | 3 | 0.41 | 0.66 | < 0.1 | 0.5 | 0.65 | 0.6 | 0.95445 | 0.04 | 0.691 | < 0.1 | 0.6 | 0.6 | 0.62 | [7 empty] | 16
| 1 | 4 | 0.41 | 0.64 | < 0.1 | 0.5 | 0.65 | 0.6 | 1.0256 | 0.05 | 0.69 | < 0.1 | 0.6 | 0.64 | 0.59 | [7 empty] | 0.6109
| 1 | 5 | 0.41 | 0.66 | < 0.1 | 0.5 | 0.64 | 0.57 | 0.98037 | 0.05 | 0.672 | < 0.1 | 0.5 | 0.64 | 0.62 | [7 empty] | 89
| 1 | 6 | 0.41 | 0.63 | < 0.1 | 0.5 | 0.62 | 0.57 | 1.026 | 0.04 | 0.667 | < 0.1 | 0.6 | 0.61 | 0.56
| Mean |  | 0.41167 | 0.64333 |  | 0.48333 | 0.64 | 0.58833 | 0.97593 | 0.083333 | 0.6825 |  | 0.6 | 0.615 | 0.60833
| Median |  | 0.41 | 0.645 |  | 0.5 | 0.645 | 0.595 | 0.96741 | 0.05 | 0.683 |  | 0.6 | 0.605 | 0.62
| Std Dev. |  | 0.0040825 | 0.01633 |  | 0.040825 | 0.021909 | 0.01472 | 0.042529 | 0.059889 | 0.012598 |  | 0.063246 | 0.019748 | 0.028577
| Rel.Std.Dev. |  | 0.9917% | 2.538% |  | 8.447% | 3.423% | 2.502% | 4.358% | 71.87% | 1.846% |  | 10.54% | 3.211% | 4.698%
| PDM3 |  | -32.61% | 5.31% |  | -20.88% | 4.764% | -3.693% | 59.75% | -86.36% | 11.72% |  | -1.784% | 0.6718% | -0.4195%
| Z-Score (Absolute) |  | 4.84 | 0.67 | NA | 2.74 | 0.59 | 0.63 | 8.59 | 12.65 | 1.61 | NA | 0.36 | 0 | 0.07
| Analytical results for Hg in OREAS 922 (Certified Value < 0.05 ppm) | [21 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 3
3
| 1 | 1 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.02 | 0.03 | 0.04 | < 1 | < 0.01 | < 0.05 | < 5 | 0.016 | < 0.01 | < 1 | 0.06 | [5 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.03 | 0.03 | 0.04 | < 1 | < 0.01 | < 0.05 | < 5 | 0.023 | < 0.01 | < 1 | 0.02 | [5 empty] | 1
| 1 | 3 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.04 | 0.03 | 0.04 | < 1 | < 0.01 | < 0.05 | < 5 | 0.023 | < 0.01 | < 1 | 0.02 | [5 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.03 | 0.03 | 0.04 | < 1 | < 0.01 | < 0.05 | < 5 | 0.018 | < 0.01 | < 1 | 0.03 | [5 empty] | < 0.05
| 1 | 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.03 | 0.02 | 0.03 | < 1 | < 0.01 | < 0.05 | < 5 | 0.026 | < 0.01 | < 1 | 0.01 | [5 empty] | 90
| 1 | 6 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | 0.03 | 0.02 | 0.04 | < 1 | < 0.01 | < 0.05 | < 5 | 0.033 | < 0.01 | < 1 | 0.02
| Mean |  | 06: 0.03; 07: 0.026667; 09: 0.038333; 16: 0.023167; 19: 0.026667
| Median |  | 06: 0.03; 07: 0.03; 09: 0.04; 16: 0.023; 19: 0.02
| Std Dev. |  | 06: 0.0063246; 07: 0.005164; 09: 0.0040825; 16: 0.006047; 19: 0.017512
| Rel.Std.Dev. |  | 06: 21.08%; 07: 19.36%; 09: 10.65%; 16: 26.1%; 19: 65.67%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | 0.64 | 0 | 2.25 | NA | NA | NA | NA | 0.67 | NA | NA | 1.28
| Analytical results for Ho in OREAS 922 (Certified Value < 0.7 ppm) | [21 empty] | Certified
Ho | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  |  | 2
2
| 1 | 1 | 02: 0.6; 03: 0.8; 04: 0.9; 05: 0.46; 06: 0.77965; 07: 0.3; 09: 0.5; 10: 0.6 |  |  | 1
| 1 | 2 | 02: 0.64; 03: 0.8; 04: 0.8; 05: 0.455; 06: 0.72664; 07: 0.27; 09: 0.5; 10: 0.6 |  |  | #N/A
| 1 | 3 | 02: 0.63; 03: 0.8; 04: 0.8; 05: 0.45; 06: 0.72218; 07: 0.31; 09: 0.5; 10: 0.5 |  |  | 16
| 1 | 4 | 02: 0.62; 03: 0.8; 04: 0.8; 05: 0.465; 06: 0.75916; 07: 0.3; 09: 0.5; 10: 0.5 |  |  | < 0.7
| 1 | 5 | 02: 0.63; 03: 0.9; 04: 0.9; 05: 0.475; 06: 0.71042; 07: 0.32; 09: 0.5; 10: 0.5 |  |  | 91
| 1 | 6 | 02: 0.56; 03: 0.9; 04: 0.8; 05: 0.475; 06: 0.77647; 07: 0.3; 09: 0.5; 10: 0.6
| Mean |  | 02: 0.61333; 03: 0.83333; 04: 0.83333; 05: 0.46333; 06: 0.74575; 07: 0.3; 09: 0.5; 10: 0.55
| Median |  | 02: 0.625; 03: 0.8; 04: 0.8; 05: 0.4625; 06: 0.7429; 07: 0.3; 09: 0.5; 10: 0.55
| Std Dev. |  | 02: 0.029439; 03: 0.05164; 04: 0.05164; 05: 0.010328; 06: 0.029803; 07: 0.016733; 09: 0; 10: 0.054772
| Rel.Std.Dev. |  | 02: 4.8%; 03: 6.197%; 04: 6.197%; 05: 2.229%; 06: 3.996%; 07: 5.578%; 09: 0%; 10: 9.959%
PDM3
| Z-Score (Absolute) |  | 02: 0.18; 03: 1.18; 04: 1.18; 05: 0.59; 06: 0.76; 07: 1.37; 09: 0.42; 10: 0.18
| Analytical results for In in OREAS 922 (Certified Value 0.24 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
3
| 1 | 1 | 0.22 | 0.244 | 0.26 | 0.24 | 0.233 | 0.277 | 0.27 | 0.22015 | 0.198 | < 5 | 0.243 | 0.22 | 0.3 | 0.24 | 0.24 | [5 empty] | 1
| 1 | 2 | 0.22 | 0.255 | 0.26 | 0.24 | 0.246 | 0.272 | 0.28 | 0.2116 | 0.191 | < 5 | 0.244 | 0.22 | 0.2 | 0.24 | 0.244 | [5 empty] | #N/A
| 1 | 3 | 0.22 | 0.252 | 0.27 | 0.24 | 0.244 | 0.27 | 0.28 | 0.22938 | 0.201 | < 5 | 0.246 | 0.22 | 0.2 | 0.24 | 0.224 | [5 empty] | 16
| 1 | 4 | 0.22 | 0.251 | 0.26 | 0.24 | 0.233 | 0.257 | 0.29 | 0.229 | 0.201 | < 5 | 0.249 | 0.21 | 0.2 | 0.24 | 0.229 | [5 empty] | 0.237
| 1 | 5 | 0.22 | 0.251 | 0.26 | 0.25 | 0.239 | 0.248 | 0.28 | 0.22076 | 0.192 | < 5 | 0.25 | 0.22 | 0.2 | 0.23 | 0.232 | [5 empty] | 92
| 1 | 6 | 0.22 | 0.252 | 0.27 | 0.25 | 0.236 | 0.249 | 0.28 | 0.23302 | 0.196 | < 5 | 0.254 | 0.22 | 0.2 | 0.24 | 0.241
| Mean |  | 0.22 | 0.25083 | 0.26333 | 0.24333 | 0.2385 | 0.26217 | 0.28 | 0.22399 | 0.1965 |  | 0.24767 | 0.21833 | 0.21667 | 0.23833 | 0.235
| Median |  | 0.22 | 0.2515 | 0.26 | 0.24 | 0.2375 | 0.2635 | 0.28 | 0.22488 | 0.197 |  | 0.2475 | 0.22 | 0.2 | 0.24 | 0.236
| Std Dev. |  | 0 | 0.003656 | 0.005164 | 0.005164 | 0.0055408 | 0.012481 | 0.0063246 | 0.0079304 | 0.0043243 |  | 0.0041312 | 0.0040825 | 0.040825 | 0.0040825 | 0.0078486
| Rel.Std.Dev. |  | 0% | 1.458% | 1.961% | 2.122% | 2.323% | 4.761% | 2.259% | 3.541% | 2.201% |  | 1.668% | 1.87% | 18.84% | 1.713% | 3.34%
| PDM3 |  | -7.173% | 5.837% | 11.11% | 2.673% | 0.6334% | 10.62% | 18.14% | -5.491% | -17.09% |  | 4.501% | -7.876% | -8.579% | 0.563% | -0.8434%
| Z-Score (Absolute) |  | 0.76 | 0.51 | 1.02 | 0.2 | 0 | 0.98 | 1.71 | 0.59 | 1.72 | NA | 0.38 | 0.82 | 1.58 | 0 | 0.14
| Analytical results for K in OREAS 922 (Certified Value 0.376 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  | 3
3
| 1 | 1 | 0.34 | 0.3634 | 0.48 | 0.46 | 0.36 | 0.38 | 0.4 | 0.44 | 0.5364 | 0.2469 | 0.43 | 0.376 | 0.41 | 0.31 | 0.359 | 0.34 | 0.34 |  |  |  | 1
| 1 | 2 | 0.33 | 0.3787 | 0.52 | 0.45 | 0.37 | 0.37 | 0.38 | 0.45 | 0.5373 | 0.2401 | 0.42 | 0.367 | 0.4 | 0.31 | 0.364 | 0.34 | 0.34 |  |  |  | #N/A
| 1 | 3 | 0.34 | 0.3652 | 0.45 | 0.47 | 0.37 | 0.38 | 0.38 | 0.46 | 0.5175 | 0.2721 | 0.42 | 0.379 | 0.41 | 0.32 | 0.354 | 0.34 | 0.35 |  |  |  | 16
| 1 | 4 | 0.33 | 0.3636 | 0.44 | 0.48 | 0.37 | 0.38 | 0.38 | 0.47 | 0.5166 | 0.2675 | 0.41 | 0.378 | 0.41 | 0.32 | 0.353 | 0.33 | 0.35 |  |  |  | 0.37562
| 1 | 5 | 0.31 | 0.3638 | 0.47 | 0.47 | 0.36 | 0.37 | 0.38 | 0.46 | 0.5382 | 0.263 | 0.4 | 0.371 | 0.41 | 0.32 | 0.345 | 0.34 | 0.35 |  |  |  | 93
| 1 | 6 | 0.31 | 0.3736 | 0.49 | 0.46 | 0.34 | 0.35 | 0.38 | 0.45 | 0.5355 | 0.2553 | 0.4 | 0.381 | 0.42 | 0.31 | 0.352 | 0.34 | 0.34
| Mean |  | 0.32667 | 0.36805 | 0.475 | 0.465 | 0.36167 | 0.37167 | 0.38333 | 0.455 | 0.53025 | 0.25748 | 0.41333 | 0.37533 | 0.41 | 0.315 | 0.3545 | 0.33833 | 0.345
| Median |  | 0.33 | 0.3645 | 0.475 | 0.465 | 0.365 | 0.375 | 0.38 | 0.455 | 0.53595 | 0.25915 | 0.415 | 0.377 | 0.41 | 0.315 | 0.3535 | 0.34 | 0.345
| Std Dev. |  | 0.013663 | 0.006509 | 0.02881 | 0.010488 | 0.01169 | 0.01169 | 0.008165 | 0.010488 | 0.010268 | 0.012356 | 0.012111 | 0.0053166 | 0.0063246 | 0.0054772 | 0.006473 | 0.0040825 | 0.0054772
| Rel.Std.Dev. |  | 4.182% | 1.769% | 6.065% | 2.256% | 3.232% | 3.145% | 2.13% | 2.305% | 1.936% | 4.799% | 2.93% | 1.417% | 1.543% | 1.739% | 1.826% | 1.207% | 1.588%
| PDM3 |  | -13.03% | -2.015% | 26.46% | 23.8% | -3.714% | -1.052% | 2.054% | 21.13% | 41.17% | -31.45% | 10.04% | -0.07604% | 9.153% | -16.14% | -5.622% | -9.926% | -8.152%
| Z-Score (Absolute) |  | 0.79 | 0.1 | 1.82 | 1.64 | 0.18 | 0 | 0.15 | 1.47 | 2.79 | 2.01 | 0.73 | 0.06 | 0.67 | 1 | 0.3 | 0.59 | 0.47
| Analytical results for La in OREAS 922 (Certified Value 32.5 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
1
| 1 | 1 | 29.2 | 29.64 | 38 | 34 | 32.7 | 32.6 | 27.8 | 28 | 38.323 | 23.14 | 34 | 37.41 | 31 | 33.7 | 33.3 | 33 |  |  |  |  | 1
| 1 | 2 | 29.1 | 29.85 | 38 | 35 | 34 | 32.2 | 27.2 | 28 | 38.5332 | 22.24 | 34 | 37.84 | 31 | 33.8 | 32.9 | 33.1 |  |  |  |  | #N/A
| 1 | 3 | 29 | 30.52 | 36 | 37 | 33.3 | 33 | 28 | 29 | 38.8578 | 25.03 | 33 | 38.66 | 30 | 33 | 33.2 | 33.5 |  |  |  |  | 16
| 1 | 4 | 29 | 29.91 | 35 | 38 | 33.4 | 31.7 | 27.2 | 29 | 36.015 | 24.66 | 33 | 38.56 | 32 | 33.1 | 33.9 | 31 |  |  |  |  | 32.4515
| 1 | 5 | 28.8 | 30.69 | 37 | 37 | 32.6 | 30.5 | 28.3 | 29 | 38.8005 | 25.04 | 33 | 38.27 | 31 | 32.6 | 33.6 | 33.9 |  |  |  |  | 94
| 1 | 6 | 28.7 | 28.39 | 39 | 37 | 31.4 | 30.2 | 28.9 | 30 | 39.5791 | 24.45 | 33 | 38.69 | 31 | 34.5 | 33.6 | 32.7
| Mean |  | 28.9667 | 29.8333 | 37.1667 | 36.3333 | 32.9 | 31.7 | 27.9 | 28.8333 | 38.3514 | 24.0933 | 33.3333 | 38.2383 | 31 | 33.45 | 33.4167 | 32.8667
| Median |  | 29 | 29.88 | 37.5 | 37 | 33 | 31.95 | 27.9 | 29 | 38.6668 | 24.555 | 33 | 38.415 | 31 | 33.4 | 33.45 | 33.05
| Std Dev. |  | 0.18619 | 0.81635 | 1.472 | 1.5055 | 0.89443 | 1.1349 | 0.65727 | 0.75277 | 1.2213 | 1.1459 | 0.5164 | 0.51596 | 0.63246 | 0.68337 | 0.35449 | 1.0053
| Rel.Std.Dev. |  | 0.6428% | 2.736% | 3.96% | 4.144% | 2.719% | 3.58% | 2.356% | 2.611% | 3.184% | 4.756% | 1.549% | 1.349% | 2.04% | 2.043% | 1.061% | 3.059%
| PDM3 |  | -10.74% | -8.068% | 14.53% | 11.96% | 1.382% | -2.316% | -14.03% | -11.15% | 18.18% | -25.76% | 2.717% | 17.83% | -4.473% | 3.077% | 2.974% | 1.279%
| Z-Score (Absolute) |  | 0.85 | 0.67 | 0.85 | 0.68 | 0.04 | 0.28 | 1.07 | 0.88 | 1.19 | 1.86 | 0.05 | 1.07 | 0.43 | 0.08 | 0.07 | 0.04
| Analytical results for Li in OREAS 922 (Certified Value 22.8 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
2
| 1 | 1 | 24 | 20.6 | 25.2 | 22.8 | 23.1 | 23.5 | 21.1 | 21.803 | 18 | 24 | 22.635 | 23 | 20.3 | 23.7 | 24 | 23.4 |  |  |  |  | 1
| 1 | 2 | 24 | 20.9 | 24.9 | 22.5 | 24.3 | 23.2 | 21.3 | 22.3163 | 17.8 | 24 | 21.773 | 23 | 20.6 | 23.4 | 24 | 23.3 |  |  |  |  | #N/A
| 1 | 3 | 24 | 21.8 | 24.7 | 22.9 | 23.6 | 23.1 | 21 | 21.5409 | 20.6 | 24 | 22.451 | 22 | 21.2 | 23.2 | 24 | 23.9 |  |  |  |  | 16
| 1 | 4 | 23 | 20.4 | 25.1 | 24 | 23.7 | 22.8 | 21.6 | 21.3574 | 18.9 | 24 | 22.47 | 23 | 20.7 | 23.1 | 24 | 22.5 |  |  |  |  | 22.7993
| 1 | 5 | 22 | 20.4 | 25.6 | 24.1 | 23.6 | 21.9 | 21.2 | 21.9758 | 19.8 | 23 | 21.873 | 23 | 20.9 | 23.9 | 24 | 23.4 |  |  |  |  | 95
| 1 | 6 | 22 | 20.1 | 25.7 | 24 | 23.1 | 22.3 | 22.5 | 21.8762 | 19.7 | 23 | 22.643 | 25 | 20.5 | 24.1 | 24 | 23.3
| Mean |  | 23.1667 | 20.7 | 25.2 | 23.3833 | 23.5667 | 22.8 | 21.45 | 21.8116 | 19.1333 | 23.6667 | 22.3075 | 23.1667 | 20.7 | 23.5667 | 24 | 23.3
| Median |  | 23.5 | 20.5 | 25.15 | 23.45 | 23.6 | 22.95 | 21.25 | 21.8396 | 19.3 | 24 | 22.4605 | 23 | 20.65 | 23.55 | 24 | 23.35
| Std Dev. |  | 0.98319 | 0.6 | 0.38987 | 0.72503 | 0.44572 | 0.6 | 0.55408 | 0.33632 | 1.0985 | 0.5164 | 0.38504 | 0.98319 | 0.31623 | 0.39833 | 0 | 0.45166
| Rel.Std.Dev. |  | 4.244% | 2.899% | 1.547% | 3.101% | 1.891% | 2.632% | 2.583% | 1.542% | 5.741% | 2.182% | 1.726% | 4.244% | 1.528% | 1.69% | 0% | 1.938%
| PDM3 |  | 1.611% | -9.208% | 10.53% | 2.562% | 3.366% | 0.003185% | -5.918% | -4.332% | -16.08% | 3.804% | -2.157% | 1.611% | -9.208% | 3.366% | 5.267% | 2.196%
| Z-Score (Absolute) |  | 0.18 | 2.48 | 2.2 | 0.4 | 0.58 | 0.18 | 1.73 | 1.16 | 3.82 | 0.68 | 0.67 | 0.18 | 2.27 | 0.58 | 1.01 | 0.31
| Analytical results for Lu in OREAS 922 (Certified Value < 0.3 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.06 | 0.237 | 0.3 | 0.3 | 0.16 | 0.30127 | 0.1 | 0.188 | 0.2 | 0.2 | 0.19 | [9 empty] | 1
| 1 | 2 | 0.06 | 0.237 | 0.3 | 0.3 | 0.15 | 0.31082 | 0.09 | 0.194 | 0.2 | 0.2 | 0.19 | [9 empty] | #N/A
| 1 | 3 | 0.06 | 0.232 | 0.3 | 0.3 | 0.15 | 0.30696 | 0.109 | 0.198 | 0.2 | 0.2 | 0.2 | [9 empty] | 16
| 1 | 4 | 0.06 | 0.22 | 0.3 | 0.3 | 0.15 | 0.30566 | 0.11 | 0.197 | 0.2 | 0.2 | 0.19 | [9 empty] | < 0.3
| 1 | 5 | 0.05 | 0.225 | 0.3 | 0.3 | 0.16 | 0.30217 | 0.114 | 0.191 | 0.2 | 0.2 | 0.19 | [9 empty] | 96
| 1 | 6 | 0.05 | 0.195 | 0.3 | 0.3 | 0.16 | 0.30369 | 0.106 | 0.192 | 0.2 | 0.2 | 0.19
| Mean |  | 0.056667 | 0.22433 | 0.3 | 0.3 | 0.155 | 0.30509 | 0.10483 | 0.19333 | 0.2 | 0.2 | 0.19167
| Median |  | 0.06 | 0.2285 | 0.3 | 0.3 | 0.155 | 0.30467 | 0.1075 | 0.193 | 0.2 | 0.2 | 0.19
| Std Dev. |  | 0.005164 | 0.01587 | 0 | 0 | 0.0054772 | 0.0035128 | 0.0086352 | 0.0037771 | 3.0405e-17 | 3.0405e-17 | 0.0040825
| Rel.Std.Dev. |  | 9.113% | 7.074% | 0% | 0% | 3.534% | 1.151% | 8.237% | 1.954% | 1.52e-14% | 1.52e-14% | 2.13%
PDM3
| Z-Score (Absolute) |  | 2.15 | 0.45 | 1.5 | 1.5 | 0.67 | 1.57 | 1.43 | 0.1 | 0 | 0 | 0.15
| Analytical results for Mg in OREAS 922 (Certified Value 1.33 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  | 2
3
| 1 | 1 | 1.43 | 1.18 | 1.36 | 1.41 | 1.32 | 1.3 | 1.44 | 1.31 | 1.321 | 1.16 | 1.41 | 1.34 | 1.19 | 1.31 | 1.289 | 1.3 | 1.34 |  |  |  | 1
| 1 | 2 | 1.42 | 1.18 | 1.35 | 1.32 | 1.36 | 1.31 | 1.41 | 1.31 | 1.343 | 1.15 | 1.4 | 1.289 | 1.2 | 1.32 | 1.319 | 1.3 | 1.34 |  |  |  | #N/A
| 1 | 3 | 1.43 | 1.19 | 1.35 | 1.31 | 1.37 | 1.31 | 1.44 | 1.32 | 1.335 | 1.25 | 1.38 | 1.321 | 1.2 | 1.29 | 1.297 | 1.31 | 1.33 |  |  |  | 16
| 1 | 4 | 1.4 | 1.18 | 1.31 | 1.31 | 1.39 | 1.29 | 1.44 | 1.34 | 1.324 | 1.23 | 1.37 | 1.305 | 1.24 | 1.33 | 1.315 | 1.29 | 1.27 |  |  |  | 1.3348
| 1 | 5 | 1.35 | 1.17 | 1.39 | 1.37 | 1.36 | 1.29 | 1.44 | 1.32 | 1.341 | 1.21 | 1.36 | 1.275 | 1.22 | 1.33 | 1.305 | 1.29 | 1.32 |  |  |  | 97
| 1 | 6 | 1.34 | 1.17 | 1.4 | 1.48 | 1.32 | 1.25 | 1.42 | 1.34 | 1.338 | 1.19 | 1.37 | 1.348 | 1.24 | 1.34 | 1.313 | 1.3 | 1.33
| Mean |  | 1.395 | 1.1783 | 1.36 | 1.3667 | 1.3533 | 1.2917 | 1.4317 | 1.3233 | 1.3337 | 1.1983 | 1.3817 | 1.313 | 1.215 | 1.32 | 1.3063 | 1.2983 | 1.3217
| Median |  | 1.41 | 1.18 | 1.355 | 1.345 | 1.36 | 1.295 | 1.44 | 1.32 | 1.3365 | 1.2 | 1.375 | 1.313 | 1.21 | 1.325 | 1.309 | 1.3 | 1.33
| Std Dev. |  | 0.040373 | 0.0075277 | 0.032249 | 0.068313 | 0.028048 | 0.022286 | 0.013292 | 0.013663 | 0.0091141 | 0.0392 | 0.019408 | 0.028643 | 0.021679 | 0.017889 | 0.01157 | 0.0075277 | 0.026394
| Rel.Std.Dev. |  | 2.894% | 0.6388% | 2.371% | 4.999% | 2.072% | 1.725% | 0.9284% | 1.032% | 0.6834% | 3.271% | 1.405% | 2.181% | 1.784% | 1.355% | 0.8857% | 0.5798% | 1.997%
| PDM3 |  | 4.512% | -11.72% | 1.89% | 2.389% | 1.39% | -3.23% | 7.259% | -0.8574% | -0.08324% | -10.22% | 3.513% | -1.632% | -8.974% | -1.107% | -2.131% | -2.73% | -0.9823%
| Z-Score (Absolute) |  | 1.53 | 3.09 | 0.78 | 0.92 | 0.64 | 0.67 | 2.31 | 0 | 0.22 | 2.66 | 1.24 | 0.22 | 2.31 | 0.07 | 0.36 | 0.53 | 0.18
| Analytical results for Mn in OREAS 922 (Certified Value 0.073 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  | 3
3
| 1 | 1 | 0.0743 | 0.0712 | 0.0772 | 0.077 | 0.0733 | 0.0734 | 0.0758 | 0.0727 | 0.081805 | 0.0696 | 0.072 | 0.078358 | 0.0696 | 0.0723 | 0.0748 | 0.0758 | 0.0707 | 0.0723 |  |  | 1
| 1 | 2 | 0.0737 | 0.0717 | 0.0774 | 0.0741 | 0.0754 | 0.0721 | 0.0748 | 0.0731 | 0.080546 | 0.0686 | 0.07 | 0.075716 | 0.0699 | 0.073 | 0.0759 | 0.0758 | 0.0707 | 0.0719 |  |  | #N/A
| 1 | 3 | 0.0747 | 0.0721 | 0.0778 | 0.0732 | 0.0747 | 0.0753 | 0.0748 | 0.0732 | 0.080546 | 0.0718 | 0.07 | 0.077784 | 0.0699 | 0.0716 | 0.0751 | 0.0753 | 0.0715 | 0.0723 |  |  | 16
| 1 | 4 | 0.0713 | 0.0711 | 0.0779 | 0.0723 | 0.0757 | 0.0737 | 0.0744 | 0.0746 | 0.079722 | 0.0738 | 0.07 | 0.078165 | 0.0727 | 0.0729 | 0.0744 | 0.0755 | 0.0701 | 0.0712 |  |  | 0.073439
| 1 | 5 | 0.0733 | 0.0708 | 0.0809 | 0.0729 | 0.0745 | 0.0726 | 0.0742 | 0.0737 | 0.081061 | 0.0724 | 0.07 | 0.076451 | 0.0714 | 0.0733 | 0.0733 | 0.0756 | 0.0701 | 0.0717 |  |  | 98
| 1 | 6 | 0.0729 | 0.0709 | 0.0785 | 0.0773 | 0.0728 | 0.0716 | 0.0754 | 0.0744 | 0.080958 | 0.0709 | 0.07 | 0.078276 | 0.0719 | 0.0733 | 0.0746 | 0.0759 | 0.071 | 0.0721
| Mean |  | 0.073367 | 0.0713 | 0.078283 | 0.074467 | 0.0744 | 0.073117 | 0.0749 | 0.073617 | 0.080773 | 0.071183 | 0.070333 | 0.077458 | 0.0709 | 0.072733 | 0.074683 | 0.07565 | 0.070683 | 0.071917
| Median |  | 0.0735 | 0.07115 | 0.07785 | 0.07365 | 0.0746 | 0.073 | 0.0748 | 0.07345 | 0.080752 | 0.07135 | 0.07 | 0.077975 | 0.07065 | 0.07295 | 0.0747 | 0.0757 | 0.0707 | 0.072
| Std Dev. |  | 0.0012044 | 0.000502 | 0.0013586 | 0.0021602 | 0.0011454 | 0.0013258 | 0.00060332 | 0.00075741 | 0.00069148 | 0.0018957 | 0.0008165 | 0.0011075 | 0.0012791 | 0.00066533 | 0.00085654 | 0.00022583 | 0.00053821 | 0.00042151
| Rel.Std.Dev. |  | 1.642% | 0.7041% | 1.735% | 2.901% | 1.54% | 1.813% | 0.8055% | 1.029% | 0.8561% | 2.663% | 1.161% | 1.43% | 1.804% | 0.9148% | 1.147% | 0.2985% | 0.7614% | 0.5861%
| PDM3 |  | -0.09893% | -2.913% | 6.596% | 1.399% | 1.308% | -0.4393% | 1.989% | 0.2415% | 9.986% | -3.072% | -4.229% | 5.473% | -3.458% | -0.9613% | 1.694% | 3.01% | -3.753% | -2.073%
| Z-Score (Absolute) |  | 0.05 | 0.79 | 1.54 | 0.35 | 0.33 | 0.14 | 0.51 | 0.05 | 2.63 | 0.83 | 1.14 | 1.43 | 0.94 | 0.27 | 0.43 | 0.78 | 1.01 | 0.57
| Analytical results for Mo in OREAS 922 (Certified Value 0.69 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  |  | 2
2
| 1 | 1 | 1.51 | 0.7 | 0.55 | 0.71 | 0.64 | 0.77 | 0.77 | < 1 | 0.76173 | 0.6 | 4 | 0.879 | < 1 | 0.55 | 0.8 | < 2 | 0.69 | 0.73 |  |  | 1
| 1 | 2 | 1.56 | 0.7 | 0.53 | 0.69 | 0.68 | 0.73 | 0.73 | < 1 | 0.67648 | 0.6 | < 2 | 0.848 | 1 | 0.55 | 0.8 | < 2 | 0.77 | 0.7 |  |  | #N/A
| 1 | 3 | 1.69 | 0.7 | 0.6 | 0.7 | 0.64 | 0.74 | 0.74 | < 1 | 0.68421 | 0.6 | < 2 | 0.799 | 1 | 0.56 | 0.7 | < 2 | 0.69 | 0.74 |  |  | 16
| 1 | 4 | 1.63 | 0.6 | 0.54 | 0.72 | 0.71 | 0.73 | 0.76 | < 1 | 0.6608 | 0.7 | < 2 | 0.846 | < 1 | 0.6 | 0.7 | < 2 | 0.74 | 0.64 |  |  | 0.69174
| 1 | 5 | 1.34 | 0.6 | 0.6 | 0.68 | 0.68 | 0.7 | 0.77 | < 1 | 0.70541 | 0.6 | < 2 | 0.82 | < 1 | 0.59 | 0.8 | < 2 | 0.72 | 0.68 |  |  | 99
| 1 | 6 | 1.41 | 0.7 | 0.58 | 0.7 | 0.66 | 0.71 | 0.79 | < 1 | 0.69829 | 0.6 | < 2 | 0.897 | < 1 | 0.56 | 0.8 | < 2 | 0.73 | 0.69
| Mean |  | 1.5233 | 0.66667 | 0.56667 | 0.7 | 0.66833 | 0.73 | 0.76 |  | 0.69782 | 0.61667 | 4 | 0.84817 | 1 | 0.56833 | 0.76667 |  | 0.72333 | 0.69667
| Median |  | 1.535 | 0.7 | 0.565 | 0.7 | 0.67 | 0.73 | 0.765 |  | 0.69125 | 0.6 | 4 | 0.847 | 1 | 0.56 | 0.8 |  | 0.725 | 0.695
| Std Dev. |  | 0.13201 | 0.05164 | 0.030768 | 0.014142 | 0.027142 | 0.024495 | 0.021909 |  | 0.035082 | 0.040825 |  | 0.036196 | 0 | 0.02137 | 0.05164 |  | 0.030768 | 0.036148
| Rel.Std.Dev. |  | 8.666% | 7.746% | 5.43% | 2.02% | 4.061% | 3.355% | 2.883% |  | 5.027% | 6.62% |  | 4.268% | 0% | 3.76% | 6.736% |  | 4.254% | 5.189%
| PDM3 |  | 120.2% | -3.625% | -18.08% | 1.194% | -3.384% | 5.531% | 9.867% |  | 0.8784% | -10.85% | 478.2% | 22.61% | 44.56% | -17.84% | 10.83% |  | 4.567% | 0.7119%
| Z-Score (Absolute) |  | 10.59 | 0.59 | 1.89 | 0.15 | 0.57 | 0.24 | 0.63 | NA | 0.18 | 1.46 | 42.92 | 1.78 | 3.76 | 1.87 | 0.72 | NA | 0.15 | 0.2
| Analytical results for Na in OREAS 922 (Certified Value 0.021 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 3
3
| 1 | 1 | 0.02 | 0.02 | 0.027 | 0.025 | 0.01 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.031 | 0.03 | 0.019 | 0.034 | 0.02 | 0.02 |  |  |  |  | 1
| 1 | 2 | 0.02 | 0.02 | 0.028 | 0.022 | 0.01 | 0.02 | 0.02 | 0.02 | 0.022 | 0.02 | 0.032 | 0.03 | 0.019 | 0.037 | 0.02 | 0.02 |  |  |  |  | #N/A
| 1 | 3 | 0.02 | 0.02 | 0.025 | 0.022 | 0.02 | 0.02 | 0.02 | 0.02 | 0.021 | 0.02 | 0.031 | 0.03 | 0.018 | 0.035 | 0.02 | 0.02 |  |  |  |  | 16
| 1 | 4 | 0.02 | 0.02 | 0.024 | 0.02 | 0.02 | 0.02 | 0.03 | 0.03 | 0.022 | 0.02 | 0.031 | 0.03 | 0.02 | 0.033 | 0.02 | 0.02 |  |  |  |  | 0.020567
| 1 | 5 | 0.02 | 0.02 | 0.025 | 0.022 | 0.01 | 0.02 | 0.02 | 0.02 | 0.022 | 0.02 | 0.031 | 0.03 | 0.02 | 0.039 | 0.02 | 0.02 |  |  |  |  | 100
| 1 | 6 | 0.02 | 0.01 | 0.025 | 0.026 | 0.01 | 0.02 | 0.03 | 0.02 | 0.023 | 0.01 | 0.032 | 0.03 | 0.02 | 0.033 | 0.02 | 0.02
| Mean |  | 0.02 | 0.018333 | 0.025667 | 0.022833 | 0.013333 | 0.02 | 0.025 | 0.021667 | 0.021667 | 0.018333 | 0.031333 | 0.03 | 0.019333 | 0.035167 | 0.02 | 0.02
| Median |  | 0.02 | 0.02 | 0.025 | 0.022 | 0.01 | 0.02 | 0.025 | 0.02 | 0.022 | 0.02 | 0.031 | 0.03 | 0.0195 | 0.0345 | 0.02 | 0.02
| Std Dev. |  | 0 | 0.0040825 | 0.0015055 | 0.0022286 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0010328 | 0.0040825 | 0.0005164 | 0 | 0.0008165 | 0.0024014 | 0 | 0
| Rel.Std.Dev. |  | 0% | 22.27% | 5.866% | 9.76% | 38.73% | 0% | 21.91% | 18.84% | 4.767% | 22.27% | 1.648% | 0% | 4.223% | 6.829% | 0% | 0%
| PDM3 |  | -2.755% | -10.86% | 24.8% | 11.02% | -35.17% | -2.755% | 21.56% | 5.348% | 5.348% | -10.86% | 52.35% | 45.87% | -5.997% | 70.99% | -2.755% | -2.755%
| Z-Score (Absolute) |  | 0 | 0 | 3.01 | 1.64 | 3.85 | 0 | 2.89 | 0 | 0.96 | 0 | 6.55 | 5.78 | 0.39 | 8.77 | 0 | 0
| Analytical results for Nb in OREAS 922 (Certified Value 0.35 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
3
| 1 | 1 | 0.32 | 0.38 | < 0.1 | 0.6 | 0.33 | 0.4 | 0.35774 | 0.14 | < 5 | 0.362 | < 0.1 | < 1 | 0.58 | 0.37 | [6 empty] | 1
| 1 | 2 | 0.34 | 0.39 | < 0.1 | 0.6 | 0.35 | 0.38 | 0.31364 | 0.14 | < 5 | 0.403 | < 0.1 | < 1 | 0.59 | 0.38 | [6 empty] | #N/A
| 1 | 3 | 0.35 | 0.33 | < 0.1 | 0.5 | 0.33 | 0.39 | 0.31679 | 0.07 | < 5 | 0.362 | < 0.1 | < 1 | 0.59 | 0.4 | [6 empty] | 16
| 1 | 4 | 0.32 | 0.3 | < 0.1 | 0.4 | 0.32 | 0.41 | 0.33285 | 0.07 | < 5 | 0.361 | < 0.1 | < 1 | 0.59 | 0.3 | [6 empty] | 0.35274
| 1 | 5 | 0.32 | 0.39 | < 0.1 | 0.4 | 0.28 | 0.4 | 0.32162 | 0.07 | < 5 | 0.363 | < 0.1 | < 1 | 0.57 | 0.39 | [6 empty] | 101
| 1 | 6 | 0.33 | 0.36 | < 0.1 | 0.5 | 0.32 | 0.38 | 0.35678 | 0.07 | < 5 | 0.39 | < 0.1 | < 1 | 0.57 | 0.35
| Mean |  | 0.33 | 0.35833 |  | 0.5 | 0.32167 | 0.39333 | 0.33324 | 0.093333 |  | 0.3735 |  |  | 0.58167 | 0.365
| Median |  | 0.325 | 0.37 |  | 0.5 | 0.325 | 0.395 | 0.32723 | 0.07 |  | 0.3625 |  |  | 0.585 | 0.375
| Std Dev. |  | 0.012649 | 0.03656 |  | 0.089443 | 0.023166 | 0.012111 | 0.01972 | 0.036148 |  | 0.018295 |  |  | 0.0098319 | 0.036194
| Rel.Std.Dev. |  | 3.833% | 10.2% |  | 17.89% | 7.202% | 3.079% | 5.918% | 38.73% |  | 4.898% |  |  | 1.69% | 9.916%
| PDM3 |  | -6.446% | 1.586% |  | 41.75% | -8.809% | 11.51% | -5.529% | -73.54% |  | 5.886% |  |  | 64.9% | 3.476%
| Z-Score (Absolute) |  | 0.67 | 0.07 | NA | 2.95 | 0.85 | 0.67 | 0.61 | 5.71 | NA | 0.13 | NA | NA | 4.68 | 0.07
| Analytical results for Nd in OREAS 922 (Certified Value 27.5 ppm) | [21 empty] | Certified
Nd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 13: Lab; 14: Lab | [6 empty] | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 13: 18; 14: 19 | [6 empty] | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 13: AR*MS; 14: AR*OES/MS | [6 empty] | 1
1
| 1 | 1 | 02: 23.7; 03: 33.3; 04: 31.7; 05: 24; 06: 35.9785; 07: 19.82; 13: 24.2; 14: 27.6 | [6 empty] | 1
| 1 | 2 | 02: 24.57; 03: 33.6; 04: 30.5; 05: 24.1; 06: 34.7563; 07: 18.52; 13: 24.5; 14: 28.2 | [6 empty] | #N/A
| 1 | 3 | 02: 25.01; 03: 33.6; 04: 30.1; 05: 23.2; 06: 34.3159; 07: 21.17; 13: 24.4; 14: 27.2 | [6 empty] | 16
| 1 | 4 | 02: 24.56; 03: 33.4; 04: 30.2; 05: 23.4; 06: 34.5244; 07: 20.44; 13: 24.4; 14: 27.1 | [6 empty] | 27.4845
| 1 | 5 | 02: 24.93; 03: 34.5; 04: 31; 05: 22.6; 06: 34.3509; 07: 20.71; 13: 24.2; 14: 27.2 | [6 empty] | 102
| 1 | 6 | 02: 23.36; 03: 35.6; 04: 32.2; 05: 23.8; 06: 36.6782; 07: 20.48; 13: 23.8; 14: 27.7
| Mean |  | 02: 24.355; 03: 34; 04: 30.95; 05: 23.5167; 06: 35.1007; 07: 20.19; 13: 24.25; 14: 27.5
| Median |  | 02: 24.565; 03: 33.6; 04: 30.75; 05: 23.6; 06: 34.6403; 07: 20.46; 13: 24.3; 14: 27.4
| Std Dev. |  | 02: 0.67334; 03: 0.89219; 04: 0.85029; 05: 0.56716; 06: 0.98872; 07: 0.92771; 13: 0.251; 14: 0.41952
| Rel.Std.Dev. |  | 02: 2.765%; 03: 2.624%; 04: 2.747%; 05: 2.412%; 06: 2.817%; 07: 4.595%; 13: 1.035%; 14: 1.526%
| PDM3 |  | 02: -11.39%; 03: 23.71%; 04: 12.61%; 05: -14.44%; 06: 27.71%; 07: -26.54%; 13: -11.77%; 14: 0.05623%
| Z-Score (Absolute) |  | 02: 0.29; 03: 1.41; 04: 0.91; 05: 0.44; 06: 1.66; 07: 0.98; 13: 0.3; 14: 0.29
| Analytical results for Ni in OREAS 922 (Certified Value 34.3 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  |  |  | 1
2
| 1 | 1 | 34.1 | 34 | 38.8 | 35.8 | 32.7 | 38.7 | 37 | 34 | 35.8304 | 36 | 34.957 | 30 | 33.4 | 31.4 | 34.5 | 33.1 | 33.3 |  |  |  | 1
| 1 | 2 | 34.3 | 33 | 38.7 | 35.6 | 34.2 | 37.3 | 37 | 34 | 34.4091 | 35 | 34.941 | 31 | 34.9 | 32.4 | 34.6 | 33.2 | 34.4 |  |  |  | #N/A
| 1 | 3 | 34.4 | 34 | 39.1 | 35.6 | 33.1 | 38.6 | 35 | 34 | 34.3444 | 35 | 34.453 | 31 | 34.8 | 31.7 | 35.3 | 34 | 34.4 |  |  |  | 16
| 1 | 4 | 34.5 | 33 | 38.2 | 37.1 | 34 | 36.3 | 36 | 35 | 33.6953 | 35 | 34.662 | 32 | 34.7 | 32 | 34.2 | 34.7 | 34.8 |  |  |  | 34.2683
| 1 | 5 | 33.8 | 32 | 40.1 | 37 | 32.7 | 35.6 | 37 | 35 | 34.9773 | 35 | 35.194 | 32 | 34.3 | 31.1 | 33.7 | 32.9 | 34.1 |  |  |  | 103
| 1 | 6 | 33.7 | 38 | 39.9 | 36.4 | 32.4 | 36.2 | 35 | 35 | 35.3121 | 34 | 34.525 | 32 | 33.8 | 31.2 | 34.5 | 33.1 | 34.1
| Mean |  | 34.1333 | 34 | 39.1333 | 36.25 | 33.1833 | 37.1167 | 36.1667 | 34.5 | 34.7614 | 35 | 34.7887 | 31.3333 | 34.3167 | 31.6333 | 34.4667 | 33.5 | 34.1833
| Median |  | 34.2 | 33.5 | 38.95 | 36.1 | 32.9 | 36.8 | 36.5 | 34.5 | 34.6932 | 35 | 34.8015 | 31.5 | 34.5 | 31.55 | 34.5 | 33.15 | 34.25
| Std Dev. |  | 0.3266 | 2.0976 | 0.73394 | 0.68629 | 0.74677 | 1.3075 | 0.98319 | 0.54772 | 0.76515 | 0.63246 | 0.28778 | 0.8165 | 0.6047 | 0.50067 | 0.52409 | 0.70143 | 0.50365
| Rel.Std.Dev. |  | 0.9568% | 6.169% | 1.875% | 1.893% | 2.25% | 3.523% | 2.719% | 1.588% | 2.201% | 1.807% | 0.8272% | 2.606% | 1.762% | 1.583% | 1.521% | 2.094% | 1.473%
| PDM3 |  | -0.394% | -0.7831% | 14.2% | 5.783% | -3.166% | 8.312% | 5.54% | 0.676% | 1.439% | 2.135% | 1.518% | -8.565% | 0.141% | -7.689% | 0.5787% | -2.242% | -0.2481%
| Z-Score (Absolute) |  | 0.19 | 0.71 | 2.61 | 1 | 0.72 | 1.48 | 0.95 | 0.02 | 0.16 | 0.3 | 0.18 | 1.75 | 0.08 | 1.58 | 0 | 0.67 | 0.16
| Analytical results for P in OREAS 922 (Certified Value 0.063 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 21 | 22 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | [5 empty] | 3
3
| 1 | 1 | 0.071 | 0.054 | 0.051 | 0.058 | 0.066 | 0.066 | 0.066 | 0.077 | 0.0608 | 0.062 | 0.061614 | 0.0622 | 0.059 | 0.06 | 0.065 | [5 empty] | 1
| 1 | 2 | 0.069 | 0.0517 | 0.052 | 0.058 | 0.067 | 0.066 | 0.064 | 0.077 | 0.0627 | 0.064 | 0.060543 | 0.0624 | 0.06 | 0.06 | 0.066 | [5 empty] | #N/A
| 1 | 3 | 0.07 | 0.0536 | 0.049 | 0.061 | 0.068 | 0.066 | 0.064 | 0.076 | 0.06365 | 0.06 | 0.063935 | 0.0631 | 0.057 | 0.06 | 0.065 | [5 empty] | 16
| 1 | 4 | 0.069 | 0.0533 | 0.048 | 0.062 | 0.069 | 0.065 | 0.066 | 0.078 | 0.0627 | 0.062 | 0.064574 | 0.0654 | 0.06 | 0.07 | 0.063 | [5 empty] | 0.063434
| 1 | 5 | 0.067 | 0.0526 | 0.051 | 0.06 | 0.067 | 0.065 | 0.066 | 0.078 | 0.0627 | 0.06 | 0.062985 | 0.0652 | 0.06 | 0.06 | 0.064 | [5 empty] | 104
| 1 | 6 | 0.063 | 0.0537 | 0.052 | 0.059 | 0.064 | 0.064 | 0.066 | 0.078 | 0.06175 | 0.06 | 0.062632 | 0.0654 | 0.06 | 0.06 | 0.065
| Mean |  | 0.068167 | 0.05315 | 0.0505 | 0.059667 | 0.066833 | 0.065333 | 0.065333 | 0.077333 | 0.062383 | 0.061333 | 0.062714 | 0.06395 | 0.059333 | 0.061667 | 0.064667
| Median |  | 0.069 | 0.05345 | 0.051 | 0.0595 | 0.067 | 0.0655 | 0.066 | 0.0775 | 0.0627 | 0.061 | 0.062808 | 0.06415 | 0.06 | 0.06 | 0.065
| Std Dev. |  | 0.0028577 | 0.00085499 | 0.0016432 | 0.001633 | 0.0017224 | 0.0008165 | 0.0010328 | 0.0008165 | 0.00098116 | 0.001633 | 0.0014802 | 0.0015463 | 0.0012111 | 0.0040825 | 0.0010328
| Rel.Std.Dev. |  | 4.192% | 1.609% | 3.254% | 2.737% | 2.577% | 1.25% | 1.581% | 1.056% | 1.573% | 2.662% | 2.36% | 2.418% | 2.041% | 6.62% | 1.597%
| PDM3 |  | 7.46% | -16.21% | -20.39% | -5.94% | 5.358% | 2.993% | 2.993% | 21.91% | -1.657% | -3.312% | -1.136% | 0.8127% | -6.465% | -2.787% | 1.942%
| Z-Score (Absolute) |  | 1.61 | 2.38 | 3.03 | 0.76 | 1.02 | 0.65 | 0.65 | 3.63 | 0.08 | 0.34 | 0 | 0.31 | 0.72 | 0.67 | 0.49
| Analytical results for Pb in OREAS 922 (Certified Value 60 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  | 0
1
| 1 | 1 | 58.6 | 53.3 | 62 | 60.3 | 61.1 | 54.5 | 62 | 58.7 | 55 | 58.1318 | 45.9 | 65 | 62.348 | 55 | 57 | 55 | 65.1 | 54.6 | 64.2 |  | 1
| 1 | 2 | 58.5 | 54.2 | 66 | 57.6 | 63 | 59.2 | 58 | 59.6 | 58 | 59.6583 | 45 | 68 | 65.452 | 58 | 62 | 59 | 64.6 | 53.3 | 63.1 |  | #N/A
| 1 | 3 | 58.3 | 55.6 | 62 | 57.7 | 59.8 | 57.4 | 60 | 59.6 | 55 | 55.9961 | 48.5 | 65 | 62.88 | 57 | 65 | 63 | 61.9 | 53.3 | 63.3 |  | 16
| 1 | 4 | 58.2 | 56.8 | 63 | 58.1 | 60.7 | 53.4 | 62 | 59 | 57 | 58.7048 | 46.5 | 68 | 63.267 | 62 | 67 | 62 | 63 | 58.4 | 62.2 |  | 59.8774
| 1 | 5 | 57.9 | 58.2 | 66 | 57.4 | 62.8 | 53.1 | 63 | 59.7 | 57 | 56.84 | 47.3 | 66 | 63.325 | 61 | 65 | 62 | 66.3 | 53.6 | 61.3 |  | 105
| 1 | 6 | 58.1 | 58.6 | 65 | 59.3 | 56.9 | 54 | 63 | 57.5 | 56 | 56.34 | 47.3 | 66 | 63.32 | 60 | 62 | 57 | 60.4 | 53.3 | 62.1
| Mean |  | 58.2667 | 56.1167 | 64 | 58.4 | 60.7167 | 55.2667 | 61.3333 | 59.0167 | 56.3333 | 57.6118 | 46.75 | 66.3333 | 63.432 | 58.8333 | 63 | 59.6667 | 63.55 | 54.4167 | 62.7
| Median |  | 58.25 | 56.2 | 64 | 57.9 | 60.9 | 54.25 | 62 | 59.3 | 56.5 | 57.4859 | 46.9 | 66 | 63.2935 | 59 | 63.5 | 60.5 | 63.8 | 53.45 | 62.65
| Std Dev. |  | 0.2582 | 2.1377 | 1.8974 | 1.1524 | 2.2409 | 2.4655 | 1.9664 | 0.84242 | 1.2111 | 1.4475 | 1.226 | 1.3663 | 1.0594 | 2.6394 | 3.5214 | 3.2042 | 2.1916 | 2.0154 | 1.0334
| Rel.Std.Dev. |  | 0.4431% | 3.809% | 2.965% | 1.973% | 3.691% | 4.461% | 3.206% | 1.427% | 2.15% | 2.513% | 2.622% | 2.06% | 1.67% | 4.486% | 5.589% | 5.37% | 3.449% | 3.704% | 1.648%
| PDM3 |  | -2.69% | -6.281% | 6.885% | -2.467% | 1.402% | -7.7% | 2.432% | -1.437% | -5.919% | -3.784% | -21.92% | 10.78% | 5.936% | -1.744% | 5.215% | -0.3519% | 6.134% | -9.12% | 4.714%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 1.16 | 0.14 | 0.4 | 0.87 | 0.54 | 0 | 0.62 | 0.33 | 2.85 | 1.7 | 0.93 | 0.04 | 0.93 | 0.15 | 1.05 | 1.25 | 0.86
| Analytical results for Pd in OREAS 922 (Indicative Value < 10 ppb) | [21 empty] | Indicative
Pd | Round | Replicate | 02: Lab; 03: Lab |  | 1
| No | No | 02: 03; 03: 09 |  | ppb
|  |  | 02: AR*MS; 03: AR*MS |  | 0
0
| 1 | 1 | 02: < 10; 03: < 10 |  | 1
| 1 | 2 | 02: 14; 03: < 10 |  | 4
| 1 | 3 | 02: 11; 03: < 10 |  | 16
| 1 | 4 | 02: < 10; 03: < 10 |  | < 10
| 1 | 5 | 02: < 10; 03: < 10 |  | 10
| 1 | 6 | 02: < 10; 03: < 10
| Mean |  | 02: 12.5
| Median |  | 02: 12.5
| Std Dev. |  | 02: 2.1213
| Rel.Std.Dev. |  | 02: 16.97%
PDM3
Z-Score (Absolute) | bar: 02 to 03
| Analytical results for Pr in OREAS 922 (Certified Value 7.33 ppm) | [21 empty] | Certified
Pr | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 14; 07: 18; 09: 19 |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 09: AR*OES/MS |  | 2
3
| 1 | 1 | 02: 7.093; 03: 9.3; 04: 8.6; 05: 6.79; 06: 5.219; 07: 6.6; 09: 7.5 |  | 1
| 1 | 2 | 02: 7.291; 03: 9.4; 04: 8.3; 05: 6.86; 06: 4.927; 07: 6.7; 09: 7.5 |  | #N/A
| 1 | 3 | 02: 7.427; 03: 9.3; 04: 8.3; 05: 6.93; 06: 5.583; 07: 6.7; 09: 7.3 |  | 16
| 1 | 4 | 02: 7.127; 03: 9.3; 04: 8.4; 05: 6.84; 06: 5.37; 07: 6.7; 09: 7.3 |  | 7.3255
| 1 | 5 | 02: 7.387; 03: 9.6; 04: 8.7; 05: 6.81; 06: 5.478; 07: 6.7; 09: 7.3 |  | 106
| 1 | 6 | 02: 6.83; 03: 10; 04: 8.8; 05: 6.88; 06: 5.443; 07: 6.6; 09: 7.5
| Mean |  | 02: 7.1925; 03: 9.4833; 04: 8.5167; 05: 6.8517; 06: 5.3367; 07: 6.6667; 09: 7.4
| Median |  | 02: 7.209; 03: 9.35; 04: 8.5; 05: 6.85; 06: 5.4065; 07: 6.7; 09: 7.4
| Std Dev. |  | 02: 0.22276; 03: 0.27869; 04: 0.2137; 05: 0.050365; 06: 0.23441; 07: 0.05164; 09: 0.10954
| Rel.Std.Dev. |  | 02: 3.097%; 03: 2.939%; 04: 2.509%; 05: 0.7351%; 06: 4.392%; 07: 0.7746%; 09: 1.48%
| PDM3 |  | 02: -1.816%; 03: 29.46%; 04: 16.26%; 05: -6.468%; 06: -27.15%; 07: -8.994%; 09: 1.017%
| Z-Score (Absolute) |  | 02: 0; 03: 2.81; 04: 1.7; 05: 0.44; 06: 2.38; 07: 0.67; 09: 0.27
| Analytical results for Pt in OREAS 922 (Indicative Value < 5 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | 02: Lab; 03: Lab |  | 1
| No | No | 02: 03; 03: 09 |  | ppb
|  |  | 02: AR*MS; 03: AR*MS |  | 0
0
| 1 | 1 | 02: < 5; 03: < 5 |  | 1
| 1 | 2 | 02: < 5; 03: < 5 |  | 3
| 1 | 3 | 02: < 5; 03: < 5 |  | 16
| 1 | 4 | 02: < 5; 03: < 5 |  | < 5
| 1 | 5 | 02: < 5; 03: < 5 |  | 9
| 1 | 6 | 02: < 5; 03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02 to 03
| Analytical results for Rb in OREAS 922 (Certified Value 22.7 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
1
| 1 | 1 | 20.3 | 26.41 | 29.2 | 30.1 | 20.1 | 20.4 | 23 | 28.9322 | 12.29 | 23.59 | 17.1 | 22.8 | 21.4 | 20 | [6 empty] | 1
| 1 | 2 | 19.8 | 28.35 | 28.7 | 29.5 | 21.4 | 19.7 | 22 | 28.1878 | 11.88 | 24.09 | 17.3 | 22.8 | 21.5 | 20.5 | [6 empty] | #N/A
| 1 | 3 | 19.8 | 26.18 | 28.9 | 30.3 | 20.7 | 20.2 | 22.4 | 29.0069 | 13.42 | 24.8 | 17.5 | 22.2 | 21.6 | 20.6 | [6 empty] | 16
| 1 | 4 | 20.3 | 25.43 | 29.2 | 31.1 | 20.6 | 19.9 | 23 | 26.8221 | 12.73 | 24.3 | 17.4 | 21.8 | 21.6 | 19.7 | [6 empty] | 22.7156
| 1 | 5 | 19.6 | 25.52 | 30.4 | 31.7 | 19.6 | 18.9 | 23.4 | 28.4365 | 12.85 | 24.19 | 17.8 | 22.5 | 21.3 | 21 | [6 empty] | 107
| 1 | 6 | 19.8 | 24.57 | 31 | 31.8 | 19.5 | 18.7 | 23.5 | 32.3302 | 12.63 | 24.43 | 17.2 | 22.8 | 21.7 | 19.9
| Mean |  | 19.9333 | 26.0767 | 29.5667 | 30.75 | 20.3167 | 19.6333 | 22.8833 | 28.9526 | 12.6333 | 24.2333 | 17.3833 | 22.4833 | 21.5167 | 20.2833
| Median |  | 19.8 | 25.85 | 29.2 | 30.7 | 20.35 | 19.8 | 23 | 28.6843 | 12.68 | 24.245 | 17.35 | 22.65 | 21.55 | 20.25
| Std Dev. |  | 0.29439 | 1.2875 | 0.91797 | 0.92898 | 0.72503 | 0.69186 | 0.58109 | 1.8331 | 0.52164 | 0.40013 | 0.24833 | 0.41191 | 0.1472 | 0.49565
| Rel.Std.Dev. |  | 1.477% | 4.937% | 3.105% | 3.021% | 3.569% | 3.524% | 2.539% | 6.332% | 4.129% | 1.651% | 1.429% | 1.832% | 0.6841% | 2.444%
| PDM3 |  | -12.25% | 14.8% | 30.16% | 35.37% | -10.56% | -13.57% | 0.7385% | 27.46% | -44.38% | 6.682% | -23.47% | -1.022% | -5.278% | -10.71%
| Z-Score (Absolute) |  | 0.61 | 1.2 | 2.22 | 2.57 | 0.49 | 0.69 | 0.26 | 1.84 | 2.75 | 0.65 | 1.35 | 0.14 | 0.14 | 0.5
| Analytical results for Re in OREAS 922 (Certified Value < 1 ppb) | [21 empty] | Certified
Re | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 06; 06: 07; 07: 13; 09: 14; 10: 18; 11: 19; 13: 22 |  | ppb
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 09: AR*MS; 10: AR*MS; 11: AR*OES/MS; 13: AR*MS |  | 0
0
| 1 | 1 | 02: < 1; 03: < 1; 04: < 1; 05: < 1; 06: < 1; 07: < 1; 09: < 1; 10: < 1; 11: < 10; 13: < 1 |  | 1
| 1 | 2 | 02: < 1; 03: < 1; 04: < 1; 05: < 1; 06: < 1; 07: < 1; 09: < 1; 10: < 1; 11: < 10; 13: < 1 |  | 8
| 1 | 3 | 02: < 1; 03: < 1; 04: < 1; 05: < 1; 06: < 1; 07: < 1; 09: < 1; 10: < 1; 11: < 10; 13: < 1 |  | 16
| 1 | 4 | 02: < 1; 03: < 1; 04: < 1; 05: 1; 06: < 1; 07: < 1; 09: < 1; 10: 1; 11: < 10; 13: < 1 |  | < 1
| 1 | 5 | 02: < 1; 03: < 1; 04: < 1; 05: < 1; 06: < 1; 07: < 1; 09: < 1; 10: < 1; 11: < 10; 13: < 1 |  | 108
| 1 | 6 | 02: < 1; 03: < 1; 04: < 1; 05: 1; 06: < 1; 07: < 1; 09: < 1; 10: < 1; 11: < 10; 13: < 1
Mean | bar: 05; 10
Median | bar: 05; 10
| Std Dev. |  | 05: 0
Rel.Std.Dev. | bar: 05
PDM3
| Z-Score (Absolute) |  | 02: NA; 03: NA; 04: NA; 05: 0.67; 06: NA; 07: NA; 09: NA; 10: 0.67; 11: NA; 13: NA
| Analytical results for Ru in OREAS 922 (Indicative Value < 0.005 ppm) | [21 empty] | Indicative
Ru | Round | Replicate | 02: Lab; 03: Lab |  | 1
| No | No | 02: 09; 03: 19 |  | ppm
|  |  | 02: AR*MS; 03: AR*OES/MS |  | 3
3
| 1 | 1 | 02: < 0.005; 03: < 0.1 |  | 1
| 1 | 2 | 02: < 0.005; 03: < 0.1 |  | 4
| 1 | 3 | 02: < 0.005; 03: < 0.1 |  | 16
| 1 | 4 | 02: < 0.005; 03: < 0.1 |  | < 0.005
| 1 | 5 | 02: < 0.005; 03: < 0.1 |  | 10
| 1 | 6 | 02: < 0.005; 03: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02 to 03
| Analytical results for S in OREAS 922 (Certified Value 0.386 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 15 | 16 | 18 | 19 | 20 | 21 | 22 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  | 3
3
| 1 | 1 | 0.43 | 0.3513 | 0.36 | 0.36 | 0.41 | 0.4 | 0.385 | 0.3975 | 0.41 | 0.3894 | 0.4 | 0.357 | 0.32 | 0.357 | 0.392 | 0.38 | 0.41 |  |  |  | 1
| 1 | 2 | 0.43 | 0.3594 | 0.37 | 0.34 | 0.42 | 0.4 | 0.385 | 0.3868 | 0.41 | 0.3916 | 0.4 | 0.351 | 0.32 | 0.368 | 0.387 | 0.38 | 0.41 |  |  |  | 6
| 1 | 3 | 0.43 | 0.3619 | 0.35 | 0.36 | 0.42 | 0.4 | 0.395 | 0.3862 | 0.41 | 0.387 | 0.4 | 0.345 | 0.31 | 0.361 | 0.4 | 0.38 | 0.41 |  |  |  | 16
| 1 | 4 | 0.43 | 0.3628 | 0.34 | 0.36 | 0.43 | 0.4 | 0.395 | 0.3798 | 0.42 | 0.3773 | 0.4 | 0.358 | 0.32 | 0.365 | 0.396 | 0.39 | 0.39 |  |  |  | 0.3863
| 1 | 5 | 0.43 | 0.365 | 0.36 | 0.35 | 0.42 | 0.39 | 0.395 | 0.399 | 0.41 | 0.385 | 0.39 | 0.35 | 0.32 | 0.365 | 0.4 | 0.38 | 0.4 |  |  |  | 109
| 1 | 6 | 0.43 | 0.3582 | 0.34 | 0.34 | 0.4 | 0.39 | 0.395 | 0.3972 | 0.41 | 0.374 | 0.39 | 0.341 | 0.32 | 0.363 | 0.391 | 0.38 | 0.41
| Mean |  | 0.43 | 0.35977 | 0.35333 | 0.35167 | 0.41667 | 0.39667 | 0.39167 | 0.39108 | 0.41167 | 0.38405 | 0.39667 | 0.35033 | 0.31833 | 0.36317 | 0.39433 | 0.38167 | 0.405
| Median |  | 0.43 | 0.36065 | 0.355 | 0.355 | 0.42 | 0.4 | 0.395 | 0.392 | 0.41 | 0.386 | 0.4 | 0.3505 | 0.32 | 0.364 | 0.394 | 0.38 | 0.41
| Std Dev. |  | 0 | 0.0048044 | 0.012111 | 0.0098319 | 0.010328 | 0.005164 | 0.005164 | 0.0078838 | 0.0040825 | 0.0069541 | 0.005164 | 0.0066232 | 0.0040825 | 0.0038166 | 0.0052409 | 0.0040825 | 0.0083666
| Rel.Std.Dev. |  | 0% | 1.335% | 3.428% | 2.796% | 2.479% | 1.302% | 1.318% | 2.016% | 0.9917% | 1.811% | 1.302% | 1.891% | 1.282% | 1.051% | 1.329% | 1.07% | 2.066%
| PDM3 |  | 11.31% | -6.868% | -8.533% | -8.964% | 7.862% | 2.685% | 1.39% | 1.239% | 6.568% | -0.5814% | 2.685% | -9.31% | -17.59% | -5.987% | 2.081% | -1.198% | 4.842%
| Z-Score (Absolute) |  | 1.27 | 1.03 | 1.24 | 1.29 | 0.84 | 0.18 | 0.02 | 0 | 0.67 | 0.23 | 0.18 | 1.33 | 2.33 | 0.91 | 0.11 | 0.31 | 0.55
| Analytical results for Sb in OREAS 922 (Certified Value 0.57 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  | 2
2
| 1 | 1 | < 0.05 | 0.5 | 0.13 | 0.66 | 0.52 | 0.72 | 0.48 | 0.63 | < 3 | 0.57369 | 0.57 | < 5 | 0.409 | < 5 | 0.21 | < 0.5 | < 20 | 0.53 | 0.69 |  | 1
| 1 | 2 | < 0.05 | 0.53 | 0.13 | 0.62 | 0.56 | 0.71 | 0.46 | 0.63 | < 3 | 0.58472 | 0.54 | < 5 | 0.438 | < 5 | 0.27 | 0.5 | < 20 | 0.55 | 0.66 |  | #N/A
| 1 | 3 | < 0.05 | 0.52 | 0.12 | 0.63 | 0.58 | 0.73 | 0.46 | 0.63 | < 3 | 0.58968 | 0.6 | < 5 | 0.458 | < 5 | 0.28 | 0.5 | < 20 | 0.55 | 0.68 |  | 16
| 1 | 4 | < 0.05 | 0.49 | 0.18 | 0.72 | 0.55 | 0.67 | 0.5 | 0.65 | < 3 | 0.61239 | 0.59 | < 5 | 0.416 | < 5 | 0.3 | 0.5 | < 20 | 0.54 | 0.62 |  | 0.56872
| 1 | 5 | < 0.05 | 0.49 | 0.16 | 0.65 | 0.54 | 0.63 | 0.5 | 0.62 | < 3 | 0.63388 | 0.54 | < 5 | 0.427 | < 5 | 0.24 | 0.5 | < 20 | 0.56 | 0.65 |  | 110
| 1 | 6 | < 0.05 | 0.5 | 0.15 | 0.68 | 0.55 | 0.65 | 0.5 | 0.66 | < 3 | 0.57769 | 0.61 | < 5 | 0.408 | < 5 | 0.26 | 0.5 | < 20 | 0.58 | 0.64
| Mean |  |  | 0.505 | 0.145 | 0.66 | 0.55 | 0.685 | 0.48333 | 0.63667 |  | 0.59534 | 0.575 |  | 0.426 |  | 0.26 | 0.5 |  | 0.55167 | 0.65667
| Median |  |  | 0.5 | 0.14 | 0.655 | 0.55 | 0.69 | 0.49 | 0.63 |  | 0.5872 | 0.58 |  | 0.4215 |  | 0.265 | 0.5 |  | 0.55 | 0.655
| Std Dev. |  |  | 0.016432 | 0.022583 | 0.036332 | 0.02 | 0.040866 | 0.019664 | 0.015055 |  | 0.023242 | 0.030166 |  | 0.019401 |  | 0.031623 | 0 |  | 0.017224 | 0.02582
| Rel.Std.Dev. |  |  | 3.254% | 15.57% | 5.505% | 3.636% | 5.966% | 4.068% | 2.365% |  | 3.904% | 5.246% |  | 4.554% |  | 12.16% | 0% |  | 3.122% | 3.932%
| PDM3 |  |  | -11.2% | -74.5% | 16.05% | -3.292% | 20.45% | -15.01% | 11.95% |  | 4.68% | 1.104% |  | -25.1% |  | -54.28% | -12.08% |  | -2.999% | 15.46%
| Z-Score (Absolute) |  | NA | 0.4 | 3.57 | 0.96 | 0.01 | 1.18 | 0.59 | 0.75 | NA | 0.39 | 0.21 | NA | 1.1 | NA | 2.56 | 0.45 | NA | 0.01 | 0.93
| Analytical results for Sc in OREAS 922 (Certified Value 3.15 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
2
| 1 | 1 | 1.8 | 3 | 3 | 4 | 3 | 3.7 | 3 | 8 | 4.4311 | 2.1 | < 5 | 3.64 | 2.7 | 3 | 3.2 | 3.3 |  |  |  |  | 1
| 1 | 2 | 1.7 | 3 | 3 | 4 | 3.2 | 3.6 | 3 | 8 | 4.4445 | 2 | < 5 | 3.72 | 2.8 | 3 | 3.3 | 3.4 |  |  |  |  | 9
| 1 | 3 | 1.8 | 3 | 3 | 4 | 3.1 | 3.6 | 3 | 8 | 4.3639 | 2.2 | < 5 | 3.78 | 2.8 | 3 | 3.2 | 3.4 |  |  |  |  | 16
| 1 | 4 | 1.7 | 3 | 3 | 4 | 3.1 | 3.6 | 3 | 8 | 4.3856 | 2.2 | < 5 | 3.81 | 2.7 | 3 | 3.3 | 3.1 |  |  |  |  | 3.1472
| 1 | 5 | 1.7 | 3 | 3 | 4 | 3 | 3.3 | 3 | 8 | 4.5655 | 2.1 | < 5 | 3.64 | 2.8 | 3 | 3.2 | 3.4 |  |  |  |  | 111
| 1 | 6 | 1.8 | 3 | 3 | 4 | 3 | 3.3 | 3 | 8 | 4.6152 | 2.2 | < 5 | 3.71 | 2.8 | 3 | 3.3 | 3.3
| Mean |  | 1.75 | 3 | 3 | 4 | 3.0667 | 3.5167 | 3 | 8 | 4.4676 | 2.1333 |  | 3.7167 | 2.7667 | 3 | 3.25 | 3.3167
| Median |  | 1.75 | 3 | 3 | 4 | 3.05 | 3.6 | 3 | 8 | 4.4378 | 2.15 |  | 3.715 | 2.8 | 3 | 3.25 | 3.35
| Std Dev. |  | 0.054772 | 0 | 0 | 0 | 0.08165 | 0.17224 | 0 | 0 | 0.10071 | 0.08165 |  | 0.070048 | 0.05164 | 0 | 0.054772 | 0.1169
| Rel.Std.Dev. |  | 3.13% | 0% | 0% | 0% | 2.662% | 4.898% | 0% | 0% | 2.254% | 3.827% |  | 1.885% | 1.866% | 0% | 1.685% | 3.525%
| PDM3 |  | -44.4% | -4.678% | -4.678% | 27.1% | -2.56% | 11.74% | -4.678% | 154.2% | 41.95% | -32.22% |  | 18.09% | -12.09% | -4.678% | 3.266% | 5.384%
| Z-Score (Absolute) |  | 2.96 | 0.15 | 0.15 | 2.1 | 0 | 1.01 | 0.15 | 11.09 | 3.15 | 2.1 | NA | 1.46 | 0.67 | 0.15 | 0.41 | 0.56
| Analytical results for Se in OREAS 922 (Certified Value 3.44 ppm) | [21 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
2
| 1 | 1 | < 1 | 4 | 2.9 | 3.2 | 4.3 | 4 | 3.2 | 3 | 3.4941 | 5 | < 5 | 3.824 | 7 | 2.8 | 4 | 3 | 3.4 |  |  |  | 1
| 1 | 2 | < 1 | 4 | 3.1 | 3.2 | 4.3 | 3.9 | 3.2 | 3 | 3.8377 | 5 | < 5 | 3.526 | < 5 | 2.9 | 4 | 3 | 3.4 |  |  |  | 10
| 1 | 3 | < 1 | 4 | 3.2 | 3.2 | 4.1 | 4 | 3.4 | 4 | 3.9147 | 5 | < 5 | 2.958 | < 5 | 3.1 | 3 | 3 | 3.2 |  |  |  | 16
| 1 | 4 | < 1 | 4 | 2.7 | 3.5 | 4.4 | 3.8 | 3.4 | 3 | 3.7488 | 5 | < 5 | 2.685 | < 5 | 3.1 | 3 | 3 | 4.2 |  |  |  | 3.4419
| 1 | 5 | < 1 | 4 | 3.3 | 3.5 | 4.1 | 3.7 | 3.2 | 3 | 3.5493 | 5 | < 5 | 2.384 | < 5 | 3.2 | 4 | 3 | 2.7 |  |  |  | 112
| 1 | 6 | < 1 | 3 | 3.2 | 3.3 | 4 | 3.8 | 3.4 | 3 | 3.4898 | 5 | < 5 | 3.853 | 6 | 3 | 4 | 3 | 3.7
| Mean |  |  | 3.8333 | 3.0667 | 3.3167 | 4.2 | 3.8667 | 3.3 | 3.1667 | 3.6724 | 5 |  | 3.205 | 6.5 | 3.0167 | 3.6667 | 3 | 3.4333
| Median |  |  | 4 | 3.15 | 3.25 | 4.2 | 3.85 | 3.3 | 3 | 3.649 | 5 |  | 3.242 | 6.5 | 3.05 | 4 | 3 | 3.4
| Std Dev. |  |  | 0.40825 | 0.22509 | 0.1472 | 0.15492 | 0.12111 | 0.10954 | 0.40825 | 0.18555 | 0 |  | 0.61831 | 0.70711 | 0.1472 | 0.5164 | 0 | 0.50067
| Rel.Std.Dev. |  |  | 10.65% | 7.34% | 4.438% | 3.689% | 3.132% | 3.32% | 12.89% | 5.053% | 0% |  | 19.29% | 10.88% | 4.879% | 14.08% | 0% | 14.58%
| PDM3 |  |  | 11.37% | -10.9% | -3.637% | 22.03% | 12.34% | -4.121% | -7.995% | 6.698% | 45.27% |  | -6.881% | 88.85% | -12.35% | 6.532% | -12.84% | -0.2474%
| Z-Score (Absolute) |  | NA | 0.92 | 0.59 | 0.19 | 1.24 | 0.7 | 0.22 | 0.7 | 0.39 | 2.54 | NA | 0.37 | 4.96 | 0.67 | 0.38 | 0.7 | 0
| Analytical results for Sm in OREAS 922 (Certified Value 4.98 ppm) | [21 empty] | Certified
Sm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  |  |  | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 10: 19 |  |  |  | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 10: AR*OES/MS |  |  |  | 2
2
| 1 | 1 | 02: 4.59; 03: 6.1; 04: 5.8; 05: 4.27; 06: 6.511; 07: 3.37; 09: 4.2; 10: 5 |  |  |  | 1
| 1 | 2 | 02: 4.83; 03: 6.1; 04: 5.4; 05: 4.26; 06: 6.5715; 07: 3.24; 09: 4.3; 10: 5 |  |  |  | #N/A
| 1 | 3 | 02: 4.86; 03: 6.1; 04: 5.4; 05: 4.25; 06: 6.3736; 07: 3.68; 09: 4.3; 10: 5.1 |  |  |  | 16
| 1 | 4 | 02: 4.7; 03: 6.1; 04: 5.3; 05: 4.31; 06: 6.4539; 07: 3.55; 09: 4.2; 10: 4.8 |  |  |  | 4.9759
| 1 | 5 | 02: 4.82; 03: 6.3; 04: 5.5; 05: 4.36; 06: 6.3855; 07: 3.56; 09: 4.2; 10: 5.1 |  |  |  | 113
| 1 | 6 | 02: 4.39; 03: 6.5; 04: 5.6; 05: 4.24; 06: 6.686; 07: 3.54; 09: 4.1; 10: 4.9
| Mean |  | 02: 4.6983; 03: 6.2; 04: 5.5; 05: 4.2817; 06: 6.4969; 07: 3.49; 09: 4.2167; 10: 4.9833
| Median |  | 02: 4.76; 03: 6.1; 04: 5.45; 05: 4.265; 06: 6.4824; 07: 3.545; 09: 4.2; 10: 5
| Std Dev. |  | 02: 0.1817; 03: 0.16733; 04: 0.17889; 05: 0.04535; 06: 0.11917; 07: 0.15748; 09: 0.075277; 10: 0.1169
| Rel.Std.Dev. |  | 02: 3.867%; 03: 2.699%; 04: 3.252%; 05: 1.059%; 06: 1.834%; 07: 4.512%; 09: 1.785%; 10: 2.346%
| PDM3 |  | 02: -5.578%; 03: 24.6%; 04: 10.53%; 05: -13.95%; 06: 30.57%; 07: -29.86%; 09: -15.26%; 10: 0.1501%
| Z-Score (Absolute) |  | 02: 0.15; 03: 1.37; 04: 0.69; 05: 0.59; 06: 1.74; 07: 1.42; 09: 0.66; 10: 0.15
| Analytical results for Sn in OREAS 922 (Certified Value 3.83 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
2
| 1 | 1 | 2.3 | 4 | 3.96 | 4.13 | 3.7 | 4 | < 5 | 4.41 | 4.4825 | 2.4 | < 10 | 3.545 | < 10 | 3.13 | 3.9 | 3.3 | 3.5 |  |  |  | 1
| 1 | 2 | 2.3 | 4.11 | 3.91 | 3.88 | 4.1 | 3.9 | < 5 | 4.17 | 4.4047 | 2.35 | < 10 | 3.618 | < 10 | 3.16 | 3.8 | 3.4 | 3.5 |  |  |  | #N/A
| 1 | 3 | 2.3 | 4.08 | 3.9 | 4.16 | 3.8 | 3.9 | < 5 | 4.33 | 4.2839 | 2.59 | < 10 | 3.669 | < 10 | 3.17 | 3.6 | 3.5 | 3.4 |  |  |  | 16
| 1 | 4 | 2.3 | 4.05 | 4.1 | 4.06 | 3.8 | 3.7 | < 5 | 4.41 | 4.2338 | 2.48 | < 10 | 3.609 | < 10 | 3.16 | 3.9 | 3.5 | 3.3 |  |  |  | 3.8314
| 1 | 5 | 2.2 | 3.89 | 4.19 | 4.15 | 3.9 | 3.6 | < 5 | 4.4 | 4.4625 | 2.46 | < 10 | 3.694 | < 10 | 3.15 | 3.9 | 3.3 | 3.5 |  |  |  | 114
| 1 | 6 | 2.2 | 3.92 | 4.34 | 4.27 | 3.8 | 3.5 | < 5 | 4.4 | 4.3928 | 2.47 | < 10 | 3.671 | < 10 | 3.11 | 3.8 | 3.4 | 3.3
| Mean |  | 2.2667 | 4.0083 | 4.0667 | 4.1083 | 3.85 | 3.7667 |  | 4.3533 | 4.3767 | 2.4583 |  | 3.6343 |  | 3.1467 | 3.8167 | 3.4 | 3.4167
| Median |  | 2.3 | 4.025 | 4.03 | 4.14 | 3.8 | 3.8 |  | 4.4 | 4.3987 | 2.465 |  | 3.6435 |  | 3.155 | 3.85 | 3.4 | 3.45
| Std Dev. |  | 0.05164 | 0.088412 | 0.17592 | 0.13075 | 0.13784 | 0.19664 |  | 0.094798 | 0.098619 | 0.081343 |  | 0.054764 |  | 0.022509 | 0.1169 | 0.089443 | 0.098319
| Rel.Std.Dev. |  | 2.278% | 2.206% | 4.326% | 3.183% | 3.58% | 5.22% |  | 2.178% | 2.253% | 3.309% |  | 1.507% |  | 0.7153% | 3.063% | 2.631% | 2.878%
| PDM3 |  | -40.84% | 4.618% | 6.141% | 7.228% | 0.4857% | -1.689% |  | 13.62% | 14.23% | -35.84% |  | -5.143% |  | -17.87% | -0.3843% | -11.26% | -10.82%
| Z-Score (Absolute) |  | 2.77 | 0.41 | 0.52 | 0.68 | 0.03 | 0.03 | NA | 1.11 | 1.09 | 2.42 | NA | 0.27 | NA | 1.16 | 0.06 | 0.7 | 0.67
| Analytical results for Sr in OREAS 922 (Certified Value 15 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
2
| 1 | 1 | 14.6 | 14.86 | 16 | 15 | 14.8 | 15.2 | 15 | 18 | 14.611 | 10.67 | 22 | 16.248 | 14 | 14 | 15.3 | 14.4 | 15.2 |  |  |  | 1
| 1 | 2 | 14.2 | 15.64 | 16 | 14.6 | 15.5 | 14.9 | 15 | 18 | 14.9245 | 10.64 | 22 | 15.905 | 14 | 15 | 15.1 | 14.5 | 15.6 |  |  |  | #N/A
| 1 | 3 | 14.4 | 15.24 | 15 | 14.8 | 15.2 | 15.2 | 15 | 18 | 14.7155 | 11.49 | 21 | 16.614 | 14 | 15 | 15 | 14.6 | 15.7 |  |  |  | 16
| 1 | 4 | 13.8 | 14.76 | 15 | 15.4 | 15.1 | 14.5 | 15 | 19 | 14.725 | 11.05 | 21 | 16.374 | 15 | 14 | 14.9 | 14.4 | 15.8 |  |  |  | 14.9972
| 1 | 5 | 13.5 | 14.72 | 16 | 16.1 | 14.8 | 14.1 | 15 | 19 | 14.934 | 11.1 | 21 | 16.23 | 15 | 15 | 15.5 | 14.5 | 15.6 |  |  |  | 115
| 1 | 6 | 13.3 | 14.54 | 16 | 16.2 | 14.5 | 14.2 | 15 | 19 | 14.877 | 11.12 | 21 | 16.449 | 15 | 14 | 15.2 | 14.5 | 15.7
| Mean |  | 13.9667 | 14.96 | 15.6667 | 15.35 | 14.9833 | 14.6833 | 15 | 18.5 | 14.7978 | 11.0117 | 21.3333 | 16.3033 | 14.5 | 14.5 | 15.1667 | 14.4833 | 15.6
| Median |  | 14 | 14.81 | 16 | 15.2 | 14.95 | 14.7 | 15 | 18.5 | 14.801 | 11.075 | 21 | 16.311 | 14.5 | 14.5 | 15.15 | 14.5 | 15.65
| Std Dev. |  | 0.5164 | 0.40615 | 0.5164 | 0.67454 | 0.35449 | 0.48751 | 0 | 0.54772 | 0.13255 | 0.31771 | 0.5164 | 0.24076 | 0.54772 | 0.54772 | 0.21602 | 0.075277 | 0.20976
| Rel.Std.Dev. |  | 3.697% | 2.715% | 3.296% | 4.394% | 2.366% | 3.32% | 0% | 2.961% | 0.8957% | 2.885% | 2.421% | 1.477% | 3.777% | 3.777% | 1.424% | 0.5198% | 1.345%
| PDM3 |  | -6.872% | -0.2482% | 4.464% | 2.352% | -0.09264% | -2.093% | 0.0185% | 23.36% | -1.33% | -26.58% | 42.25% | 8.709% | -3.315% | -3.315% | 1.13% | -3.427% | 4.019%
| Z-Score (Absolute) |  | 1.42 | 0.03 | 0.95 | 0.51 | 0 | 0.42 | 0.02 | 4.91 | 0.26 | 5.54 | 8.86 | 1.84 | 0.67 | 0.67 | 0.26 | 0.7 | 0.86
| Analytical results for Ta in OREAS 922 (Certified Value < 0.05 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
3
| 1 | 1 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01 | [6 empty] | 1
| 1 | 2 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01 | [6 empty] | #N/A
| 1 | 3 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01 | [6 empty] | 16
| 1 | 4 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01 | [6 empty] | < 0.05
| 1 | 5 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01 | [6 empty] | 116
| 1 | 6 | 0.06 | < 0.01 | < 0.05 | < 0.05 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.01
| Mean |  | 02: 0.06; 06: 0.01; 07: 0.01; 22: 0.01
| Median |  | 02: 0.06; 06: 0.01; 07: 0.01; 22: 0.01
| Std Dev. |  | 02: 0; 06: 0; 07: 0; 22: 0
Rel.Std.Dev. | bar: 02; 06 to 07; 22
PDM3
Z-Score (Absolute) | bar: 02 to 22
| Analytical results for Tb in OREAS 922 (Certified Value 0.62 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.59 | 0.566 | 0.7 | 0.8 | 0.465 | 0.80426 | 0.311 | 0.57 | 0.5 | 0.6 | 0.63 | [9 empty] | 1
| 1 | 2 | 0.59 | 0.586 | 0.7 | 0.8 | 0.465 | 0.81137 | 0.283 | 0.582 | 0.5 | 0.5 | 0.63 | [9 empty] | #N/A
| 1 | 3 | 0.58 | 0.588 | 0.8 | 0.8 | 0.455 | 0.78378 | 0.35 | 0.591 | 0.5 | 0.5 | 0.62 | [9 empty] | 16
| 1 | 4 | 0.59 | 0.576 | 0.7 | 0.8 | 0.46 | 0.79273 | 0.319 | 0.589 | 0.5 | 0.5 | 0.63 | [9 empty] | 0.61837
| 1 | 5 | 0.58 | 0.594 | 0.8 | 0.8 | 0.085 | 0.80014 | 0.338 | 0.58 | 0.5 | 0.5 | 0.63 | [9 empty] | 117
| 1 | 6 | 0.57 | 0.539 | 0.8 | 0.8 | 0.48 | 0.82069 | 0.326 | 0.578 | 0.5 | 0.5 | 0.62
| Mean |  | 0.58333 | 0.57483 | 0.75 | 0.8 | 0.40167 | 0.80216 | 0.32117 | 0.58167 | 0.5 | 0.51667 | 0.62667
| Median |  | 0.585 | 0.581 | 0.75 | 0.8 | 0.4625 | 0.8022 | 0.3225 | 0.581 | 0.5 | 0.5 | 0.63
| Std Dev. |  | 0.008165 | 0.020144 | 0.054772 | 1.2162e-16 | 0.15536 | 0.01314 | 0.02325 | 0.0076594 | 0 | 0.040825 | 0.005164
| Rel.Std.Dev. |  | 1.4% | 3.504% | 7.303% | 1.52e-14% | 38.68% | 1.638% | 7.239% | 1.317% | 0% | 7.902% | 0.824%
| PDM3 |  | -5.665% | -7.04% | 21.29% | 29.37% | -35.04% | 29.72% | -48.06% | -5.935% | -19.14% | -16.45% | 1.342%
| Z-Score (Absolute) |  | 0.01 | 0.06 | 1.39 | 1.8 | 0.96 | 1.82 | 2.15 | 0 | 0.67 | 0.67 | 0.37
| Analytical results for Te in OREAS 922 (Certified Value < 0.03 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.01 | < 0.02 | < 0.02 | 0.02 | 0.03 | < 0.1 | 0.030086 | < 0.01 | < 5 | < 0.05 | < 0.02 | < 1 | < 0.05 | 0.03 | [5 empty] | 1
| 1 | 2 | < 0.05 | 0.01 | < 0.02 | 0.03 | 0.01 | 0.02 | < 0.1 | 0.037144 | 0.02 | < 5 | 0.09 | < 0.02 | < 1 | < 0.05 | 0.04 | [5 empty] | 8
| 1 | 3 | < 0.05 | 0.01 | < 0.02 | 0.03 | 0.03 | 0.02 | < 0.1 | 0.026121 | < 0.01 | < 5 | < 0.05 | < 0.02 | < 1 | < 0.05 | 0.04 | [5 empty] | 16
| 1 | 4 | < 0.05 | 0.01 | < 0.02 | 0.03 | 0.02 | 0.01 | < 0.1 | 0.038042 | 0.01 | < 5 | < 0.05 | < 0.02 | < 1 | < 0.05 | < 0.01 | [5 empty] | < 0.03
| 1 | 5 | < 0.05 | < 0.01 | < 0.02 | 0.02 | 0.02 | 0.02 | < 0.1 | 0.028063 | 0.01 | < 5 | < 0.05 | < 0.02 | < 1 | < 0.05 | 0.03 | [5 empty] | 118
| 1 | 6 | < 0.05 | < 0.01 | < 0.02 | < 0.02 | 0.02 | 0.03 | < 0.1 | 0.026416 | < 0.01 | < 5 | < 0.05 | < 0.02 | < 1 | < 0.05 | 0.02
| Mean |  | 03: 0.01; 05: 0.0275; 06: 0.02; 07: 0.021667; 13: 0.030979; 14: 0.013333; 16: 0.09; 22: 0.032
| Median |  | 03: 0.01; 05: 0.03; 06: 0.02; 07: 0.02; 13: 0.029074; 14: 0.01; 16: 0.09; 22: 0.03
| Std Dev. |  | 03: 0; 05: 0.005; 06: 0.0063246; 07: 0.0075277; 13: 0.0053218; 14: 0.0057735; 22: 0.0083666
| Rel.Std.Dev. |  | 03: 0%; 05: 18.18%; 06: 31.62%; 07: 34.74%; 13: 17.18%; 14: 43.3%; 22: 26.15%
PDM3
| Z-Score (Absolute) |  | NA | 2.08 | NA | 0 | 0.25 | 0 | NA | 1.37 | 1.84 | NA | 0.49 | NA | NA | NA | 0.86
| Analytical results for Th in OREAS 922 (Certified Value 14.5 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 1
2
| 1 | 1 | 12.4 | 14.33 | 15.9 | 17 | 14.9 | 14.5 | 13.6 | 14 | 15.0451 | 10.73 | 15.326 | 13 | 15.56 | 14.2 | 14.4 | [5 empty] | 1
| 1 | 2 | 12.4 | 14.75 | 15.9 | 16.1 | 15.6 | 14 | 13 | 14 | 15.0517 | 10.24 | 15.196 | 13.2 | 15.52 | 14.1 | 14.5 | [5 empty] | #N/A
| 1 | 3 | 12.3 | 15.27 | 16.4 | 16.1 | 15.2 | 14.4 | 13.1 | 14 | 14.418 | 11.19 | 15.705 | 13.4 | 15.44 | 13.9 | 14.9 | [5 empty] | 16
| 1 | 4 | 12.4 | 14.73 | 15.8 | 15.8 | 15.4 | 14 | 13.5 | 15 | 14.5872 | 10.93 | 15.413 | 13.3 | 15.21 | 14.1 | 14.6 | [5 empty] | 14.5424
| 1 | 5 | 12.3 | 14.92 | 16.4 | 15.8 | 15 | 13.6 | 14.7 | 14 | 14.63 | 11.09 | 15.285 | 13 | 15.44 | 14.2 | 14.1 | [5 empty] | 119
| 1 | 6 | 12.2 | 14.5 | 17 | 16.1 | 14.7 | 13.5 | 14 | 15 | 14.9588 | 10.97 | 15.417 | 13 | 15.28 | 14 | 13.7
| Mean |  | 12.3333 | 14.75 | 16.2333 | 16.15 | 15.1333 | 14 | 13.65 | 14.3333 | 14.7818 | 10.8583 | 15.3903 | 13.15 | 15.4083 | 14.0833 | 14.3667
| Median |  | 12.35 | 14.74 | 16.15 | 16.1 | 15.1 | 14 | 13.55 | 14 | 14.7944 | 10.95 | 15.3695 | 13.1 | 15.44 | 14.1 | 14.45
| Std Dev. |  | 0.08165 | 0.32821 | 0.45898 | 0.44159 | 0.33267 | 0.40497 | 0.62849 | 0.5164 | 0.27084 | 0.34061 | 0.17507 | 0.17607 | 0.13659 | 0.1169 | 0.41793
| Rel.Std.Dev. |  | 0.662% | 2.225% | 2.827% | 2.734% | 2.198% | 2.893% | 4.604% | 3.603% | 1.832% | 3.137% | 1.138% | 1.339% | 0.8865% | 0.8301% | 2.909%
| PDM3 |  | -15.19% | 1.427% | 11.63% | 11.05% | 4.063% | -3.73% | -6.137% | -1.438% | 1.646% | -25.33% | 5.831% | -9.575% | 5.954% | -3.157% | -1.209%
| Z-Score (Absolute) |  | 1.79 | 0.34 | 1.64 | 1.42 | 0.67 | 0.32 | 0.63 | 0.03 | 0.37 | 2.98 | 0.9 | 1.07 | 0.92 | 0.25 | 0
| Analytical results for Ti in OREAS 922 (Certified Value < 0.15 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 3
3
| 1 | 1 | 0.09 | 0.1361 | 0.18 | 0.19 | 0.111 | 0.097 | 0.065 | 0.073 | 0.2391 | 0.16 | 0.116 | 0.1333 | 0.1 | 0.081 | 0.1 | 0.1 |  |  |  |  | 1
| 1 | 2 | 0.09 | 0.1418 | 0.18 | 0.19 | 0.119 | 0.095 | 0.065 | 0.074 | 0.2401 | 0.15 | 0.112 | 0.1287 | 0.1 | 0.08 | 0.1 | 0.1 |  |  |  |  | #N/A
| 1 | 3 | 0.09 | 0.1342 | 0.17 | 0.2 | 0.115 | 0.096 | 0.065 | 0.078 | 0.2362 | 0.15 | 0.118 | 0.1385 | 0.1 | 0.077 | 0.1 | 0.098 |  |  |  |  | 16
| 1 | 4 | 0.09 | 0.1286 | 0.17 | 0.19 | 0.114 | 0.097 | 0.065 | 0.078 | 0.2352 | 0.15 | 0.117 | 0.1352 | 0.1 | 0.094 | 0.1 | 0.099 |  |  |  |  | < 0.15
| 1 | 5 | 0.09 | 0.1365 | 0.18 | 0.2 | 0.107 | 0.094 | 0.065 | 0.076 | 0.2429 | 0.14 | 0.11 | 0.1306 | 0.1 | 0.079 | 0.1 | 0.103 |  |  |  |  | 120
| 1 | 6 | 0.08 | 0.1158 | 0.19 | 0.18 | 0.104 | 0.086 | 0.065 | 0.076 | 0.2408 | 0.14 | 0.118 | 0.1323 | 0.1 | 0.08 | 0.1 | 0.096
| Mean |  | 0.088333 | 0.13217 | 0.17833 | 0.19167 | 0.11167 | 0.094167 | 0.065 | 0.075833 | 0.23905 | 0.14833 | 0.11517 | 0.1331 | 0.1 | 0.081833 | 0.1 | 0.099333
| Median |  | 0.09 | 0.13515 | 0.18 | 0.19 | 0.1125 | 0.0955 | 0.065 | 0.076 | 0.2396 | 0.15 | 0.1165 | 0.1328 | 0.1 | 0.08 | 0.1 | 0.0995
| Std Dev. |  | 0.0040825 | 0.0090749 | 0.0075277 | 0.0075277 | 0.0055015 | 0.0041673 | 0 | 0.0020412 | 0.002896 | 0.0075277 | 0.0033714 | 0.0034601 | 1.5202e-17 | 0.0061128 | 1.5202e-17 | 0.0023381
| Rel.Std.Dev. |  | 4.622% | 6.866% | 4.221% | 3.928% | 4.927% | 4.425% | 0% | 2.692% | 1.211% | 5.075% | 2.927% | 2.6% | 1.52e-14% | 7.47% | 1.52e-14% | 2.354%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.74 | 1.82 | 2.16 | 0.15 | 0.25 | 1.03 | 0.75 | 3.35 | 1.07 | 0.23 | 0.68 | 0.15 | 0.66 | 0.15 | 0.16
| Analytical results for Tl in OREAS 922 (Certified Value 0.14 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
2
| 1 | 1 | 0.05 | 0.16 | 0.18 | 0.22 | 0.14 | 0.15 | 0.14 | < 5 | 0.1928 | 0.09 | < 10 | 0.126 | 0.09 | 0.16 | 0.14 | 0.15 |  |  |  |  | 1
| 1 | 2 | 0.05 | 0.16 | 0.17 | 0.21 | 0.15 | 0.15 | 0.13 | < 5 | 0.20171 | 0.09 | < 10 | 0.134 | 0.09 | 0.16 | 0.14 | 0.14 |  |  |  |  | #N/A
| 1 | 3 | 0.04 | 0.15 | 0.17 | 0.22 | 0.14 | 0.15 | 0.13 | < 5 | 0.20012 | 0.1 | < 10 | 0.134 | 0.09 | 0.15 | 0.14 | 0.14 |  |  |  |  | 16
| 1 | 4 | 0.05 | 0.15 | 0.18 | 0.22 | 0.14 | 0.15 | 0.14 | < 5 | 0.19553 | 0.1 | < 10 | 0.131 | 0.1 | 0.15 | 0.14 | 0.16 |  |  |  |  | 0.14227
| 1 | 5 | 0.04 | 0.15 | 0.19 | 0.22 | 0.13 | 0.14 | 0.15 | < 5 | 0.20031 | 0.1 | < 10 | 0.13 | 0.09 | 0.15 | 0.14 | 0.14 |  |  |  |  | 121
| 1 | 6 | 0.04 | 0.14 | 0.18 | 0.22 | 0.13 | 0.14 | 0.15 | < 5 | 0.19898 | 0.1 | < 10 | 0.131 | 0.09 | 0.17 | 0.14 | 0.13
| Mean |  | 0.045 | 0.15167 | 0.17833 | 0.21833 | 0.13833 | 0.14667 | 0.14 |  | 0.19824 | 0.096667 |  | 0.131 | 0.091667 | 0.15667 | 0.14 | 0.14333
| Median |  | 0.045 | 0.15 | 0.18 | 0.22 | 0.14 | 0.15 | 0.14 |  | 0.19955 | 0.1 |  | 0.131 | 0.09 | 0.155 | 0.14 | 0.14
| Std Dev. |  | 0.0054772 | 0.0075277 | 0.0075277 | 0.0040825 | 0.0075277 | 0.005164 | 0.0089443 |  | 0.003388 | 0.005164 |  | 0.0029665 | 0.0040825 | 0.008165 | 0 | 0.010328
| Rel.Std.Dev. |  | 12.17% | 4.963% | 4.221% | 1.87% | 5.442% | 3.521% | 6.389% |  | 1.709% | 5.342% |  | 2.264% | 4.454% | 5.212% | 0% | 7.206%
| PDM3 |  | -68.37% | 6.607% | 25.35% | 53.47% | -2.765% | 3.093% | -1.593% |  | 39.34% | -32.05% |  | -7.919% | -35.57% | 10.12% | -1.593% | 0.7498%
| Z-Score (Absolute) |  | 5.08 | 0.53 | 1.93 | 4.12 | 0.18 | 0.26 | 0.09 | NA | 2.97 | 2.36 | NA | 0.56 | 2.71 | 0.79 | 0.09 | 0.09
| Analytical results for Tm in OREAS 922 (Certified Value < 0.4 ppm) | [21 empty] | Certified
Tm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 13: Lab | [5 empty] | 1
| No | No | 02: 03; 03: 04; 04: 05; 05: 09; 06: 13; 07: 14; 09: 18; 13: 19 | [5 empty] | ppm
|  |  | 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 09: AR*MS; 13: AR*OES/MS | [5 empty] | 2
2
| 1 | 1 | 02: 0.23; 03: 0.3; 04: 0.3; 05: 0.16; 06: 0.33159; 07: 0.11; 09: 0.2; 13: 0.2 | [5 empty] | 1
| 1 | 2 | 02: 0.25; 03: 0.3; 04: 0.3; 05: 0.16; 06: 0.32995; 07: 0.1; 09: 0.2; 13: 0.2 | [5 empty] | #N/A
| 1 | 3 | 02: 0.24; 03: 0.3; 04: 0.3; 05: 0.165; 06: 0.33778; 07: 0.12; 09: 0.2; 13: 0.2 | [5 empty] | 16
| 1 | 4 | 02: 0.23; 03: 0.3; 04: 0.3; 05: 0.175; 06: 0.32707; 07: 0.11; 09: 0.2; 13: 0.2 | [5 empty] | < 0.4
| 1 | 5 | 02: 0.24; 03: 0.3; 04: 0.3; 05: 0.175; 06: 0.32895; 07: 0.11; 09: 0.2; 13: 0.2 | [5 empty] | 122
| 1 | 6 | 02: 0.21; 03: 0.3; 04: 0.3; 05: 0.18; 06: 0.34868; 07: 0.11; 09: 0.2; 13: 0.2
| Mean |  | 02: 0.23333; 03: 0.3; 04: 0.3; 05: 0.16917; 06: 0.33401; 07: 0.11; 09: 0.2; 13: 0.2
| Median |  | 02: 0.235; 03: 0.3; 04: 0.3; 05: 0.17; 06: 0.33077; 07: 0.11; 09: 0.2; 13: 0.2
| Std Dev. |  | 02: 0.013663; 03: 0; 04: 0; 05: 0.008612; 06: 0.0080694; 07: 0.0063246; 09: 3.0405e-17; 13: 3.0405e-17
| Rel.Std.Dev. |  | 02: 5.855%; 03: 0%; 04: 0%; 05: 5.091%; 06: 2.416%; 07: 5.75%; 09: 1.52e-14%; 13: 1.52e-14%
PDM3
| Z-Score (Absolute) |  | 02: 0.17; 03: 0.86; 04: 0.86; 05: 0.49; 06: 1.18; 07: 1.1; 09: 0.17; 13: 0.17
| Analytical results for U in OREAS 922 (Certified Value 1.98 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | 22 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
2
| 1 | 1 | 2.16 | 1.93 | 2.3 | 2.6 | 1.86 | 1.88 | 1.76 | 2.2348 | 1.45 | 2.061 | 1.8 | 2.06 | 1.84 | 1.81 | [6 empty] | 1
| 1 | 2 | 2.14 | 2.03 | 2.4 | 2.5 | 1.98 | 1.83 | 1.72 | 2.3072 | 1.36 | 2.032 | 1.8 | 2.05 | 1.86 | 1.82 | [6 empty] | #N/A
| 1 | 3 | 2.12 | 2.05 | 2.4 | 2.4 | 1.91 | 1.88 | 1.7 | 2.1949 | 1.52 | 2.119 | 1.9 | 1.94 | 1.85 | 1.84 | [6 empty] | 16
| 1 | 4 | 2.15 | 2.02 | 2.4 | 2.4 | 1.93 | 1.79 | 1.74 | 2.2336 | 1.47 | 2.131 | 1.8 | 1.92 | 1.87 | 1.82 | [6 empty] | 1.9777
| 1 | 5 | 2.11 | 2.04 | 2.4 | 2.4 | 1.89 | 1.76 | 1.8 | 2.2298 | 1.56 | 2.071 | 1.8 | 1.95 | 1.88 | 1.76 | [6 empty] | 123
| 1 | 6 | 2.1 | 1.86 | 2.5 | 2.4 | 1.83 | 1.72 | 1.8 | 2.2694 | 1.48 | 2.071 | 1.8 | 2.04 | 1.85 | 1.67
| Mean |  | 2.13 | 1.9883 | 2.4 | 2.45 | 1.9 | 1.81 | 1.7533 | 2.2449 | 1.4733 | 2.0808 | 1.8167 | 1.9933 | 1.8583 | 1.7867
| Median |  | 2.13 | 2.025 | 2.4 | 2.4 | 1.9 | 1.81 | 1.75 | 2.2342 | 1.475 | 2.071 | 1.8 | 1.995 | 1.855 | 1.815
| Std Dev. |  | 0.023664 | 0.076267 | 0.063246 | 0.083666 | 0.052915 | 0.065115 | 0.041312 | 0.038558 | 0.06802 | 0.037269 | 0.040825 | 0.06314 | 0.01472 | 0.06314
| Rel.Std.Dev. |  | 1.111% | 3.836% | 2.635% | 3.415% | 2.785% | 3.598% | 2.356% | 1.718% | 4.617% | 1.791% | 2.247% | 3.168% | 0.7921% | 3.534%
| PDM3 |  | 7.7% | 0.5365% | 21.35% | 23.88% | -3.93% | -8.481% | -11.35% | 13.51% | -25.5% | 5.214% | -8.143% | 0.7893% | -6.037% | -9.66%
| Z-Score (Absolute) |  | 0.87 | 0.32 | 2.16 | 2.25 | 0.22 | 0.65 | 0.92 | 1.42 | 2.25 | 0.64 | 0.7 | 0.22 | 0.42 | 0.65
| Analytical results for V in OREAS 922 (Certified Value 29.4 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
2
| 1 | 1 | 34 | 30 | 34 | 35 | 28 | 29 | 28 | 28 | 35.1 | 33 | 27.066 | 30 | 30 | 28.9 | 26 | 28 |  |  |  |  | 1
| 1 | 2 | 34 | 30 | 35 | 34 | 30 | 29 | 28 | 28 | 36.9 | 32 | 26.721 | 30 | 30 | 28.1 | 27 | 28 |  |  |  |  | #N/A
| 1 | 3 | 34 | 30 | 33 | 36 | 29 | 30 | 28 | 28 | 36 | 32 | 26.816 | 30 | 29 | 28.3 | 27 | 28 |  |  |  |  | 16
| 1 | 4 | 33 | 29 | 32 | 36 | 29 | 29 | 28 | 29 | 36 | 32 | 27.367 | 31 | 30 | 28.1 | 27 | 27 |  |  |  |  | 29.3908
| 1 | 5 | 32 | 29 | 34 | 36 | 29 | 28 | 26 | 28 | 36.9 | 32 | 26.782 | 30 | 30 | 28.5 | 26 | 28 |  |  |  |  | 124
| 1 | 6 | 32 | 28 | 34 | 35 | 28 | 28 | 26 | 28 | 36.9 | 31 | 26.279 | 31 | 30 | 28.9 | 27 | 28
| Mean |  | 33.1667 | 29.3333 | 33.6667 | 35.3333 | 28.8333 | 28.8333 | 27.3333 | 28.1667 | 36.3 | 32 | 26.8385 | 30.3333 | 29.8333 | 28.4667 | 26.6667 | 27.8333
| Median |  | 33.5 | 29.5 | 34 | 35.5 | 29 | 29 | 28 | 28 | 36.45 | 32 | 26.799 | 30 | 30 | 28.4 | 27 | 28
| Std Dev. |  | 0.98319 | 0.8165 | 1.0328 | 0.8165 | 0.75277 | 0.75277 | 1.0328 | 0.40825 | 0.73485 | 0.63246 | 0.36376 | 0.5164 | 0.40825 | 0.36697 | 0.5164 | 0.40825
| Rel.Std.Dev. |  | 2.964% | 2.784% | 3.068% | 2.311% | 2.611% | 2.611% | 3.779% | 1.449% | 2.024% | 1.976% | 1.355% | 1.702% | 1.368% | 1.289% | 1.936% | 1.467%
| PDM3 |  | 12.85% | -0.1957% | 14.55% | 20.22% | -1.897% | -1.897% | -7.001% | -4.165% | 23.51% | 8.877% | -8.684% | 3.207% | 1.506% | -3.144% | -9.269% | -5.299%
| Z-Score (Absolute) |  | 1.84 | 0.11 | 2.06 | 2.81 | 0.11 | 0.11 | 0.79 | 0.49 | 3.24 | 1.31 | 1.01 | 0.56 | 0.41 | 0.28 | 1.09 | 0.49
| Analytical results for W in OREAS 922 (Certified Value 1.12 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
2
| 1 | 1 | 0.6 | 1.08 | 0.4 | 1.3 | 0.93 | 1.1 | 0.8 | < 2 | 1.3687 | 0.68 | < 5 | 0.936 | < 10 | 0.6 | 1 | 1.1 | 1.09 |  |  |  | 1
| 1 | 2 | 0.7 | 1.25 | 0.4 | 1.1 | 1.13 | 1.52 | 0.7 | < 2 | 1.3253 | 0.67 | < 5 | 1.054 | < 10 | 0.6 | 1 | 1.1 | 1.61 |  |  |  | #N/A
| 1 | 3 | 0.7 | 1.04 | 0.4 | 1.2 | 1.27 | 1.14 | 0.7 | < 2 | 1.305 | 0.55 | < 5 | 1.165 | < 10 | 0.6 | 1 | 1.1 | 1.29 |  |  |  | 16
| 1 | 4 | 0.7 | 1.08 | 0.5 | 1.1 | 0.97 | 1.26 | 0.7 | < 2 | 1.3071 | 0.53 | < 5 | 1.107 | < 10 | 0.7 | 1 | 1.2 | 1.29 |  |  |  | 1.1178
| 1 | 5 | 0.7 | 1.2 | 0.4 | 1.1 | 1.57 | 1.11 | 0.8 | < 2 | 1.4848 | 0.54 | < 5 | 1.075 | < 10 | 0.6 | 1 | 1.5 | 1.06 |  |  |  | 125
| 1 | 6 | 0.6 | 1.29 | 0.6 | 1.2 | 0.98 | 1.04 | 0.9 | < 2 | 1.2912 | 0.59 | < 5 | 1.364 | < 10 | 0.5 | 1 | 1.2 | 1.1
| Mean |  | 0.66667 | 1.1567 | 0.45 | 1.1667 | 1.1417 | 1.195 | 0.76667 |  | 1.347 | 0.59333 |  | 1.1168 |  | 0.6 | 1 | 1.2 | 1.24
| Median |  | 0.7 | 1.14 | 0.4 | 1.15 | 1.055 | 1.125 | 0.75 |  | 1.3162 | 0.57 |  | 1.091 |  | 0.6 | 1 | 1.15 | 1.195
| Std Dev. |  | 0.05164 | 0.10367 | 0.083666 | 0.08165 | 0.24515 | 0.17501 | 0.08165 |  | 0.072681 | 0.066533 |  | 0.14276 |  | 0.063246 | 0 | 0.15492 | 0.20804
| Rel.Std.Dev. |  | 7.746% | 8.962% | 18.59% | 6.999% | 21.47% | 14.65% | 10.65% |  | 5.396% | 11.21% |  | 12.78% |  | 10.54% | 0% | 12.91% | 16.78%
| PDM3 |  | -40.36% | 3.477% | -59.74% | 4.372% | 2.136% | 6.907% | -31.41% |  | 20.51% | -46.92% |  | -0.0861% |  | -46.32% | -10.54% | 7.354% | 10.93%
| Z-Score (Absolute) |  | 2.57 | 0.19 | 3.95 | 0.24 | 0.1 | 0.04 | 2 | NA | 1.1 | 2.98 | NA | 0.04 | NA | 2.94 | 0.69 | 0.09 | 0.66
| Analytical results for Y in OREAS 922 (Certified Value 16 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
1
| 1 | 1 | 11.4 | 18.51 | 20 | 20 | 14.9 | 14.9 | 12.5 | 20.2 | 7.58 | 19 | 16.636 | 17 | 14 | 14.5 | 14.9 | 13.4 |  |  |  |  | 1
| 1 | 2 | 11.3 | 19.64 | 20 | 21 | 15.9 | 14.35 | 12.4 | 20.53 | 7.37 | 18 | 17.242 | 16 | 14 | 14.5 | 15.2 | 13.95 |  |  |  |  | #N/A
| 1 | 3 | 11.3 | 18.86 | 19 | 21 | 15.4 | 14.55 | 12.5 | 19.92 | 8.82 | 18 | 17.769 | 17 | 14 | 14.2 | 15.3 | 14 |  |  |  |  | 16
| 1 | 4 | 11.2 | 17.88 | 18 | 21 | 15.3 | 14.2 | 12.5 | 20 | 8.37 | 18 | 17.367 | 17 | 14 | 14.1 | 15.4 | 13.9 |  |  |  |  | 16.0264
| 1 | 5 | 11.1 | 18.3 | 19 | 21 | 14.45 | 13.6 | 12.4 | 20.32 | 8.6 | 17 | 16.97 | 16 | 14 | 14 | 14.9 | 13.85 |  |  |  |  | 126
| 1 | 6 | 11.2 | 16.65 | 19 | 20 | 14.35 | 13.2 | 12.8 | 20.46 | 8.21 | 18 | 17.368 | 17 | 14 | 14.2 | 15.3 | 13.45
| Mean |  | 11.25 | 18.3067 | 19.1667 | 20.6667 | 15.05 | 14.1333 | 12.5167 | 20.2383 | 8.1583 | 18 | 17.2253 | 16.6667 | 14 | 14.25 | 15.1667 | 13.7583
| Median |  | 11.25 | 18.405 | 19 | 21 | 15.1 | 14.275 | 12.5 | 20.26 | 8.29 | 18 | 17.3045 | 17 | 14 | 14.2 | 15.25 | 13.875
| Std Dev. |  | 0.10488 | 1.0052 | 0.75277 | 0.5164 | 0.59666 | 0.62743 | 0.1472 | 0.24515 | 0.57206 | 0.63246 | 0.38698 | 0.5164 | 0 | 0.20736 | 0.21602 | 0.26347
| Rel.Std.Dev. |  | 0.9323% | 5.491% | 3.928% | 2.499% | 3.965% | 4.439% | 1.176% | 1.211% | 7.012% | 3.514% | 2.247% | 3.098% | 0% | 1.455% | 1.424% | 1.915%
| PDM3 |  | -29.8% | 14.23% | 19.59% | 28.95% | -6.092% | -11.81% | -21.9% | 26.28% | -49.09% | 12.31% | 7.481% | 3.995% | -12.64% | -11.08% | -5.364% | -14.15%
| Z-Score (Absolute) |  | 1.11 | 0.92 | 1.16 | 1.59 | 0.02 | 0.28 | 0.74 | 1.47 | 1.99 | 0.83 | 0.61 | 0.45 | 0.32 | 0.25 | 0.02 | 0.39
| Analytical results for Yb in OREAS 922 (Certified Value < 2 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 09 | 13 | 14 | 16 | 18 | 19 | 21 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.2 | 1.51 | 1.8 | 2.1 | 1.08 | 2.1494 | 0.68 | 1.405 | 1.2 | 0.5 | 1.3 | [9 empty] | 1
| 1 | 2 | 1.2 | 1.59 | 1.8 | 1.9 | 1.04 | 2.2701 | 0.59 | 1.416 | 1.2 | 0.5 | 1.3 | [9 empty] | #N/A
| 1 | 3 | 1.2 | 1.66 | 1.8 | 1.9 | 1.05 | 2.214 | 0.7 | 1.463 | 1.2 | 0.5 | 1.3 | [9 empty] | 16
| 1 | 4 | 1.2 | 1.51 | 1.8 | 1.9 | 1.04 | 2.2681 | 0.69 | 1.455 | 1.2 | 0.5 | 1.3 | [9 empty] | < 2
| 1 | 5 | 1.2 | 1.58 | 1.9 | 2 | 1.1 | 2.247 | 0.77 | 1.405 | 1.2 | 0.5 | 1.3 | [9 empty] | 127
| 1 | 6 | 1.2 | 1.33 | 1.9 | 2.1 | 1.08 | 2.3223 | 0.7 | 1.431 | 1.2 | 0.5 | 1.3
| Mean |  | 1.2 | 1.53 | 1.8333 | 1.9833 | 1.065 | 2.2452 | 0.68833 | 1.4292 | 1.2 | 0.5 | 1.3
| Median |  | 1.2 | 1.545 | 1.8 | 1.95 | 1.065 | 2.2575 | 0.695 | 1.4235 | 1.2 | 0.5 | 1.3
| Std Dev. |  | 0 | 0.11296 | 0.05164 | 0.098319 | 0.0251 | 0.058721 | 0.057764 | 0.025127 | 0 | 0 | 0
| Rel.Std.Dev. |  | 0% | 7.383% | 2.817% | 4.957% | 2.357% | 2.615% | 8.392% | 1.758% | 0% | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 0.29 | 0.66 | 1.53 | 1.96 | 0.67 | 2.71 | 1.76 | 0.37 | 0.29 | 2.3 | 0
| Analytical results for Zn in OREAS 922 (Certified Value 256 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS |  | 0
0
| 1 | 1 | 260 | 242 | 271 | 260 | 262 | 271 | 302.3534 | 246 | 268 | 231 | 269.287 | 240 | 258.831 | 237 | 268 | 257.1 | 274 | 240 | 259 |  | 1
| 1 | 2 | 264 | 241 | 273 | 237 | 268 | 269 | 298.9459 | 244 | 269 | 230 | 273.6238 | 260 | 254.398 | 244 | 271 | 255.5 | 266 | 250 | 259 |  | #N/A
| 1 | 3 | 261 | 247 | 271 | 237 | 264 | 277 | 303.186 | 243 | 246 | 225 | 271.0366 | 240 | 258.22 | 243 | 261 | 251.6 | 276 | 252 | 260 |  | 16
| 1 | 4 | 257 | 245 | 264 | 249 | 268 | 271 | 294.7599 | 252 | 242 | 240 | 261.0334 | 240 | 253.236 | 253 | 266 | 254.3 | 275 | 259 | 253 |  | 256.0471
| 1 | 5 | 245 | 245 | 268 | 275 | 269 | 263 | 299.0076 | 248 | 261 | 234 | 274.498 | 240 | 244.374 | 244 | 268 | 259.1 | 280 | 255 | 257 |  | 128
| 1 | 6 | 248 | 242 | 274 | 270 | 258 | 257 | 295.7926 | 246 | 256 | 231 | 247 | 240 | 257.577 | 248 | 265 | 259.4 | 273 | 246 | 261
| Mean |  | 255.8333 | 243.6667 | 270.1667 | 254.6667 | 264.8333 | 268 | 299.0076 | 246.5 | 257 | 231.8333 | 266.0798 | 243.3333 | 254.4393 | 244.8333 | 266.5 | 256.1667 | 274 | 250.3333 | 258.1667
| Median |  | 258.5 | 243.5 | 271 | 254.5 | 266 | 270 | 298.9768 | 246 | 258.5 | 231 | 270.1618 | 240 | 255.9875 | 244 | 267 | 256.3 | 274.5 | 251 | 259
| Std Dev. |  | 7.6267 | 2.3381 | 3.656 | 16.3299 | 4.3089 | 7.0143 | 3.3774 | 3.2094 | 11.2071 | 4.9565 | 10.5085 | 8.165 | 5.4071 | 5.3448 | 3.3912 | 2.9904 | 4.6043 | 6.7132 | 2.8577
| Rel.Std.Dev. |  | 2.981% | 0.9595% | 1.353% | 6.412% | 1.627% | 2.617% | 1.13% | 1.302% | 4.361% | 2.138% | 3.949% | 3.355% | 2.125% | 2.183% | 1.272% | 1.167% | 1.68% | 2.682% | 1.107%
| PDM3 |  | -0.0835% | -4.835% | 5.514% | -0.5391% | 3.431% | 4.668% | 16.78% | -3.729% | 0.3722% | -9.457% | 3.918% | -4.965% | -0.6279% | -4.38% | 4.082% | 0.04669% | 7.012% | -2.232% | 0.8278%
| Z-Score (Absolute) |  | 0.04 | 0.86 | 0.92 | 0.12 | 0.56 | 0.77 | 2.86 | 0.67 | 0.04 | 1.65 | 0.9 | 1.1 | 0 | 0.78 | 0.67 | 0.02 | 1.18 | 0.41 | 0.12
| Analytical results for Zr in OREAS 922 (Certified Value 22.3 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 09 | 13 | 14 | 15 | 16 | 18 | 19 | 21 | 22 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | [5 empty] | 1
1
| 1 | 1 | 21.2 | 22.2 | 16 | 22.2 | 23.6 | 23.2 | 19 | 41.1325 | 6.5 | 26 | 21.106 | 10 | 21 | 22.8 | 22.7 | [5 empty] | 1
| 1 | 2 | 19.8 | 23.6 | 14 | 24.3 | 25.5 | 22.3 | 19 | 41.1472 | 5.8 | 27 | 19.884 | 11 | 21 | 23 | 23.3 | [5 empty] | 9
| 1 | 3 | 20.8 | 22.3 | 17 | 26.4 | 24.8 | 23 | 18 | 40.4968 | 2.4 | 27 | 19.369 | 10 | 20 | 23.3 | 23.9 | [5 empty] | 16
| 1 | 4 | 20 | 21.7 | 23 | 28.2 | 24.8 | 22.5 | 19 | 40.4598 | 2.7 | 28 | 20.179 | 9 | 21 | 24.1 | 21.1 | [5 empty] | 22.304
| 1 | 5 | 19.2 | 21.6 | 24 | 28.1 | 23.1 | 21.5 | 18 | 40.0127 | 5 | 28 | 21.121 | 10 | 20 | 22.6 | 23.5 | [5 empty] | 129
| 1 | 6 | 17.6 | 18.4 | 26 | 27.5 | 23.3 | 21.1 | 20 | 40.5229 | 2.4 | 28 | 20.35 | 10 | 20 | 23.2 | 22.7
| Mean |  | 19.7667 | 21.6333 | 20 | 26.1167 | 24.1833 | 22.2667 | 18.8333 | 40.6286 | 4.1333 | 27.3333 | 20.3348 | 10 | 20.5 | 23.1667 | 22.8667
| Median |  | 19.9 | 21.95 | 20 | 26.95 | 24.2 | 22.4 | 19 | 40.5098 | 3.85 | 27.5 | 20.2645 | 10 | 20.5 | 23.1 | 23
| Std Dev. |  | 1.2801 | 1.7374 | 4.9396 | 2.4045 | 0.9786 | 0.82624 | 0.75277 | 0.43801 | 1.8544 | 0.8165 | 0.68896 | 0.63246 | 0.54772 | 0.52409 | 0.98319
| Rel.Std.Dev. |  | 6.476% | 8.031% | 24.7% | 9.207% | 4.047% | 3.711% | 3.997% | 1.078% | 44.86% | 2.987% | 3.388% | 6.325% | 2.672% | 2.262% | 4.3%
| PDM3 |  | -11.38% | -3.007% | -10.33% | 17.09% | 8.426% | -0.1674% | -15.56% | 82.16% | -81.47% | 22.55% | -8.829% | -55.17% | -8.088% | 3.868% | 2.523%
| Z-Score (Absolute) |  | 0.74 | 0 | 0.67 | 1.15 | 0.57 | 0 | 1.02 | 5.46 | 5.39 | 1.51 | 0.57 | 3.65 | 0.53 | 0.27 | 0.18
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1169: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1187: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for S in OREAS 922 (Indicative Value 0.398 wt.%) |  |  |  |  | Indicative
S | Round | Replicate | Lab | Lab | Lab | 1
| No | No | 01 | 10 | 20 | wt.%
|  |  | IRC | IRC | IRC | 3
3
| 1 | 1 | NA | 0.38 | 0.396 | 1
| 1 | 2 | NA | 0.4 | 0.4 | 6
| 1 | 3 | NA | 0.41 | 0.391 | 16
| 1 | 4 | NA | 0.41 | 0.387 | 0.39783
| 1 | 5 | NA | 0.4 | 0.401 | 12
| 1 | 6 | NA | 0.4 | 0.399
| Mean |  |  | 0.4 | 0.39567
| Median |  |  | 0.4 | 0.3975
| Std Dev. |  |  | 0.010954 | 0.0055737
| Rel.Std.Dev. |  |  | 2.739% | 1.409%
| PDM3 |  |  | 0.5446% | -0.5446%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 922 (Indicative Value 2.83 ppm) | [15 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 09 | 10 | 16 | 21 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 7 | < 5 | < 10 | 1.208 | 1 | [9 empty] | 1
| 1 | 2 | 7 | < 5 | < 10 | < 1 | 1 | [9 empty] | 8
| 1 | 3 | 7 | < 5 | < 10 | < 1 | 2 | [9 empty] | 16
| 1 | 4 | 7 | < 5 | < 10 | < 1 | < 1 | [9 empty] | 2.8333
| 1 | 5 | 7 | < 5 | < 10 | < 1 | 1 | [9 empty] | 14
| 1 | 6 | 7 | < 5 | < 10 | < 1 | < 1
| Mean |  | 7 |  |  | 1.208 | 1.25
| Median |  | 7 |  |  | 1.208 | 1
| Std Dev. |  | 0 |  |  |  | 0.5
| Rel.Std.Dev. |  | 0% |  |  |  | 40%
| PDM3 |  | 147.1% |  |  | -57.36% | -55.88%
| Z-Score (Absolute) |  | 8.09 | NA | NA | 0.67 | 0
| Analytical results for Al in OREAS 922 (Certified Value 7.59 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 2
3
| 1 | 1 | 7.15 | 7.6 | 7.68 | 8.0179 | 6.39 | 7.427 | 8.9176 | 7.65 | 7.62 | 7.56 | 7.51 |  |  |  | 1
| 1 | 2 | 6.78 | 7.46 | 7.72 | 8.0913 | 6.75 | 7.301 | 8.1372 | 7.55 | 7.61 | 7.52 | 7.52 |  |  |  | #N/A
| 1 | 3 | 7.14 | 7.42 | 7.62 | 7.9711 | 6.84 | 7.315 | 8.3714 | 7.78 | 7.7 | 7.5 | 7.6 |  |  |  | 16
| 1 | 4 | 6.87 | 7.25 | 7.67 | 8.0441 | 6.78 | 7.233 | 8.1189 | 7.8 | 7.68 | 7.53 | 7.48 |  |  |  | 7.5925
| 1 | 5 | 6.84 | 7.33 | 7.58 | 8.1187 | 6.7 | 7.486 | 8.1793 | 7.31 | 7.61 | 7.15 | 7.63 |  |  |  | 131
| 1 | 6 | 6.73 | 7.3 | 7.56 | 8.1846 | 6.48 | 7.476 | 8.1661 | 7.54 | 7.58 | 7.26 | 7.57
| Mean |  | 6.9183 | 7.3933 | 7.6383 | 8.0713 | 6.6567 | 7.373 | 8.3151 | 7.605 | 7.6333 | 7.42 | 7.5517
| Median |  | 6.855 | 7.375 | 7.645 | 8.0677 | 6.725 | 7.371 | 8.1727 | 7.6 | 7.615 | 7.51 | 7.545
| Std Dev. |  | 0.18214 | 0.12738 | 0.062102 | 0.076318 | 0.17985 | 0.10435 | 0.30885 | 0.18163 | 0.046332 | 0.17123 | 0.057764
| Rel.Std.Dev. |  | 2.633% | 1.723% | 0.813% | 0.9456% | 2.702% | 1.415% | 3.714% | 2.388% | 0.607% | 2.308% | 0.7649%
| PDM3 |  | -8.879% | -2.623% | 0.6037% | 6.306% | -12.33% | -2.891% | 9.517% | 0.1647% | 0.5379% | -2.272% | -0.5377%
| Z-Score (Absolute) |  | 2.7 | 0.67 | 0.37 | 2.21 | 3.81 | 0.76 | 2.74 | 0.23 | 0.35 | 0.33 | 0
| Analytical results for As in OREAS 922 (Certified Value < 8 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 10 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*AAS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  | 2
2
| 1 | 1 | 11 | < 5 | < 5 | < 20 | < 20 | < 100 | 7.171 | < 1000 | < 5 | < 100 | 7 |  |  |  | 1
| 1 | 2 | 11 | < 5 | < 5 | < 20 | < 20 | < 100 | 7.675 | < 1000 | < 5 | < 100 | 6 |  |  |  | #N/A
| 1 | 3 | 12 | < 5 | < 5 | < 20 | < 20 | < 100 | 6.891 | < 1000 | < 5 | < 100 | 6 |  |  |  | 16
| 1 | 4 | 12 | < 5 | < 5 | < 20 | < 20 | < 100 | 7.615 | < 1000 | < 5 | < 100 | 6 |  |  |  | < 8
| 1 | 5 | 12 | < 5 | < 5 | < 20 | < 20 | < 100 | 7.201 | < 1000 | < 5 | < 100 | 7 |  |  |  | 132
| 1 | 6 | 11 | < 5 | < 5 | < 20 | < 20 | < 100 | 7.437 | < 1000 | < 5 | < 100 | 7
| Mean |  | 02: 11.5; 17: 7.3317; 21: 6.5
| Median |  | 02: 11.5; 17: 7.319; 21: 6.5
| Std Dev. |  | 02: 0.54772; 17: 0.29877; 21: 0.54772
| Rel.Std.Dev. |  | 02: 4.763%; 17: 4.075%; 21: 8.427%
PDM3
| Z-Score (Absolute) |  | 3.38 | NA | NA | NA | NA | NA | 0 | NA | NA | NA | 0.67
| Analytical results for B in OREAS 922 (Indicative Value 42.5 ppm) | [15 empty] | Indicative
B | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab; 15: Lab |  |  |  | 1
| No | No | 02: 04; 04: 05; 05: 09; 13: 13; 15: 18 |  |  |  | ppm
|  |  | 02: PF*MS; 04: PF*MS; 05: PF*OES; 13: PF*OES; 15: PF*MS |  |  |  | 1
1
| 1 | 1 | 02: 40; 04: 50; 05: 40; 13: < 50; 15: 40 |  |  |  | 1
| 1 | 2 | 02: 40; 04: 50; 05: 40; 13: < 50; 15: 50 |  |  |  | 1
| 1 | 3 | 02: 40; 04: 50; 05: 40; 13: < 50; 15: 50 |  |  |  | 16
| 1 | 4 | 02: 40; 04: 50; 05: 20; 13: < 50; 15: 40 |  |  |  | 42.5
| 1 | 5 | 02: 40; 04: 50; 05: 40; 13: < 50; 15: 50 |  |  |  | 15
| 1 | 6 | 02: 40; 04: 50; 05: 20; 13: < 50; 15: 50
| Mean |  | 02: 40; 04: 50; 05: 33.3333; 15: 46.6667
| Median |  | 02: 40; 04: 50; 05: 40; 15: 50
| Std Dev. |  | 02: 0; 04: 0; 05: 10.328; 15: 5.164
| Rel.Std.Dev. |  | 02: 0%; 04: 0%; 05: 30.98%; 15: 11.07%
| PDM3 |  | 02: -5.882%; 04: 17.65%; 05: -21.57%; 15: 9.804%
| Z-Score (Absolute) |  | 02: 0.45; 04: 0.9; 05: 1.35; 13: NA; 15: 0.45
| Analytical results for Ba in OREAS 922 (Certified Value 481 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES/MS |  |  |  |  | 0
0
| 1 | 1 | 441 | 494 | 502 | 480 | 507.8647 | 471.89 | 473 | 497.8 | 450 | 470 |  |  |  |  | 1
| 1 | 2 | 426 | 491 | 474 | 490 | 505.692 | 473.73 | 477 | 498 | 442 | 460 |  |  |  |  | #N/A
| 1 | 3 | 457 | 498 | 486 | 480 | 505.1651 | 467.22 | 480 | 468.7 | 446 | 480 |  |  |  |  | 16
| 1 | 4 | 441 | 495 | 482 | 490 | 499.1297 | 466.71 | 644 | 486.7 | 450 | 470 |  |  |  |  | 480.8062
| 1 | 5 | 427 | 498 | 466 | 490 | 509.5841 | 477.52 | 493 | 464.8 | 420 | 500 |  |  |  |  | 133
| 1 | 6 | 434 | 490 | 480 | 470 | 501.9645 | 461.99 | 467 | 464.8 | 438 | 490
| Mean |  | 437.6667 | 494.3333 | 481.6667 | 483.3333 | 504.9 | 469.8433 | 505.6667 | 480.1333 | 441 | 478.3333
| Median |  | 437.5 | 494.5 | 481 | 485 | 505.4286 | 469.555 | 478.5 | 477.7 | 444 | 475
| Std Dev. |  | 11.4833 | 3.3862 | 12.16 | 8.165 | 3.8299 | 5.5964 | 68.3218 | 15.9712 | 11.296 | 14.7196
| Rel.Std.Dev. |  | 2.624% | 0.685% | 2.525% | 1.689% | 0.7586% | 1.191% | 13.51% | 3.326% | 2.561% | 3.077%
| PDM3 |  | -8.972% | 2.813% | 0.179% | 0.5256% | 5.011% | -2.28% | 5.171% | -0.1399% | -8.279% | -0.5143%
| Z-Score (Absolute) |  | 4.16 | 1.51 | 0.24 | 0.41 | 2.57 | 0.94 | 0.12 | 0.09 | 3.4 | 0.09
| Analytical results for Be in OREAS 922 (Certified Value < 3 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | < 5 | < 3 | < 3 | 2 | 2.7191 | < 5 | < 3 | 2.8 | < 10 | < 5 |  |  |  |  | 1
| 1 | 2 | < 5 | < 3 | < 3 | 2 | 2.7151 | < 5 | < 3 | 2.2 | < 10 | < 5 |  |  |  |  | #N/A
| 1 | 3 | < 5 | < 3 | < 3 | 2 | 2.4326 | < 5 | 4 | 2.6 | < 10 | < 5 |  |  |  |  | 16
| 1 | 4 | < 5 | < 3 | < 3 | 2 | 2.468 | < 5 | 4 | 2.6 | < 10 | < 5 |  |  |  |  | < 3
| 1 | 5 | < 5 | < 3 | < 3 | 2 | 2.5241 | < 5 | 4 | 2.3 | < 10 | < 5 |  |  |  |  | 134
| 1 | 6 | < 5 | < 3 | < 3 | 3 | 2.3394 | < 5 | 4 | 2 | < 10 | < 5
| Mean |  | 13: 2.1667; 15: 2.5331; 17: 4; 18: 2.4167
| Median |  | 13: 2; 15: 2.4961; 17: 4; 18: 2.45
| Std Dev. |  | 13: 0.40825; 15: 0.15469; 17: 0; 18: 0.29944
| Rel.Std.Dev. |  | 13: 18.84%; 15: 6.107%; 17: 0%; 18: 12.39%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | 1.2 | 0.15 | NA | 1.75 | 0.15 | NA | NA
| Analytical results for Bi in OREAS 922 (Certified Value 10.8 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 16 | 18 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [5 empty] | 1
2
| 1 | 1 | 12.1 | 11 | 8 | 10 | 11 | 10.9091 | 10.67 | 12 | 11.2 | [5 empty] | 1
| 1 | 2 | 12.4 | 7 | 8 | 10 | 10.1 | 10.0253 | 10.561 | 10 | 11.4 | [5 empty] | #N/A
| 1 | 3 | 12.7 | 10 | 8 | 10 | 12.6 | 10.6048 | 10.176 | 11 | 10.7 | [5 empty] | 16
| 1 | 4 | 12.5 | 10 | 7 | 10 | 9.3 | 10.0598 | 12.385 | 14 | 10.7 | [5 empty] | 10.7943
| 1 | 5 | 12.1 | 8 | 11 | 10 | 10 | 11.8937 | 10.228 | 11 | 11 | [5 empty] | 135
| 1 | 6 | 12.5 | 10 | 9 | 10 | 11.4 | 9.2606 | 11.851 | 11 | 9
| Mean |  | 12.3833 | 9.3333 | 8.5 | 10 | 10.7333 | 10.4589 | 10.9785 | 11.5 | 10.6667
| Median |  | 12.45 | 10 | 8 | 10 | 10.55 | 10.3323 | 10.6155 | 11 | 10.85
| Std Dev. |  | 0.24014 | 1.5055 | 1.3784 | 0 | 1.1827 | 0.90174 | 0.91828 | 1.3784 | 0.86178
| Rel.Std.Dev. |  | 1.939% | 16.13% | 16.22% | 0% | 11.02% | 8.622% | 8.364% | 11.99% | 8.079%
| PDM3 |  | 14.72% | -13.53% | -21.25% | -7.358% | -0.5644% | -3.107% | 1.707% | 6.538% | -1.182%
| Z-Score (Absolute) |  | 4.05 | 2.29 | 6.72 | 1.8 | 0 | 0.67 | 0.6 | 0.66 | 0.66
| Analytical results for Ca in OREAS 922 (Certified Value 0.486 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 18 | 19 | 20 | 21 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.42 | 0.49 | 0.73 | 0.5 | 0.51553 | 0.77 | 0.524 | 0.67 | 0.47 | 0.503 | 0.5 |  |  |  | 1
| 1 | 2 | 0.41 | 0.46 | 0.46 | 0.6 | 0.53308 | 0.91 | 0.509 | 0.67 | 0.5 | 0.498 | 0.5 |  |  |  | #N/A
| 1 | 3 | 0.51 | 0.52 | 0.46 | 0.5 | 0.48211 | 0.91 | 0.501 | 0.64 | 0.49 | 0.49 | 0.5 |  |  |  | 16
| 1 | 4 | 0.45 | 0.49 | 0.48 | 0.4 | 0.46737 | 0.85 | 0.508 | 0.63 | 0.48 | 0.496 | 0.5 |  |  |  | 0.48635
| 1 | 5 | 0.49 | 0.48 | 0.5 | 0.4 | 0.58543 | 0.96 | 0.538 | 0.62 | 0.47 | 0.455 | 0.5 |  |  |  | 136
| 1 | 6 | 0.44 | 0.41 | 0.44 | 0.5 | 0.5632 | 0.9 | 0.505 | 0.64 | 0.46 | 0.47 | 0.5
| Mean |  | 0.45333 | 0.475 | 0.51167 | 0.48333 | 0.52445 | 0.88333 | 0.51417 | 0.645 | 0.47833 | 0.48533 | 0.5
| Median |  | 0.445 | 0.485 | 0.47 | 0.5 | 0.5243 | 0.905 | 0.5085 | 0.64 | 0.475 | 0.493 | 0.5
| Std Dev. |  | 0.039328 | 0.037283 | 0.10889 | 0.075277 | 0.045662 | 0.065625 | 0.014049 | 0.020736 | 0.01472 | 0.018779 | 0
| Rel.Std.Dev. |  | 8.675% | 7.849% | 21.28% | 15.57% | 8.707% | 7.429% | 2.732% | 3.215% | 3.077% | 3.869% | 0%
| PDM3 |  | -6.789% | -2.335% | 5.205% | -0.6211% | 7.834% | 81.62% | 5.719% | 32.62% | -1.649% | -0.2099% | 2.806%
| Z-Score (Absolute) |  | 1.24 | 0.4 | 0.67 | 0.08 | 1.52 | 16.36 | 0.94 | 6.21 | 0.27 | 0 | 0.57
| Analytical results for Cd in OREAS 922 (Certified Value < 0.5 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | < 0.2 | < 2 | < 2 | < 10 | < 10 | 0.269 | < 2 | < 1 | 0.3 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 2 | < 2 | < 10 | < 10 | 0.289 | < 2 | < 1 | 0.3 | [5 empty] | #N/A
| 1 | 3 | < 0.2 | < 2 | < 2 | < 10 | < 10 | 0.274 | < 2 | < 1 | 0.2 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 2 | < 2 | < 10 | < 10 | < 0.2 | < 2 | < 1 | 0.3 | [5 empty] | < 0.5
| 1 | 5 | < 0.2 | < 2 | < 2 | < 10 | < 10 | 0.223 | < 2 | < 1 | 0.3 | [5 empty] | 137
| 1 | 6 | < 0.2 | < 2 | < 2 | < 10 | < 10 | 0.269 | < 2 | < 1 | 0.3
| Mean |  | 16: 0.2648; 19: 0.28333
| Median |  | 16: 0.269; 19: 0.3
| Std Dev. |  | 16: 0.024763; 19: 0.040825
| Rel.Std.Dev. |  | 16: 9.352%; 19: 14.41%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.67 | NA | NA | 0.67
| Analytical results for Ce in OREAS 922 (Certified Value 88 ppm) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 0
1
| 1 | 1 | 96.1 | 84.7 | 93 | 89 | 88.16 | 80.8 | 83.8 | 89.2 | [6 empty] | 1
| 1 | 2 | 95.6 | 82.4 | 93.4 | 90 | 84.31 | 79.7 | 79.8 | 89.7 | [6 empty] | #N/A
| 1 | 3 | 96.4 | 83.3 | 93.6 | 88 | 84.04 | 79.6 | 86.5 | 88.6 | [6 empty] | 16
| 1 | 4 | 97 | 86.8 | 93.7 | 90 | 85.07 | 79.6 | 83.1 | 90 | [6 empty] | 87.5198
| 1 | 5 | 96.8 | 84.2 | 92.8 | 90.5 | 88.8 | 79.2 | 83.4 | 89.3 | [6 empty] | 138
| 1 | 6 | 96.7 | 78.6 | 90.6 | 90 | 83.87 | 78.9 | 81.5 | 90.8
| Mean |  | 96.4333 | 83.3333 | 92.85 | 89.5833 | 85.7083 | 79.6333 | 83.0167 | 89.6
| Median |  | 96.55 | 83.75 | 93.2 | 90 | 84.69 | 79.6 | 83.25 | 89.5
| Std Dev. |  | 0.5164 | 2.7537 | 1.1554 | 0.91742 | 2.1952 | 0.64704 | 2.2605 | 0.75631
| Rel.Std.Dev. |  | 0.5355% | 3.304% | 1.244% | 1.024% | 2.561% | 0.8125% | 2.723% | 0.8441%
| PDM3 |  | 10.18% | -4.783% | 6.09% | 2.358% | -2.07% | -9.011% | -5.145% | 2.377%
| Z-Score (Absolute) |  | 1.33 | 0.65 | 0.78 | 0.29 | 0.29 | 1.21 | 0.7 | 0.29
| Analytical results for Co in OREAS 922 (Certified Value 20.9 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*AAS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  | 1
2
| 1 | 1 | 21.2 | 22.3 | 21.1 | 20 | 20 | 21.5974 | < 20 | 21.265 | < 100 | 20.2 | 21.1 | 20.4 |  |  | 1
| 1 | 2 | 21.1 | 21.7 | 19 | 20 | 21 | 22.5168 | < 20 | 21.72 | < 100 | 19.5 | 20.4 | 19.3 |  |  | #N/A
| 1 | 3 | 21.3 | 23.1 | 20 | 20 | 20 | 21.3132 | < 20 | 20.823 | < 100 | 20.7 | 15.7 | 19.4 |  |  | 16
| 1 | 4 | 21.4 | 21.2 | 20 | 20 | 21 | 21.3063 | < 20 | 21.326 | < 100 | 20.9 | 19.3 | 19.8 |  |  | 20.9039
| 1 | 5 | 21.3 | 22.2 | 23 | 20 | 22 | 22.6546 | < 20 | 20.619 | < 100 | 20.7 | 25 | 20 |  |  | 139
| 1 | 6 | 21.2 | 23.9 | 21.4 | 20 | 22 | 23.2026 | < 20 | 20.792 | < 100 | 19.9 | 11.1 | 20.1
| Mean |  | 21.25 | 22.4 | 20.75 | 20 | 21 | 22.0985 |  | 21.0908 |  | 20.3167 | 18.7667 | 19.8333
| Median |  | 21.25 | 22.25 | 20.55 | 20 | 21 | 22.0571 |  | 21.044 |  | 20.45 | 19.85 | 19.9
| Std Dev. |  | 0.10488 | 0.9716 | 1.3996 | 0 | 0.89443 | 0.79979 |  | 0.41597 |  | 0.54559 | 4.8049 | 0.42269
| Rel.Std.Dev. |  | 0.4936% | 4.337% | 6.745% | 0% | 4.259% | 3.619% |  | 1.972% |  | 2.685% | 25.6% | 2.131%
| PDM3 |  | 1.656% | 7.157% | -0.7364% | -4.324% | 0.4596% | 5.714% |  | 0.8941% |  | -2.809% | -10.22% | -5.122%
| Z-Score (Absolute) |  | 0.45 | 1.81 | 0.15 | 1.04 | 0.15 | 1.46 | NA | 0.26 | NA | 0.66 | 0.68 | 1.24
| Analytical results for Cr in OREAS 922 (Certified Value 90 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 95 | 100 | 90 | 100 | 106.5838 | 100 | 86.196 | 159.625 | 90 | 77.9 | 80 |  |  |  | 1
| 1 | 2 | 95 | 70 | 80 | 100 | 103.8343 | 100 | 86.662 | 139.462 | 90 | 67.9 | 80 |  |  |  | #N/A
| 1 | 3 | 95 | 70 | 80 | 100 | 107.9586 | 100 | 81.967 | 137.502 | 100 | 84.3 | 80 |  |  |  | 16
| 1 | 4 | 98 | 70 | 90 | 100 | 103.4818 | 100 | 85.753 | 123.8134 | 100 | 85.4 | 80 |  |  |  | 90.2367
| 1 | 5 | 96 | 60 | 110 | 100 | 105.8555 | 100 | 85.513 | 177.0912 | 140 | 66.4 | 90 |  |  |  | 140
| 1 | 6 | 94 | 90 | 70 | 100 | 111.7335 | 100 | 85.863 | 132.6764 | 90 | 83.9 | 80
| Mean |  | 95.5 | 76.6667 | 86.6667 | 100 | 106.5746 | 100 | 85.3257 | 145.0283 | 101.6667 | 77.6333 | 81.6667
| Median |  | 95 | 70 | 85 | 100 | 106.2196 | 100 | 85.808 | 138.482 | 95 | 80.9 | 80
| Std Dev. |  | 1.3784 | 15.0555 | 13.6626 | 0 | 3.0372 | 0 | 1.6929 | 19.6538 | 19.4079 | 8.5444 | 4.0825
| Rel.Std.Dev. |  | 1.443% | 19.64% | 15.76% | 0% | 2.85% | 0% | 1.984% | 13.55% | 19.09% | 11.01% | 4.999%
| PDM3 |  | 5.833% | -15.04% | -3.956% | 10.82% | 18.11% | 10.82% | -5.442% | 60.72% | 12.67% | -13.97% | -9.497%
| Z-Score (Absolute) |  | 0.12 | 1.35 | 0.57 | 0.47 | 0.98 | 0.47 | 0.67 | 3.97 | 0 | 1.27 | 1.09
| Analytical results for Cs in OREAS 922 (Certified Value 7.5 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 8.2 | 7.9 | 7.8 | 7 | 7.2354 | 7.42 | 7 | 7.3 | 7.5 | [5 empty] | 1
| 1 | 2 | 8.1 | 7.8 | 7.5 | 8 | 7.3969 | 7.33 | 7.1 | 7.4 | 7.6 | [5 empty] | #N/A
| 1 | 3 | 8.2 | 7.7 | 7.7 | 8 | 7.6896 | 7.34 | 7 | 7.7 | 7.7 | [5 empty] | 16
| 1 | 4 | 8.2 | 7.8 | 7.4 | 8 | 7.4103 | 7.29 | 7 | 7.1 | 7.6 | [5 empty] | 7.495
| 1 | 5 | 8.2 | 7.7 | 7.1 | 8 | 7.3716 | 7.39 | 7.1 | 7.2 | 7.6 | [5 empty] | 141
| 1 | 6 | 8.2 | 7.7 | 7.3 | 8 | 7.3364 | 7.25 | 6.7 | 7.6 | 7.7
| Mean |  | 8.1833 | 7.7667 | 7.4667 | 7.8333 | 7.4067 | 7.3367 | 6.9833 | 7.3833 | 7.6167
| Median |  | 8.2 | 7.75 | 7.45 | 8 | 7.3842 | 7.335 | 7 | 7.35 | 7.6
| Std Dev. |  | 0.040825 | 0.08165 | 0.2582 | 0.40825 | 0.15207 | 0.062503 | 0.1472 | 0.23166 | 0.075277
| Rel.Std.Dev. |  | 0.4989% | 1.051% | 3.458% | 5.212% | 2.053% | 0.8519% | 2.108% | 3.138% | 0.9883%
| PDM3 |  | 9.184% | 3.625% | -0.3781% | 4.514% | -1.178% | -2.113% | -6.827% | -1.49% | 1.623%
| Z-Score (Absolute) |  | 3.22 | 1.35 | 0 | 2.4 | 0.27 | 0.58 | 2.17 | 0.37 | 0.67
| Analytical results for Cu in OREAS 922 (Certified Value 2215 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 2165 | 2280 | 2310 | 2200 | 2299 | 2192 | 2200 | 2145.764 | 2477.92 | 2270 | 2250 | 2060 |  |  | 1
| 1 | 2 | 2031 | 2310 | 2180 | 2200 | 2283 | 2217.6062 | 2100 | 2199.917 | 2296.24 | 2250 | 2240 | 2080 |  |  | #N/A
| 1 | 3 | 2258 | 2340 | 2240 | 2150 | 2276 | 2201.1859 | 2200 | 2130.21 | 2361.58 | 2270 | 2260 | 2090 |  |  | 16
| 1 | 4 | 2054 | 2350 | 2530 | 2350 | 2309 | 2186.7909 | 2100 | 2256.858 | 2309.12 | 2240 | 2240 | 2090 |  |  | 2214.5991
| 1 | 5 | 2025 | 2300 | 2150 | 2250 | 2303 | 2223.8558 | 2100 | 2175.197 | 2316.86 | 2250 | 2120 | 2130 |  |  | 142
| 1 | 6 | 2097 | 2280 | 2190 | 2200 | 2236 | 2230.7349 | 2100 | 2280.458 | 2278.84 | 2230 | 2170 | 2100
| Mean |  | 2105 | 2310 | 2266.6667 | 2225 | 2284.3333 | 2208.6956 | 2133.3333 | 2198.0673 | 2340.0933 | 2251.6667 | 2213.3333 | 2091.6667
| Median |  | 2075.5 | 2305 | 2215 | 2200 | 2291 | 2209.396 | 2100 | 2187.557 | 2312.99 | 2250 | 2240 | 2090
| Std Dev. |  | 91.1592 | 29.6648 | 140.6651 | 68.9202 | 26.7557 | 17.9439 | 51.6398 | 60.1852 | 72.9823 | 16.0208 | 55.7375 | 23.1661
| Rel.Std.Dev. |  | 4.331% | 1.284% | 6.206% | 3.098% | 1.171% | 0.8124% | 2.421% | 2.738% | 3.119% | 0.7115% | 2.518% | 1.108%
| PDM3 |  | -4.949% | 4.308% | 2.351% | 0.4697% | 3.149% | -0.2666% | -3.67% | -0.7465% | 5.667% | 1.674% | -0.05716% | -5.551%
| Z-Score (Absolute) |  | 1.35 | 1.2 | 0 | 0.14 | 0.88 | 0.06 | 1 | 0.19 | 1.57 | 0.47 | 0 | 1.51
| Analytical results for Dy in OREAS 922 (Certified Value 5.75 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 5.88 | 5.5 | 6.2 | 6 | 5.749 | 6.7 | 5.8 | 5.8 | [6 empty] | 1
| 1 | 2 | 5.84 | 5.5 | 5.9 | 5.5 | 5.683 | 6.8 | 5.6 | 5.98 | [6 empty] | #N/A
| 1 | 3 | 5.87 | 5.4 | 6.1 | 5.5 | 5.616 | 7 | 6.1 | 5.92 | [6 empty] | 16
| 1 | 4 | 5.89 | 5.5 | 6 | 5.5 | 5.639 | 6.7 | 5.5 | 5.91 | [6 empty] | 5.7482
| 1 | 5 | 5.89 | 5.5 | 5.9 | 5.5 | 5.919 | 6.9 | 5.8 | 5.98 | [6 empty] | 143
| 1 | 6 | 5.87 | 5.4 | 6.1 | 5.5 | 5.637 | 6.7 | 5.7 | 5.85
| Mean |  | 5.8733 | 5.4667 | 6.0333 | 5.5833 | 5.7072 | 6.8 | 5.75 | 5.9067
| Median |  | 5.875 | 5.5 | 6.05 | 5.5 | 5.661 | 6.75 | 5.75 | 5.915
| Std Dev. |  | 0.018619 | 0.05164 | 0.12111 | 0.20412 | 0.1141 | 0.12649 | 0.20736 | 0.071461
| Rel.Std.Dev. |  | 0.317% | 0.9446% | 2.007% | 3.656% | 1.999% | 1.86% | 3.606% | 1.21%
| PDM3 |  | 2.178% | -4.897% | 4.961% | -2.868% | -0.7133% | 18.3% | 0.03189% | 2.757%
| Z-Score (Absolute) |  | 0.25 | 1.43 | 0.92 | 1.29 | 0.43 | 4.09 | 0.25 | 0.39
| Analytical results for Er in OREAS 922 (Certified Value 3.38 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 3.3 | 3.7 | 3.5 | 3.5 | 2.916 | 3.8 | 3.1 | 3.27 | [6 empty] | 1
| 1 | 2 | 3.32 | 3.6 | 3.3 | 3.5 | 2.98 | 3.7 | 3 | 3.3 | [6 empty] | #N/A
| 1 | 3 | 3.31 | 3.7 | 3.5 | 3.5 | 2.921 | 3.9 | 3.4 | 3.35 | [6 empty] | 16
| 1 | 4 | 3.33 | 3.6 | 3.5 | 3.5 | 2.982 | 3.9 | 2.9 | 3.2 | [6 empty] | 3.3803
| 1 | 5 | 3.35 | 3.7 | 3.3 | 3.5 | 3.196 | 3.9 | 3 | 3.35 | [6 empty] | 144
| 1 | 6 | 3.37 | 3.6 | 3.4 | 3.5 | 2.939 | 3.7 | 3.2 | 3.22
| Mean |  | 3.33 | 3.65 | 3.4167 | 3.5 | 2.989 | 3.8167 | 3.1 | 3.2817
| Median |  | 3.325 | 3.65 | 3.45 | 3.5 | 2.9595 | 3.85 | 3.05 | 3.285
| Std Dev. |  | 0.026077 | 0.054772 | 0.098319 | 0 | 0.10529 | 0.098319 | 0.17889 | 0.063692
| Rel.Std.Dev. |  | 0.7831% | 1.501% | 2.878% | 0% | 3.523% | 2.576% | 5.77% | 1.941%
| PDM3 |  | -1.489% | 7.978% | 1.075% | 3.54% | -11.58% | 12.91% | -8.293% | -2.919%
| Z-Score (Absolute) |  | 0.15 | 0.93 | 0.15 | 0.43 | 1.44 | 1.49 | 0.92 | 0.31
| Analytical results for Eu in OREAS 922 (Certified Value 1.52 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.62 | 1.6 | 1.5 | 1.6 | 1.453 | 1.6 | 1.4 | 1.45 | [6 empty] | 1
| 1 | 2 | 1.63 | 1.5 | 1.4 | 1.6 | 1.403 | 1.6 | 1.5 | 1.45 | [6 empty] | #N/A
| 1 | 3 | 1.63 | 1.6 | 1.5 | 1.6 | 1.413 | 1.7 | 1.4 | 1.49 | [6 empty] | 16
| 1 | 4 | 1.63 | 1.5 | 1.5 | 1.6 | 1.414 | 1.6 | 1.3 | 1.5 | [6 empty] | 1.5229
| 1 | 5 | 1.64 | 1.6 | 1.5 | 1.6 | 1.459 | 1.6 | 1.3 | 1.47 | [6 empty] | 145
| 1 | 6 | 1.63 | 1.5 | 1.5 | 1.6 | 1.398 | 1.7 | 1.3 | 1.52
| Mean |  | 1.63 | 1.55 | 1.4833 | 1.6 | 1.4233 | 1.6333 | 1.3667 | 1.48
| Median |  | 1.63 | 1.55 | 1.5 | 1.6 | 1.4135 | 1.6 | 1.35 | 1.48
| Std Dev. |  | 0.0063246 | 0.054772 | 0.040825 | 2.4324e-16 | 0.026082 | 0.05164 | 0.08165 | 0.028284
| Rel.Std.Dev. |  | 0.388% | 3.534% | 2.752% | 1.52e-14% | 1.832% | 3.162% | 5.974% | 1.911%
| PDM3 |  | 7.031% | 1.778% | -2.599% | 5.062% | -6.539% | 7.25% | -10.26% | -2.818%
| Z-Score (Absolute) |  | 0.8 | 0.19 | 0.19 | 0.57 | 0.78 | 0.83 | 1.21 | 0.34
| Analytical results for Fe in OREAS 922 (Certified Value 5.71 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 08 | 09 | 10 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 5.38 | 5.96 | 5.57 | 6.105 | 5.8 | 5.99 | 5.77 | 5.11 | 5.532 | 6.7535 | 5.85 | 5.8 | 5.71 | 5.63 | 1
| 1 | 2 | 4.82 | 5.92 | 5.61 | 6.0659 | 5.88 | 5.97 | 5.8 | 5.39 | 5.533 | 6.239 | 5.86 | 5.79 | 5.68 | 5.54 | #N/A
| 1 | 3 | 5.76 | 5.83 | 5.6 | 6.0275 | 5.86 | 6.01 | 5.8 | 5.47 | 5.44 | 6.7996 | 5.71 | 6.02 | 5.69 | 5.75 | 16
| 1 | 4 | 5.45 | 5.54 | 5.55 | 6.0465 | 5.86 | 6 | 5.77 | 5.37 | 5.48 | 6.1548 | 5.67 | 5.83 | 5.66 | 5.62 | 5.7119
| 1 | 5 | 5.64 | 5.63 | 5.59 | 6.0172 | 5.89 | 5.95 | 5.81 | 5.32 | 5.523 | 6.3196 | 5.65 | 5.79 | 5.43 | 5.8 | 146
| 1 | 6 | 5.7 | 5.53 | 5.55 | 6.0862 | 5.78 | 5.89 | 5.86 | 5.24 | 5.51 | 6.2077 | 5.69 | 5.78 | 5.54 | 5.68
| Mean |  | 5.4583 | 5.735 | 5.5783 | 6.0581 | 5.845 | 5.9683 | 5.8017 | 5.3167 | 5.503 | 6.4124 | 5.7383 | 5.835 | 5.6183 | 5.67
| Median |  | 5.545 | 5.73 | 5.58 | 6.0562 | 5.86 | 5.98 | 5.8 | 5.345 | 5.5165 | 6.2793 | 5.7 | 5.795 | 5.67 | 5.655
| Std Dev. |  | 0.34528 | 0.19233 | 0.025626 | 0.034058 | 0.044609 | 0.044008 | 0.033116 | 0.12675 | 0.036573 | 0.28748 | 0.09261 | 0.09225 | 0.11017 | 0.094234
| Rel.Std.Dev. |  | 6.326% | 3.354% | 0.4594% | 0.5622% | 0.7632% | 0.7374% | 0.5708% | 2.384% | 0.6646% | 4.483% | 1.614% | 1.581% | 1.961% | 1.662%
| PDM3 |  | -4.439% | 0.4048% | -2.338% | 6.061% | 2.331% | 4.49% | 1.572% | -6.919% | -3.657% | 12.26% | 0.4632% | 2.156% | -1.638% | -0.7332%
| Z-Score (Absolute) |  | 0.78 | 0.01 | 0.82 | 1.67 | 0.56 | 1.21 | 0.34 | 2.19 | 1.22 | 3.52 | 0.01 | 0.32 | 0.42 | 0.35
| Analytical results for Ga in OREAS 922 (Certified Value 21.2 ppm) | [15 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 1
2
| 1 | 1 | 19 | 20.8 | 21.7 | 24 | 22.0576 | 21.357 | 21.6 | 22.5 | 20 | [5 empty] | 1
| 1 | 2 | 19 | 20.6 | 20.7 | 28 | 23.4909 | 21.166 | 21.5 | 22.4 | 20 | [5 empty] | #N/A
| 1 | 3 | 19 | 20.6 | 21.9 | 22 | 22.428 | 20.915 | 21.7 | 23.7 | 20 | [5 empty] | 16
| 1 | 4 | 19 | 20.9 | 21.1 | 20 | 22.4205 | 21.015 | 21.5 | 21.4 | 20 | [5 empty] | 21.2177
| 1 | 5 | 19 | 20.7 | 20.3 | 26 | 22.1345 | 20.995 | 21.6 | 21.8 | 20 | [5 empty] | 147
| 1 | 6 | 19 | 20.7 | 20.8 | 24 | 22.1976 | 20.721 | 21 | 21.4 | 20
| Mean |  | 19 | 20.7167 | 21.0833 | 24 | 22.4549 | 21.0282 | 21.4833 | 22.2 | 20
| Median |  | 19 | 20.7 | 20.95 | 24 | 22.3091 | 21.005 | 21.55 | 22.1 | 20
| Std Dev. |  | 0 | 0.1169 | 0.61455 | 2.8284 | 0.52953 | 0.21705 | 0.24833 | 0.87407 | 0
| Rel.Std.Dev. |  | 0% | 0.5643% | 2.915% | 11.79% | 2.358% | 1.032% | 1.156% | 3.937% | 0%
| PDM3 |  | -10.45% | -2.361% | -0.6332% | 13.11% | 5.831% | -0.8932% | 1.252% | 4.63% | -5.739%
| Z-Score (Absolute) |  | 1.3 | 0.23 | 0 | 1.32 | 0.72 | 0.03 | 0.25 | 0.7 | 0.67
| Analytical results for Gd in OREAS 922 (Certified Value 6.94 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 7.05 | 7.4 | 7.4 | 6 | 6.425 | 7.5 | 5.8 | 6.12 | [6 empty] | 1
| 1 | 2 | 7 | 7.3 | 7.1 | 8 | 6.494 | 7.3 | 5.9 | 6.15 | [6 empty] | #N/A
| 1 | 3 | 7.07 | 7.4 | 7.2 | 8 | 6.133 | 7.4 | 6.4 | 6.28 | [6 empty] | 16
| 1 | 4 | 7.15 | 7.4 | 7.1 | 8 | 6.189 | 7.5 | 6.2 | 6.29 | [6 empty] | 6.9411
| 1 | 5 | 7.12 | 7.4 | 6.9 | 8 | 6.403 | 7.7 | 6.5 | 6.18 | [6 empty] | 148
| 1 | 6 | 7.07 | 7 | 6.9 | 8 | 6.109 | 7.2 | 5.5 | 6.16
| Mean |  | 7.0767 | 7.3167 | 7.1 | 7.6667 | 6.2922 | 7.4333 | 6.05 | 6.1967
| Median |  | 7.07 | 7.4 | 7.1 | 8 | 6.296 | 7.45 | 6.05 | 6.17
| Std Dev. |  | 0.052789 | 0.16021 | 0.18974 | 0.8165 | 0.16745 | 0.17512 | 0.38341 | 0.071181
| Rel.Std.Dev. |  | 0.746% | 2.19% | 2.672% | 10.65% | 2.661% | 2.356% | 6.337% | 1.149%
| PDM3 |  | 1.953% | 5.411% | 2.289% | 10.45% | -9.349% | 7.092% | -12.84% | -10.73%
| Z-Score (Absolute) |  | 0.01 | 0.34 | 0.01 | 1.08 | 0.94 | 0.41 | 1.23 | 1.05
| Analytical results for Ge in OREAS 922 (Indicative Value 2.36 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 2 | 3.6 | 11 | < 20 | 2.3634 | 2.243 | 3.3 | 2 | [6 empty] | 1
| 1 | 2 | 2 | 3.9 | 9.2 | < 20 | 2.642 | 2.074 | 3.7 | 2 | [6 empty] | 10
| 1 | 3 | 2 | 4 | 11.4 | < 20 | 2.022 | 2.156 | 3.1 | 2 | [6 empty] | 16
| 1 | 4 | 2 | 5 | 6.4 | < 20 | 2.1112 | 2.426 | 3.9 | 2 | [6 empty] | 2.3611
| 1 | 5 | 2 | 4.2 | 5.5 | < 20 | 2.2883 | 2.321 | 3 | 2 | [6 empty] | 16
| 1 | 6 | 2 | 3.3 | 6.7 | < 20 | 1.7847 | 2.202 | 3.2 | 2
| Mean |  | 2 | 4 | 8.3667 |  | 2.2019 | 2.237 | 3.3667 | 2
| Median |  | 2 | 3.95 | 7.95 |  | 2.1997 | 2.2225 | 3.25 | 2
| Std Dev. |  | 0 | 0.5831 | 2.5177 |  | 0.29728 | 0.12423 | 0.3559 | 0
| Rel.Std.Dev. |  | 0% | 14.58% | 30.09% |  | 13.5% | 5.554% | 10.57% | 0%
| PDM3 |  | -15.29% | 69.41% | 254.4% |  | -6.742% | -5.257% | 42.59% | -15.29%
| Z-Score (Absolute) |  | 0.67 | 5.02 | 17.44 | NA | 0.1 | 0 | 3.21 | 0.67
| Analytical results for Hf in OREAS 922 (Certified Value 5.93 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 7 | < 10 | < 10 | 6 | 6.561 | < 10 | 6 | 5 | [6 empty] | 1
| 1 | 2 | 7 | < 10 | < 10 | 6 | 6.036 | < 10 | 6 | 5 | [6 empty] | #N/A
| 1 | 3 | 7 | < 10 | < 10 | 6 | 6.305 | < 10 | 6.4 | 6 | [6 empty] | 16
| 1 | 4 | 7 | < 10 | < 10 | 6 | 6.241 | < 10 | 5.4 | 5 | [6 empty] | 5.929
| 1 | 5 | 7 | < 10 | 10 | 6 | 6.386 | < 10 | 6.2 | 5 | [6 empty] | 149
| 1 | 6 | 7 | < 10 | < 10 | 6 | 6.115 | < 10 | 5.8 | 6
| Mean |  | 02: 7; 05: 10; 13: 6; 15: 6.274; 17: 5.9667; 18: 5.3333
| Median |  | 02: 7; 05: 10; 13: 6; 15: 6.273; 17: 6; 18: 5
| Std Dev. |  | 02: 0; 13: 0; 15: 0.18915; 17: 0.34448; 18: 0.5164
| Rel.Std.Dev. |  | 02: 0%; 13: 0%; 15: 3.015%; 17: 5.773%; 18: 9.682%
| PDM3 |  | 02: 18.06%; 05: 68.66%; 13: 1.198%; 15: 5.819%; 17: 0.6353%; 18: -10.05%
| Z-Score (Absolute) |  | 1.46 | NA | 1.41 | 0.02 | 0.42 | NA | 0.02 | 0.93
| Analytical results for Ho in OREAS 922 (Certified Value 1.2 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.25 | 1.3 | 1.2 | 1.2 | 1.106 | 1.3 | 1.1 | 1.19 | [6 empty] | 1
| 1 | 2 | 1.24 | 1.3 | 1.2 | 1 | 1.114 | 1.3 | 1.1 | 1.21 | [6 empty] | #N/A
| 1 | 3 | 1.26 | 1.3 | 1.2 | 1.2 | 1.119 | 1.3 | 1.2 | 1.2 | [6 empty] | 16
| 1 | 4 | 1.26 | 1.3 | 1.2 | 1.2 | 1.088 | 1.4 | 1.1 | 1.18 | [6 empty] | 1.2003
| 1 | 5 | 1.25 | 1.3 | 1.2 | 1 | 1.125 | 1.3 | 1.1 | 1.2 | [6 empty] | 150
| 1 | 6 | 1.25 | 1.2 | 1.2 | 1.2 | 1.092 | 1.3 | 1.1 | 1.18
| Mean |  | 1.2517 | 1.2833 | 1.2 | 1.1333 | 1.1073 | 1.3167 | 1.1167 | 1.1933
| Median |  | 1.25 | 1.3 | 1.2 | 1.2 | 1.11 | 1.3 | 1.1 | 1.195
| Std Dev. |  | 0.0075277 | 0.040825 | 0 | 0.10328 | 0.014855 | 0.040825 | 0.040825 | 0.012111
| Rel.Std.Dev. |  | 0.6014% | 3.181% | 0% | 9.113% | 1.341% | 3.101% | 3.656% | 1.015%
| PDM3 |  | 4.28% | 6.918% | -0.0243% | -5.579% | -7.745% | 9.696% | -6.967% | -0.5797%
| Z-Score (Absolute) |  | 0.52 | 0.82 | 0.03 | 0.6 | 0.84 | 1.13 | 0.75 | 0.03
| Analytical results for In in OREAS 922 (Certified Value 0.34 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.34373 | 0.323 | 0.3 | 0.4 | 0.3 | [5 empty] | 1
| 1 | 2 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.2728 | 0.331 | 0.3 | 0.3 | 0.3 | [5 empty] | #N/A
| 1 | 3 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.31523 | 0.339 | 0.3 | 0.4 | 0.4 | [5 empty] | 16
| 1 | 4 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.25534 | 0.323 | 0.3 | 0.3 | 0.3 | [5 empty] | 0.33525
| 1 | 5 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.36018 | 0.319 | 0.3 | 0.3 | 0.3 | [5 empty] | 151
| 1 | 6 | 0.4 | < 0.2 | 0.3 | 0.4 | 0.37769 | 0.332 | 0.3 | 0.3 | 0.3
| Mean |  | 0.4 |  | 0.3 | 0.4 | 0.32083 | 0.32783 | 0.3 | 0.33333 | 0.31667
| Median |  | 0.4 |  | 0.3 | 0.4 | 0.32948 | 0.327 | 0.3 | 0.3 | 0.3
| Std Dev. |  | 6.0809e-17 |  | 0 | 6.0809e-17 | 0.048853 | 0.0074409 | 0 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 1.52e-14% |  | 0% | 1.52e-14% | 15.23% | 2.27% | 0% | 15.49% | 12.89%
| PDM3 |  | 19.31% |  | -10.51% | 19.31% | -4.302% | -2.212% | -10.51% | -0.5715% | -5.543%
| Z-Score (Absolute) |  | 2.1 | NA | 0.67 | 2.1 | 0.1 | 0.1 | 0.67 | 0.25 | 0.67
| Analytical results for K in OREAS 922 (Certified Value 2.6 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.66 | 2.7 | 2.6 | 2.6 | 3.0677 | 2.58 | 2.414 | 3.3408 | 3 | 2.59 | 2.73 | 2.7 |  |  | 1
| 1 | 2 | 2.45 | 2.7 | 2.6 | 2.6 | 3.1276 | 2.58 | 2.453 | 3.0382 | 3 | 2.6 | 2.72 | 2.7 |  |  | #N/A
| 1 | 3 | 2.75 | 2.6 | 2.5 | 2.6 | 3.0044 | 2.59 | 2.348 | 3.1555 | 3 | 2.62 | 2.72 | 2.7 |  |  | 16
| 1 | 4 | 2.48 | 2.6 | 2.6 | 2.7 | 2.9946 | 2.58 | 2.421 | 3.0319 | 3 | 2.63 | 2.72 | 2.7 |  |  | 2.5969
| 1 | 5 | 2.49 | 2.6 | 2.5 | 2.8 | 2.9842 | 2.51 | 2.528 | 3.0303 | 2.9 | 2.63 | 2.58 | 2.7 |  |  | 152
| 1 | 6 | 2.56 | 2.6 | 2.5 | 2.6 | 3.0731 | 2.42 | 2.431 | 3.0337 | 2.9 | 2.61 | 2.64 | 2.7
| Mean |  | 2.565 | 2.6333 | 2.55 | 2.65 | 3.042 | 2.5433 | 2.4325 | 3.1051 | 2.9667 | 2.6133 | 2.685 | 2.7
| Median |  | 2.525 | 2.6 | 2.55 | 2.6 | 3.0361 | 2.58 | 2.426 | 3.0359 | 3 | 2.615 | 2.72 | 2.7
| Std Dev. |  | 0.11777 | 0.05164 | 0.054772 | 0.083666 | 0.056476 | 0.067132 | 0.058579 | 0.12541 | 0.05164 | 0.01633 | 0.061237 | 4.8648e-16
| Rel.Std.Dev. |  | 4.591% | 1.961% | 2.148% | 3.157% | 1.857% | 2.64% | 2.408% | 4.039% | 1.741% | 0.6249% | 2.281% | 1.802e-14%
| PDM3 |  | -1.23% | 1.401% | -1.808% | 2.043% | 17.14% | -2.064% | -6.332% | 19.57% | 14.24% | 0.6311% | 3.391% | 3.968%
| Z-Score (Absolute) |  | 0.61 | 0.07 | 0.73 | 0.07 | 3.21 | 0.79 | 1.68 | 3.33 | 2.6 | 0.23 | 0.35 | 0.47
| Analytical results for La in OREAS 922 (Certified Value 45.6 ppm) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 20 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | [5 empty] | 1
2
| 1 | 1 | 47.7 | 47.3 | 46.2 | 42 | 45.17 | 43.8 | 47.8 | 60.3 | 46.3 | [5 empty] | 1
| 1 | 2 | 47.4 | 46.8 | 47 | 43 | 43.1 | 43.3 | 44.1 | 58.2 | 46.9 | [5 empty] | #N/A
| 1 | 3 | 47.9 | 46.9 | 47 | 42.5 | 42.93 | 43.3 | 46.4 | 38.1 | 46.1 | [5 empty] | 16
| 1 | 4 | 48.3 | 48.6 | 47 | 43 | 43.52 | 43.2 | 45.1 | 34.6 | 46.8 | [5 empty] | 45.6006
| 1 | 5 | 48 | 47.7 | 49.4 | 43 | 45.35 | 43.2 | 46.1 | 61.3 | 46.2 | [5 empty] | 153
| 1 | 6 | 48.1 | 45.6 | 45.3 | 41.5 | 42.71 | 42.9 | 45.5 | 54.8 | 46.9
| Mean |  | 47.9 | 47.15 | 46.9833 | 42.5 | 43.7967 | 43.2833 | 45.8333 | 51.2167 | 46.5333
| Median |  | 47.95 | 47.1 | 47 | 42.75 | 43.31 | 43.25 | 45.8 | 56.5 | 46.55
| Std Dev. |  | 0.31623 | 1.0015 | 1.363 | 0.63246 | 1.1656 | 0.29269 | 1.258 | 11.7809 | 0.37238
| Rel.Std.Dev. |  | 0.6602% | 2.124% | 2.901% | 1.488% | 2.661% | 0.6762% | 2.745% | 23% | 0.8002%
| PDM3 |  | 5.043% | 3.398% | 3.032% | -6.799% | -3.956% | -5.082% | 0.5105% | 12.32% | 2.046%
| Z-Score (Absolute) |  | 1.37 | 0.67 | 0.52 | 3.65 | 2.44 | 2.92 | 0.55 | 0 | 0.1
| Analytical results for Li in OREAS 922 (Certified Value 28.8 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 20 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | [5 empty] | 1
1
| 1 | 1 | 27 | 30 | 32 | 31 | 28.1752 | 32.464 | 28 | 28.6 | 20 | [5 empty] | 1
| 1 | 2 | 25 | 30 | 31 | 31 | 27.785 | 32.28 | 41 | 26.5 | 20 | [5 empty] | #N/A
| 1 | 3 | 27 | 30 | 32 | 30 | 28.5157 | 31.564 | 50 | 26.4 | 30 | [5 empty] | 16
| 1 | 4 | 25 | 30 | 31 | 33 | 29.8264 | 32.268 | 49 | 22.9 | 30 | [5 empty] | 28.8037
| 1 | 5 | 26 | 31 | 31 | 31 | 26.7077 | 33.599 | 50 | 24.5 | 30 | [5 empty] | 154
| 1 | 6 | 25 | 30 | 32 | 31 | 27.7666 | 31.826 | 49 | 21.1 | 30
| Mean |  | 25.8333 | 30.1667 | 31.5 | 31.1667 | 28.1294 | 32.3335 | 44.5 | 25 | 26.6667
| Median |  | 25.5 | 30 | 31.5 | 31 | 27.9801 | 32.274 | 49 | 25.45 | 30
| Std Dev. |  | 0.98319 | 0.40825 | 0.54772 | 0.98319 | 1.0298 | 0.70361 | 8.7807 | 2.7218 | 5.164
| Rel.Std.Dev. |  | 3.806% | 1.353% | 1.739% | 3.155% | 3.661% | 2.176% | 19.73% | 10.89% | 19.36%
| PDM3 |  | -10.31% | 4.732% | 9.361% | 8.204% | -2.341% | 12.25% | 54.49% | -13.21% | -7.419%
| Z-Score (Absolute) |  | 1.35 | 0 | 0.41 | 0.2 | 0.63 | 0.67 | 5.49 | 1.61 | 1.09
| Analytical results for Lu in OREAS 922 (Certified Value 0.49 ppm) | [15 empty] | Certified
Lu | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab; 15: Lab |  |  |  | 1
| No | No | 02: 02; 04: 09; 05: 16; 13: 19; 15: 21 |  |  |  | ppm
|  |  | 02: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 13: PF*MS; 15: PF*OES/MS |  |  |  | 2
2
| 1 | 1 | 02: 0.59; 04: 0.4; 05: 0.474; 13: 0.4; 15: 0.51 |  |  |  | 1
| 1 | 2 | 02: 0.59; 04: 0.4; 05: 0.521; 13: 0.4; 15: 0.51 |  |  |  | #N/A
| 1 | 3 | 02: 0.59; 04: 0.6; 05: 0.478; 13: 0.5; 15: 0.5 |  |  |  | 16
| 1 | 4 | 02: 0.6; 04: 0.4; 05: 0.49; 13: 0.4; 15: 0.48 |  |  |  | 0.48553
| 1 | 5 | 02: 0.59; 04: 0.4; 05: 0.511; 13: 0.5; 15: 0.48 |  |  |  | 155
| 1 | 6 | 02: 0.59; 04: 0.4; 05: 0.472; 13: 0.5; 15: 0.49
| Mean |  | 02: 0.59167; 04: 0.43333; 05: 0.491; 13: 0.45; 15: 0.495
| Median |  | 02: 0.59; 04: 0.4; 05: 0.484; 13: 0.45; 15: 0.495
| Std Dev. |  | 02: 0.0040825; 04: 0.08165; 05: 0.020591; 13: 0.054772; 15: 0.013784
| Rel.Std.Dev. |  | 02: 0.69%; 04: 18.84%; 05: 4.194%; 13: 12.17%; 15: 2.785%
| PDM3 |  | 02: 21.86%; 04: -10.75%; 05: 1.126%; 13: -7.318%; 15: 1.95%
| Z-Score (Absolute) |  | 02: 1.66; 04: 1.5; 05: 0; 13: 0.67; 15: 0.07
| Analytical results for Mg in OREAS 922 (Certified Value 1.61 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 2
3
| 1 | 1 | 1.46 | 1.67 | 1.67 | 1.67 | 1.568 | 1.48 | 1.589 | 1.8315 | 1.61 | 1.68 | 1.58 | 1.61 |  |  | 1
| 1 | 2 | 1.43 | 1.65 | 1.61 | 1.69 | 1.585 | 1.52 | 1.587 | 1.6836 | 1.6 | 1.66 | 1.57 | 1.6 |  |  | #N/A
| 1 | 3 | 1.5 | 1.63 | 1.59 | 1.67 | 1.591 | 1.52 | 1.555 | 1.7456 | 1.64 | 1.68 | 1.57 | 1.63 |  |  | 16
| 1 | 4 | 1.46 | 1.61 | 1.6 | 1.7 | 1.581 | 1.55 | 1.58 | 1.6731 | 1.64 | 1.71 | 1.55 | 1.62 |  |  | 1.6127
| 1 | 5 | 1.42 | 1.62 | 1.58 | 1.69 | 1.618 | 1.52 | 1.601 | 1.6564 | 1.59 | 1.7 | 1.52 | 1.63 |  |  | 156
| 1 | 6 | 1.44 | 1.57 | 1.59 | 1.68 | 1.605 | 1.46 | 1.583 | 1.6567 | 1.59 | 1.71 | 1.53 | 1.61
| Mean |  | 1.4517 | 1.625 | 1.6067 | 1.6833 | 1.5913 | 1.5083 | 1.5825 | 1.7078 | 1.6117 | 1.69 | 1.5533 | 1.6167
| Median |  | 1.45 | 1.625 | 1.595 | 1.685 | 1.588 | 1.52 | 1.585 | 1.6784 | 1.605 | 1.69 | 1.56 | 1.615
| Std Dev. |  | 0.028577 | 0.034496 | 0.03266 | 0.012111 | 0.017829 | 0.032506 | 0.015281 | 0.068954 | 0.023166 | 0.02 | 0.024221 | 0.012111
| Rel.Std.Dev. |  | 1.969% | 2.123% | 2.033% | 0.7194% | 1.12% | 2.155% | 0.9656% | 4.038% | 1.437% | 1.183% | 1.559% | 0.7491%
| PDM3 |  | -9.983% | 0.7653% | -0.3716% | 4.382% | -1.322% | -6.469% | -1.87% | 5.901% | -0.06152% | 4.796% | -3.679% | 0.2485%
| Z-Score (Absolute) |  | 2.84 | 0.42 | 0.17 | 1.51 | 0.22 | 1.78 | 0.38 | 1.51 | 0.17 | 1.64 | 0.93 | 0.26
| Analytical results for Mn in OREAS 922 (Certified Value 0.088 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 17 | 18 | 20 | 21 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.08 | 0.0857 | 0.0853 | 0.086 | 0.0907 | 0.09 | 0.089085 | 0.10483 | 0.0873 | 0.0884 | 0.087 |  |  |  | 1
| 1 | 2 | 0.08 | 0.0849 | 0.0823 | 0.086 | 0.0912 | 0.09 | 0.08884 | 0.095841 | 0.0878 | 0.0879 | 0.087 |  |  |  | #N/A
| 1 | 3 | 0.08 | 0.0842 | 0.0862 | 0.084 | 0.0922 | 0.09 | 0.087797 | 0.10124 | 0.089 | 0.0876 | 0.095 |  |  |  | 16
| 1 | 4 | 0.08 | 0.0857 | 0.0836 | 0.088 | 0.0917 | 0.09 | 0.088128 | 0.094012 | 0.0888 | 0.0872 | 0.089 |  |  |  | 0.087587
| 1 | 5 | 0.08 | 0.0847 | 0.0809 | 0.088 | 0.0929 | 0.09 | 0.088691 | 0.095565 | 0.0884 | 0.0837 | 0.09 |  |  |  | 157
| 1 | 6 | 0.08 | 0.0851 | 0.0827 | 0.084 | 0.0945 | 0.09 | 0.088057 | 0.095146 | 0.0867 | 0.0854 | 0.089
| Mean |  | 0.08 | 0.08505 | 0.0835 | 0.086 | 0.0922 | 0.09 | 0.088433 | 0.097772 | 0.088 | 0.0867 | 0.0895
| Median |  | 0.08 | 0.085 | 0.08315 | 0.086 | 0.09195 | 0.09 | 0.088409 | 0.095703 | 0.0881 | 0.0874 | 0.089
| Std Dev. |  | 0 | 0.00058566 | 0.0019688 | 0.0017889 | 0.0013624 | 1.5202e-17 | 0.00050937 | 0.004276 | 0.00089666 | 0.0017933 | 0.0029496
| Rel.Std.Dev. |  | 0% | 0.6886% | 2.358% | 2.08% | 1.478% | 1.689e-14% | 0.576% | 4.373% | 1.019% | 2.068% | 3.296%
| PDM3 |  | -8.662% | -2.897% | -4.666% | -1.812% | 5.267% | 2.755% | 0.966% | 11.63% | 0.4715% | -1.013% | 2.184%
| Z-Score (Absolute) |  | 2.7 | 0.99 | 1.52 | 0.67 | 1.42 | 0.67 | 0.15 | 2.82 | 0 | 0.44 | 0.13
| Analytical results for Mo in OREAS 922 (Certified Value < 1 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 20 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 2 | 7 | < 1 | < 5 | < 1 | < 2 | < 1 | < 50 | < 2 | [5 empty] | 1
| 1 | 2 | 2 | < 1 | < 1 | < 5 | < 1 | < 2 | < 1 | < 50 | 3 | [5 empty] | #N/A
| 1 | 3 | 2 | < 1 | < 1 | < 5 | < 1 | < 2 | < 1 | < 50 | < 2 | [5 empty] | 16
| 1 | 4 | 2 | < 1 | < 1 | < 5 | < 1 | < 2 | < 1 | < 50 | < 2 | [5 empty] | < 1
| 1 | 5 | 2 | < 1 | < 1 | < 5 | < 1 | < 2 | 2 | < 50 | 2 | [5 empty] | 158
| 1 | 6 | 2 | 3 | < 1 | < 5 | < 1 | < 2 | < 1 | < 50 | < 2
| Mean |  | 02: 2; 04: 5; 17: 2; 19: 2.5
| Median |  | 02: 2; 04: 5; 17: 2; 19: 2.5
| Std Dev. |  | 02: 0; 04: 2.8284; 19: 0.70711
| Rel.Std.Dev. |  | 02: 0%; 04: 56.57%; 19: 28.28%
PDM3
| Z-Score (Absolute) |  | 1.73 | 0.19 | NA | NA | NA | NA | 1.16 | NA | 0.19
| Analytical results for Nb in OREAS 922 (Certified Value 15.2 ppm) | [15 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 1
1
| 1 | 1 | 17 | 13.7 | 15.2 | 15 | 11.4444 | 15.865 | 16 | 18.3 | 8 | [5 empty] | 1
| 1 | 2 | 17 | 13.6 | 14.4 | 15 | 11.2149 | 16.042 | 15.8 | 18 | 8 | [5 empty] | #N/A
| 1 | 3 | 17 | 34.1 | 14.4 | 15 | 10.6489 | 16.033 | 16.4 | 18.4 | 9 | [5 empty] | 16
| 1 | 4 | 17 | 14 | 14.6 | 15 | 11.2213 | 15.812 | 16 | 17.8 | 9 | [5 empty] | 15.227
| 1 | 5 | 17 | 14.2 | 14.3 | 15 | 11.5991 | 16.043 | 16.1 | 17.1 | 8 | [5 empty] | 159
| 1 | 6 | 17 | 15.3 | 14 | 15 | 12.3216 | 16.051 | 15.6 | 18.1 | 10
| Mean |  | 17 | 17.4833 | 14.4833 | 15 | 11.4084 | 15.9743 | 15.9833 | 17.95 | 8.6667
| Median |  | 17 | 14.1 | 14.4 | 15 | 11.3329 | 16.0375 | 16 | 18.05 | 8.5
| Std Dev. |  | 0 | 8.1632 | 0.40208 | 0 | 0.55153 | 0.1067 | 0.27142 | 0.46797 | 0.8165
| Rel.Std.Dev. |  | 0% | 46.69% | 2.776% | 0% | 4.834% | 0.6679% | 1.698% | 2.607% | 9.421%
| PDM3 |  | 11.64% | 14.82% | -4.884% | -1.491% | -25.08% | 4.908% | 4.967% | 17.88% | -43.08%
| Z-Score (Absolute) |  | 1.37 | 0.58 | 0.45 | 0 | 2.46 | 0.67 | 0.67 | 2.02 | 4.34
| Analytical results for Nd in OREAS 922 (Certified Value 38.9 ppm) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 39.1 | 39.1 | 38.8 | 36 | 39.606 | 42.1 | 41.4 | 37.4 | [6 empty] | 1
| 1 | 2 | 38.9 | 38.2 | 37.6 | 35.5 | 37.914 | 41.3 | 39.8 | 38.2 | [6 empty] | #N/A
| 1 | 3 | 39.3 | 38.5 | 39.3 | 35 | 37.795 | 42.2 | 40.6 | 37.8 | [6 empty] | 16
| 1 | 4 | 39.4 | 40.4 | 39.5 | 35 | 38.293 | 41.9 | 38.8 | 38.4 | [6 empty] | 38.8761
| 1 | 5 | 39.4 | 39.3 | 38.5 | 36 | 39.353 | 41.9 | 38.7 | 38.4 | [6 empty] | 160
| 1 | 6 | 39.3 | 36.3 | 38.3 | 35 | 37.98 | 41.5 | 41.7 | 38.2
| Mean |  | 39.2333 | 38.6333 | 38.6667 | 35.4167 | 38.4902 | 41.8167 | 40.1667 | 38.0667
| Median |  | 39.3 | 38.8 | 38.65 | 35.25 | 38.1365 | 41.9 | 40.2 | 38.2
| Std Dev. |  | 0.19664 | 1.3736 | 0.69474 | 0.4916 | 0.7879 | 0.34881 | 1.2817 | 0.39328
| Rel.Std.Dev. |  | 0.5012% | 3.555% | 1.797% | 1.388% | 2.047% | 0.8341% | 3.191% | 1.033%
| PDM3 |  | 0.9188% | -0.6246% | -0.5388% | -8.899% | -0.9928% | 7.564% | 3.32% | -2.082%
| Z-Score (Absolute) |  | 0.67 | 0.02 | 0.02 | 3.74 | 0.18 | 3.66 | 1.75 | 0.67
| Analytical results for Ni in OREAS 922 (Certified Value 43.4 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 18 | 20 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 55 | 30 | 40 | < 50 | 37.6 | < 50 | 41.708 | 50 | < 50 | 42 |  |  |  |  | 1
| 1 | 2 | 54 | 30 | 30 | 50 | 36 | < 50 | 41.692 | 40 | 56.7 | 41 |  |  |  |  | #N/A
| 1 | 3 | 55 | 40 | 30 | < 50 | 37.8 | < 50 | 41.144 | 30 | 51.8 | 41 |  |  |  |  | 16
| 1 | 4 | 55 | 30 | 40 | 50 | 37.44 | < 50 | 42.388 | 30 | < 50 | 41 |  |  |  |  | 43.3695
| 1 | 5 | 55 | 40 | 50 | 50 | 34 | < 50 | 41.959 | 50 | < 50 | 41 |  |  |  |  | 161
| 1 | 6 | 56 | 40 | 40 | 50 | 35.7 | < 50 | 41.824 | 30 | < 50 | 45
| Mean |  | 55 | 35 | 38.3333 | 50 | 36.4233 |  | 41.7858 | 38.3333 | 54.25 | 41.8333
| Median |  | 55 | 35 | 40 | 50 | 36.72 |  | 41.766 | 35 | 54.25 | 41
| Std Dev. |  | 0.63246 | 5.4772 | 7.5277 | 0 | 1.4757 |  | 0.40518 | 9.8319 | 3.4648 | 1.6021
| Rel.Std.Dev. |  | 1.15% | 15.65% | 19.64% | 0% | 4.052% |  | 0.9697% | 25.65% | 6.387% | 3.83%
| PDM3 |  | 26.82% | -19.3% | -11.61% | 15.29% | -16.02% |  | -3.652% | -11.61% | 25.09% | -3.542%
| Z-Score (Absolute) |  | 1.95 | 0.88 | 0.4 | 1.24 | 0.67 | NA | 0.08 | 0.4 | 1.84 | 0
| Analytical results for P in OREAS 922 (Certified Value 0.066 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 20 | 21 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.07 | 0.066 | 0.069 | 0.08 | 0.066 | 0.062 | 0.0689 | 0.07 | [6 empty] | 1
| 1 | 2 | 0.06 | 0.065 | 0.069 | 0.08 | 0.069 | 0.058 | 0.0677 | 0.06 | [6 empty] | #N/A
| 1 | 3 | 0.07 | 0.064 | 0.071 | 0.08 | 0.074 | 0.058 | 0.0719 | 0.06 | [6 empty] | 16
| 1 | 4 | 0.06 | 0.062 | 0.07 | 0.08 | 0.074 | 0.059 | 0.072 | 0.06 | [6 empty] | 0.065664
| 1 | 5 | 0.07 | 0.061 | 0.068 | 0.08 | 0.069 | 0.06 | 0.0671 | 0.07 | [6 empty] | 162
| 1 | 6 | 0.06 | 0.06 | 0.068 | 0.08 | 0.062 | 0.062 | 0.0643 | 0.07
| Mean |  | 0.065 | 0.063 | 0.069167 | 0.08 | 0.069 | 0.059833 | 0.06865 | 0.065
| Median |  | 0.065 | 0.063 | 0.069 | 0.08 | 0.069 | 0.0595 | 0.0683 | 0.065
| Std Dev. |  | 0.0054772 | 0.0023664 | 0.001169 | 0 | 0.0046476 | 0.0018348 | 0.002969 | 0.0054772
| Rel.Std.Dev. |  | 8.427% | 3.756% | 1.69% | 0% | 6.736% | 3.067% | 4.325% | 8.427%
| PDM3 |  | -1.012% | -4.057% | 5.334% | 21.83% | 5.08% | -8.88% | 4.547% | -1.012%
| Z-Score (Absolute) |  | 0.54 | 1.14 | 0.7 | 3.93 | 0.65 | 2.09 | 0.54 | 0.54
| Analytical results for Pb in OREAS 922 (Certified Value 64 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 15 | 16 | 18 | 20 | 21 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*AAS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  | 0
1
| 1 | 1 | 92 | 58 | 59.7 | 60 | 66 | 69.6975 | 100 | 65.116 | 72.1 | 112 | 60 |  |  |  | 1
| 1 | 2 | 91 | 52.8 | 65 | 60 | 69 | 68.1455 | 100 | 60.704 | 65.3 | 141 | 62 |  |  |  | #N/A
| 1 | 3 | 92 | 54.3 | 63.9 | 60 | 63 | 62.079 | 100 | 70.953 | 75.9 | 146 | 62 |  |  |  | 16
| 1 | 4 | 92 | 117 | 57.2 | 60 | 66 | 63.8347 | 100 | 65.54 | 77.2 | < 100 | 65 |  |  |  | 63.6088
| 1 | 5 | 93 | 77 | 57.8 | 60 | 67 | 65.651 | 100 | 61.088 | 68.6 | 131 | 62 |  |  |  | 163
| 1 | 6 | 92 | 56.4 | 59.7 | 60 | 66 | 60.5197 | 100 | 61.096 | 65.2 | 168 | 65
| Mean |  | 92 | 69.25 | 60.55 | 60 | 66.1667 | 64.9879 | 100 | 64.0828 | 70.7167 | 139.6 | 62.6667
| Median |  | 92 | 57.2 | 59.7 | 60 | 66 | 64.7429 | 100 | 63.106 | 70.35 | 141 | 62
| Std Dev. |  | 0.63246 | 25.0037 | 3.2017 | 0 | 1.9408 | 3.5312 | 0 | 3.9924 | 5.1975 | 20.5256 | 1.9664
| Rel.Std.Dev. |  | 0.6875% | 36.11% | 5.288% | 0% | 2.933% | 5.434% | 0% | 6.23% | 7.35% | 14.7% | 3.138%
| PDM3 |  | 44.63% | 8.869% | -4.809% | -5.673% | 4.021% | 2.168% | 57.21% | 0.7452% | 11.17% | 119.5% | -1.481%
| Z-Score (Absolute) |  | 3.65 | 0.71 | 0.6 | 0.67 | 0.16 | 0 | 4.73 | 0.12 | 0.77 | 10.09 | 0.31
| Analytical results for Pr in OREAS 922 (Certified Value 10.6 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 10.8 | 10.8 | 10.7 | 10 | 10.667 | 11.2 | 10.1 | 11 | [6 empty] | 1
| 1 | 2 | 10.7 | 10.5 | 10.4 | 10.4 | 10.268 | 11.1 | 10.1 | 10.9 | [6 empty] | #N/A
| 1 | 3 | 10.9 | 10.7 | 10.8 | 10.2 | 10.126 | 11.1 | 10.7 | 10.9 | [6 empty] | 16
| 1 | 4 | 10.9 | 11.1 | 10.7 | 10.2 | 10.381 | 11 | 10.6 | 11 | [6 empty] | 10.6446
| 1 | 5 | 10.9 | 10.8 | 10.5 | 10.4 | 10.667 | 11.1 | 10.4 | 11 | [6 empty] | 164
| 1 | 6 | 10.9 | 10.2 | 10.5 | 10 | 10.033 | 11.1 | 10.4 | 11.1
| Mean |  | 10.85 | 10.6833 | 10.6 | 10.2 | 10.357 | 11.1 | 10.3833 | 10.9833
| Median |  | 10.9 | 10.75 | 10.6 | 10.2 | 10.3245 | 11.1 | 10.4 | 11
| Std Dev. |  | 0.083666 | 0.30605 | 0.15492 | 0.17889 | 0.26797 | 0.063246 | 0.24833 | 0.075277
| Rel.Std.Dev. |  | 0.7711% | 2.865% | 1.462% | 1.754% | 2.587% | 0.5698% | 2.392% | 0.6854%
| PDM3 |  | 1.929% | 0.3636% | -0.4192% | -4.177% | -2.702% | 4.278% | -2.455% | 3.182%
| Z-Score (Absolute) |  | 0.52 | 0.1 | 0.1 | 1.1 | 0.71 | 1.14 | 0.64 | 0.85
| Analytical results for Rb in OREAS 922 (Certified Value 167 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 169 | 166 | 167 | 211 | 171.18 | 175 | 172.7 | 164 | [6 empty] | 1
| 1 | 2 | 168 | 166 | 159 | 219 | 169.54 | 174 | 167.3 | 163 | [6 empty] | #N/A
| 1 | 3 | 169 | 165 | 164 | 214 | 169.36 | 179 | 171.5 | 164 | [6 empty] | 16
| 1 | 4 | 170 | 167 | 159 | 226 | 166.64 | 178 | 164.7 | 163 | [6 empty] | 167.485
| 1 | 5 | 170 | 165 | 157 | 220 | 169.67 | 178 | 163.2 | 163 | [6 empty] | 165
| 1 | 6 | 170 | 165 | 157 | 214 | 168.58 | 175 | 166 | 165
| Mean |  | 169.3333 | 165.6667 | 160.5 | 217.3333 | 169.1617 | 176.5 | 167.5667 | 163.6667
| Median |  | 169.5 | 165.5 | 159 | 216.5 | 169.45 | 176.5 | 166.65 | 163.5
| Std Dev. |  | 0.8165 | 0.8165 | 4.0866 | 5.4283 | 1.4974 | 2.0736 | 3.7851 | 0.8165
| Rel.Std.Dev. |  | 0.4822% | 0.4929% | 2.546% | 2.498% | 0.8852% | 1.175% | 2.259% | 0.4989%
| PDM3 |  | 1.104% | -1.086% | -4.171% | 29.76% | 1.001% | 5.383% | 0.04876% | -2.28%
| Z-Score (Absolute) |  | 0.18 | 0.49 | 1.43 | 8.93 | 0.15 | 1.48 | 0.15 | 0.86
| Analytical results for Re in OREAS 922 (Indicative Value < 0.1 ppm) | [15 empty] | Indicative
Re | Round | Replicate | 02: Lab; 04: Lab |  |  |  | 1
| No | No | 02: 09; 04: 19 |  |  |  | ppm
|  |  | 02: PF*MS; 04: PF*MS |  |  |  | 2
2
| 1 | 1 | 02: < 1; 04: < 0.1 |  |  |  | 1
| 1 | 2 | 02: < 1; 04: < 0.1 |  |  |  | 8
| 1 | 3 | 02: < 1; 04: < 0.1 |  |  |  | 16
| 1 | 4 | 02: < 1; 04: < 0.1 |  |  |  | < 0.1
| 1 | 5 | 02: < 1; 04: < 0.1 |  |  |  | 14
| 1 | 6 | 02: < 1; 04: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02 to 04
| Analytical results for S in OREAS 922 (Certified Value 0.389 wt.%) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 04 | 05 | 09 | 10 | 13 | 17 | 18 | 20 | [6 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | [6 empty] | 3
3
| 1 | 1 | 0.39 | 0.43 | 0.42 | 0.4 | 0.386 | 0.62896 | 0.38 | 0.547 | [6 empty] | 1
| 1 | 2 | 0.38 | 0.39 | 0.42 | 0.38 | 0.375 | 0.62044 | 0.38 | 0.603 | [6 empty] | 6
| 1 | 3 | 0.45 | 0.37 | 0.4 | 0.37 | 0.401 | 0.67777 | 0.41 | 0.526 | [6 empty] | 16
| 1 | 4 | 0.38 | 0.36 | 0.4 | 0.4 | 0.409 | 0.89517 | 0.4 | 0.549 | [6 empty] | 0.38894
| 1 | 5 | 0.38 | 0.35 | 0.4 | 0.38 | 0.377 | 0.80352 | 0.38 | 0.529 | [6 empty] | 166
| 1 | 6 | 0.37 | 0.38 | 0.4 | 0.39 | 0.394 | 0.55862 | 0.39 | 0.527
| Mean |  | 0.39167 | 0.38 | 0.40667 | 0.38667 | 0.39033 | 0.69741 | 0.39 | 0.54683
| Median |  | 0.38 | 0.375 | 0.4 | 0.385 | 0.39 | 0.65336 | 0.385 | 0.538
| Std Dev. |  | 0.029269 | 0.028284 | 0.010328 | 0.012111 | 0.013471 | 0.12699 | 0.012649 | 0.029342
| Rel.Std.Dev. |  | 7.473% | 7.443% | 2.54% | 3.132% | 3.451% | 18.21% | 3.243% | 5.366%
| PDM3 |  | 0.6999% | -2.3% | 4.556% | -0.5856% | 0.3571% | 79.31% | 0.2714% | 40.59%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 1.09 | 0.23 | 0.01 | 20.38 | 0.01 | 9.65
| Analytical results for Sb in OREAS 922 (Certified Value < 2 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 09 | 10 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | < 2 | < 2 | 2 | 1.6 | 2.4014 | 1.526 | < 2 | 0.71156 | 0.9 | [5 empty] | 1
| 1 | 2 | < 2 | < 2 | 2 | 1.7 | 2.1745 | 1.503 | < 2 | 0.90281 | 1.1 | [5 empty] | #N/A
| 1 | 3 | < 2 | < 2 | 2 | 1.6 | 2.1295 | 1.397 | < 2 | 0.77719 | 1.1 | [5 empty] | 16
| 1 | 4 | < 2 | < 2 | 2 | 1.7 | 1.7236 | 1.513 | < 2 | 1.0688 | 0.9 | [5 empty] | < 2
| 1 | 5 | < 2 | < 2 | 2 | 1.6 | 1.8693 | 1.472 | < 2 | 1.0603 | 0.8 | [5 empty] | 167
| 1 | 6 | < 2 | < 2 | 2 | 1.4 | 2.1873 | 1.54 | < 2 | 1.0753 | 1
| Mean |  | 05: 2; 13: 1.6; 15: 2.081; 16: 1.4918; 18: 0.93266; 19: 0.96667
| Median |  | 05: 2; 13: 1.6; 15: 2.152; 16: 1.508; 18: 0.98156; 19: 0.95
| Std Dev. |  | 05: 0; 13: 0.10954; 15: 0.24394; 16: 0.051859; 18: 0.16069; 19: 0.12111
| Rel.Std.Dev. |  | 05: 0%; 13: 6.847%; 15: 11.72%; 16: 3.476%; 18: 17.23%; 19: 12.53%
PDM3
| Z-Score (Absolute) |  | NA | NA | 0.62 | 0.07 | 0.73 | 0.07 | NA | 0.84 | 0.79
| Analytical results for Sc in OREAS 922 (Indicative Value 12 ppm) | [15 empty] | Indicative
Sc | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab; 15: Lab; 16: Lab |  |  |  | 1
| No | No | 02: 09; 04: 13; 05: 16; 13: 19; 15: 20; 16: 21 |  |  |  | ppm
|  |  | 02: PF*OES; 04: PF*OES/MS; 05: PF*OES/MS; 13: PF*OES; 15: PF*OES; 16: PF*OES/MS |  |  |  | 1
1
| 1 | 1 | 02: 10; 04: < 2; 05: 12.963; 13: 13; 15: < 20; 16: 12 |  |  |  | 1
| 1 | 2 | 02: 10; 04: < 2; 05: 12.895; 13: 13; 15: < 20; 16: 12 |  |  |  | 9
| 1 | 3 | 02: 10; 04: < 2; 05: 12.656; 13: 13; 15: < 20; 16: 12 |  |  |  | 16
| 1 | 4 | 02: 10; 04: < 2; 05: 12.793; 13: 13; 15: < 20; 16: 12 |  |  |  | 11.9592
| 1 | 5 | 02: 10; 04: < 2; 05: 13.019; 13: 13; 15: < 20; 16: 13 |  |  |  | 15
| 1 | 6 | 02: 10; 04: < 2; 05: 12.695; 13: 13; 15: < 20; 16: 12
| Mean |  | 02: 10; 05: 12.8368; 13: 13; 16: 12.1667
| Median |  | 02: 10; 05: 12.844; 13: 13; 16: 12
| Std Dev. |  | 02: 0; 05: 0.14644; 13: 0; 16: 0.40825
| Rel.Std.Dev. |  | 02: 0%; 05: 1.141%; 13: 0%; 16: 3.355%
| PDM3 |  | 02: -16.38%; 05: 7.338%; 13: 8.703%; 16: 1.735%
| Z-Score (Absolute) |  | 02: 3.26; 04: NA; 05: 0.56; 13: 0.78; 15: NA; 16: 0.56
| Analytical results for Se in OREAS 922 (Indicative Value < 20 ppm) | [15 empty] | Indicative
Se | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab; 15: Lab; 16: Lab |  |  |  | 1
| No | No | 02: 04; 04: 05; 05: 09; 13: 10; 15: 13; 16: 18 |  |  |  | ppm
|  |  | 02: PF*MS; 04: PF*MS; 05: PF*MS; 13: PF*MS; 15: PF*MS; 16: PF*MS |  |  |  | 1
1
| 1 | 1 | 02: < 0.8; 04: 14.8; 05: < 20; 13: < 20; 15: < 20; 16: < 0.8 |  |  |  | 1
| 1 | 2 | 02: < 0.8; 04: 14.6; 05: < 20; 13: < 20; 15: < 20; 16: < 0.8 |  |  |  | 10
| 1 | 3 | 02: < 0.8; 04: 8.2; 05: < 20; 13: < 20; 15: < 20; 16: < 0.8 |  |  |  | 16
| 1 | 4 | 02: < 0.8; 04: 4; 05: < 20; 13: < 20; 15: < 20; 16: < 0.8 |  |  |  | < 20
| 1 | 5 | 02: < 0.8; 04: 4.3; 05: < 20; 13: < 20; 15: < 20; 16: 1.5 |  |  |  | 16
| 1 | 6 | 02: < 0.8; 04: 12.4; 05: < 20; 13: < 20; 15: < 20; 16: < 0.8
| Mean |  | 04: 9.7167; 16: 1.5
| Median |  | 04: 10.3; 16: 1.5
| Std Dev. |  | 04: 4.9236
| Rel.Std.Dev. |  | 04: 50.67%
PDM3
| Z-Score (Absolute) |  | 02: NA; 04: 0.67; 05: NA; 13: NA; 15: NA; 16: 0.67
| Analytical results for Si in OREAS 922 (Certified Value 30.51 wt.%) | [15 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 04 | 05 | 09 | 13 | 15 | 18 | 20 | [7 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 32.2 | 30.7 | 30.2 | 31.5 | NA | 31.6 | 30.3 | [7 empty] | 1
| 1 | 2 | 31.8 | 30.4 | 30 | 32.5 | NA | 31 | 31.2 | [7 empty] | #N/A
| 1 | 3 | 31.3 | 29.9 | 30.2 | 32.2 | NA | 31.5 | 30.7 | [7 empty] | 16
| 1 | 4 | 29 | 29.8 | 30.5 | 32.4 | NA | 33.1 | 30 | [7 empty] | 30.5133
| 1 | 5 | 29.3 | 30.2 | 30.8 | 33.2 | NA | 30.1 | 30.2 | [7 empty] | 168
| 1 | 6 | 29.3 | 29.5 | 30.4 | 33.4 | NA | 30.6 | 29.6
| Mean |  | 02: 30.4833; 04: 30.0833; 05: 30.35; 13: 32.5333; 16: 31.3167; 17: 30.3333
| Median |  | 02: 30.3; 04: 30.05; 05: 30.3; 13: 32.45; 16: 31.25; 17: 30.25
| Std Dev. |  | 02: 1.4386; 04: 0.43551; 05: 0.28107; 13: 0.69186; 16: 1.0381; 17: 0.55737
| Rel.Std.Dev. |  | 02: 4.719%; 04: 1.448%; 05: 0.9261%; 13: 2.127%; 16: 3.315%; 17: 1.837%
| PDM3 |  | 02: -0.09832%; 04: -1.409%; 05: -0.5353%; 13: 6.62%; 16: 2.633%; 17: -0.5899%
| Z-Score (Absolute) |  | 0.22 | 1.08 | 0.22 | 6.85 | NA | 2.91 | 0.27
| Analytical results for Sm in OREAS 922 (Certified Value 7.31 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 7.5 | 7.8 | 7.5 | 7 | 7.16 | 7.6 | 7.3 | 7 | [6 empty] | 1
| 1 | 2 | 7.6 | 7.7 | 7.4 | 7 | 6.9 | 7.4 | 6.7 | 7.1 | [6 empty] | #N/A
| 1 | 3 | 7.6 | 7.6 | 7.6 | 7 | 7.07 | 7.4 | 8 | 7.1 | [6 empty] | 16
| 1 | 4 | 7.6 | 7.7 | 7.5 | 7 | 7.05 | 7.5 | 7.2 | 7.2 | [6 empty] | 7.3112
| 1 | 5 | 7.6 | 7.8 | 7.4 | 7.5 | 7.28 | 7.3 | 7.3 | 6.9 | [6 empty] | 169
| 1 | 6 | 7.6 | 7.2 | 7.3 | 7 | 6.94 | 7.5 | 7.2 | 7.2
| Mean |  | 7.5833 | 7.6333 | 7.45 | 7.0833 | 7.0667 | 7.45 | 7.2833 | 7.0833
| Median |  | 7.6 | 7.7 | 7.45 | 7 | 7.06 | 7.45 | 7.25 | 7.1
| Std Dev. |  | 0.040825 | 0.22509 | 0.10488 | 0.20412 | 0.14024 | 0.10488 | 0.41673 | 0.1169
| Rel.Std.Dev. |  | 0.5383% | 2.949% | 1.408% | 2.882% | 1.984% | 1.408% | 5.722% | 1.65%
| PDM3 |  | 3.721% | 4.405% | 1.898% | -3.117% | -3.345% | 1.898% | -0.3818% | -3.117%
| Z-Score (Absolute) |  | 0.92 | 1.08 | 0.49 | 0.67 | 0.73 | 0.49 | 0.49 | 0.67
| Analytical results for Sn in OREAS 922 (Certified Value 10 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 16 | 18 | 19 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 14 | 8.1 | 9.5 | 10 | < 100 | < 100 | 11.526 | 7.7 | 11 | 9 |  |  |  |  | 1
| 1 | 2 | 14 | 7 | 9.1 | 10 | < 100 | < 100 | 11.828 | 11 | 10 | 9 |  |  |  |  | #N/A
| 1 | 3 | 14 | 7.2 | 9.9 | 10 | < 100 | < 100 | 11.821 | 11.6 | 11 | 12 |  |  |  |  | 16
| 1 | 4 | 14 | 7.6 | 12.1 | 10 | < 100 | < 100 | 13.201 | 11.4 | 9 | 9 |  |  |  |  | 10.018
| 1 | 5 | 14 | 7.4 | 8.7 | 10 | < 100 | < 100 | 12.998 | 11.3 | 10 | 10 |  |  |  |  | 170
| 1 | 6 | 14 | 7 | 9.4 | 10 | < 100 | < 100 | 11.841 | 10.8 | 9 | 11
| Mean |  | 14 | 7.3833 | 9.7833 | 10 |  |  | 12.2025 | 10.6333 | 10 | 10
| Median |  | 14 | 7.3 | 9.45 | 10 |  |  | 11.8345 | 11.15 | 10 | 9.5
| Std Dev. |  | 0 | 0.42151 | 1.204 | 0 |  |  | 0.70767 | 1.4652 | 0.89443 | 1.2649
| Rel.Std.Dev. |  | 0% | 5.709% | 12.31% | 0% |  |  | 5.799% | 13.78% | 8.944% | 12.65%
| PDM3 |  | 39.75% | -26.3% | -2.342% | -0.1794% |  |  | 21.81% | 6.143% | -0.1794% | -0.1794%
| Z-Score (Absolute) |  | 2.84 | 1.86 | 0.48 | 0 | NA | NA | 1.56 | 0.87 | 0 | 0
| Analytical results for Sr in OREAS 922 (Certified Value 58 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 20 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 0
1
| 1 | 1 | 57 | 59 | 62 | 65 | 62.425 | 54 | 57 | 59.4 | 70 | [5 empty] | 1
| 1 | 2 | 56 | 56 | 44 | 67 | 60.156 | 54 | 59 | 58.7 | 70 | [5 empty] | #N/A
| 1 | 3 | 56 | 58 | 47 | 59 | 62.06 | 57 | 59 | 57.2 | 70 | [5 empty] | 16
| 1 | 4 | 57 | 62 | 46 | 58 | 61.106 | 79 | 59 | 59.3 | 70 | [5 empty] | 58.3976
| 1 | 5 | 57 | 58 | 44 | 65 | 63.973 | 55 | 57 | 53.1 | 70 | [5 empty] | 171
| 1 | 6 | 54 | 56 | 42 | 59 | 63.179 | 56 | 56 | 54.9 | 70
| Mean |  | 56.1667 | 58.1667 | 47.5 | 62.1667 | 62.1498 | 59.1667 | 57.8333 | 57.1 | 70
| Median |  | 56.5 | 58 | 45 | 62 | 62.2425 | 55.5 | 58 | 57.95 | 70
| Std Dev. |  | 1.169 | 2.2286 | 7.3144 | 3.92 | 1.3807 | 9.786 | 1.3292 | 2.59 | 0
| Rel.Std.Dev. |  | 2.081% | 3.831% | 15.4% | 6.306% | 2.222% | 16.54% | 2.298% | 4.536% | 0%
| PDM3 |  | -3.82% | -0.3954% | -18.66% | 6.454% | 6.425% | 1.317% | -0.9662% | -2.222% | 19.87%
| Z-Score (Absolute) |  | 0.43 | 0.09 | 3.39 | 1.11 | 1.11 | 0.67 | 0 | 0.19 | 3.12
| Analytical results for Ta in OREAS 922 (Certified Value 1.33 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 8.2 | 1.3 | 1 | 1.5 | 1.6 | 1.329 | 1.2 | 1.5 | 1.1 | [5 empty] | 1
| 1 | 2 | 8.3 | 1.3 | 0.8 | 2 | 1.2895 | 1.372 | 1.1 | 1.5 | 1 | [5 empty] | #N/A
| 1 | 3 | 8.3 | 1.3 | 0.8 | 1.5 | 1.1951 | 1.338 | 1.2 | 1.5 | 1.1 | [5 empty] | 16
| 1 | 4 | 8.2 | 1.3 | 0.7 | 1.5 | 1.2788 | 1.322 | 1.2 | 1.4 | 1.1 | [5 empty] | 1.3276
| 1 | 5 | 8.2 | 1.4 | 0.9 | 1.5 | 1.5554 | 1.371 | 1.2 | 1.4 | 1 | [5 empty] | 172
| 1 | 6 | 8.2 | 11.4 | 0.7 | 1.5 | 1.5562 | 1.332 | 1.2 | 1.4 | 1.2
| Mean |  | 8.2333 | 3 | 0.81667 | 1.5833 | 1.4125 | 1.344 | 1.1833 | 1.45 | 1.0833
| Median |  | 8.2 | 1.3 | 0.8 | 1.5 | 1.4224 | 1.335 | 1.2 | 1.45 | 1.1
| Std Dev. |  | 0.05164 | 4.1153 | 0.1169 | 0.20412 | 0.17692 | 0.021918 | 0.040825 | 0.054772 | 0.075277
| Rel.Std.Dev. |  | 0.6272% | 137.2% | 14.31% | 12.89% | 12.53% | 1.631% | 3.45% | 3.777% | 6.949%
| PDM3 |  | 520.2% | 126% | -38.49% | 19.26% | 6.395% | 1.236% | -10.87% | 9.22% | -18.4%
| Z-Score (Absolute) |  | 29.78 | 0.1 | 2.28 | 0.67 | 0.3 | 0 | 0.69 | 0.46 | 1.13
| Analytical results for Tb in OREAS 922 (Certified Value 1.02 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.11 | 1 | 1 | 1 | 0.951 | 1.1 | 1 | 1 | [6 empty] | 1
| 1 | 2 | 1.11 | 1 | 1 | 1 | 0.943 | 1.1 | 0.9 | 1.07 | [6 empty] | #N/A
| 1 | 3 | 1.12 | 1.1 | 1 | 1 | 0.958 | 1.1 | 0.9 | 1.04 | [6 empty] | 16
| 1 | 4 | 1.13 | 1 | 1 | 1 | 0.95 | 1.1 | 0.9 | 1.07 | [6 empty] | 1.0234
| 1 | 5 | 1.12 | 1.1 | 1 | 1 | 0.98 | 1.2 | 0.8 | 1.05 | [6 empty] | 173
| 1 | 6 | 1.12 | 1 | 1 | 1 | 0.952 | 1.1 | 1 | 1.01
| Mean |  | 1.1183 | 1.0333 | 1 | 1 | 0.95567 | 1.1167 | 0.91667 | 1.04
| Median |  | 1.12 | 1 | 1 | 1 | 0.9515 | 1.1 | 0.9 | 1.045
| Std Dev. |  | 0.0075277 | 0.05164 | 0 | 0 | 0.012848 | 0.040825 | 0.075277 | 0.029665
| Rel.Std.Dev. |  | 0.6731% | 4.997% | 0% | 0% | 1.344% | 3.656% | 8.212% | 2.852%
| PDM3 |  | 9.274% | 0.969% | -2.288% | -2.288% | -6.62% | 9.112% | -10.43% | 1.62%
| Z-Score (Absolute) |  | 1.63 | 0.27 | 0.27 | 0.27 | 0.98 | 1.33 | 1.23 | 0.37
| Analytical results for Te in OREAS 922 (Indicative Value < 6 ppm) | [15 empty] | Indicative
Te | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab; 15: Lab; 16: Lab |  |  |  | 1
| No | No | 02: 04; 04: 05; 05: 09; 13: 13; 15: 18; 16: 19 |  |  |  | ppm
|  |  | 02: PF*MS; 04: PF*MS; 05: PF*MS; 13: PF*MS; 15: PF*MS; 16: PF*MS |  |  |  | 2
2
| 1 | 1 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: < 0.1 |  |  |  | 1
| 1 | 2 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: < 0.1 |  |  |  | 8
| 1 | 3 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: 0.2203 |  |  |  | 16
| 1 | 4 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: < 0.1 |  |  |  | < 6
| 1 | 5 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: < 0.1 |  |  |  | 14
| 1 | 6 | 02: < 6; 04: < 6; 05: < 50; 13: < 1; 15: < 6; 16: < 0.1
| Mean |  | 16: 0.2203
| Median |  | 16: 0.2203
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02 to 16
| Analytical results for Th in OREAS 922 (Certified Value 17.7 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 1
2
| 1 | 1 | 20.6 | 17.5 | 18.4 | 15.5 | 14.9949 | 18.991 | 18 | 17.5 | 16.8 | [5 empty] | 1
| 1 | 2 | 20.5 | 17.2 | 17.6 | 16 | 14.494 | 18.321 | 17.9 | 17.65 | 17.1 | [5 empty] | #N/A
| 1 | 3 | 20.6 | 17.2 | 18.6 | 16 | 14.0598 | 18.058 | 18 | 18 | 17.9 | [5 empty] | 16
| 1 | 4 | 20.7 | 18.1 | 18.4 | 16.5 | 14.3698 | 18.918 | 18.3 | 18.01 | 17.8 | [5 empty] | 17.6598
| 1 | 5 | 20.7 | 17.8 | 18 | 17 | 14.7303 | 18.92 | 18.3 | 17.83 | 17.9 | [5 empty] | 174
| 1 | 6 | 20.6 | 16.9 | 17.7 | 16.5 | 14.9883 | 17.945 | 17.7 | 17.67 | 17.3
| Mean |  | 20.6167 | 17.45 | 18.1167 | 16.25 | 14.6062 | 18.5255 | 18.0333 | 17.7767 | 17.4667
| Median |  | 20.6 | 17.35 | 18.2 | 16.25 | 14.6121 | 18.6195 | 18 | 17.75 | 17.55
| Std Dev. |  | 0.075277 | 0.44159 | 0.41191 | 0.5244 | 0.36864 | 0.47407 | 0.23381 | 0.20549 | 0.46762
| Rel.Std.Dev. |  | 0.3651% | 2.531% | 2.274% | 3.227% | 2.524% | 2.559% | 1.297% | 1.156% | 2.677%
| PDM3 |  | 16.74% | -1.188% | 2.587% | -7.983% | -17.29% | 4.902% | 2.115% | 0.6616% | -1.094%
| Z-Score (Absolute) |  | 5.63 | 0.65 | 0.67 | 3.03 | 6.29 | 1.49 | 0.51 | 0 | 0.61
| Analytical results for Ti in OREAS 922 (Certified Value 0.439 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 21 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 3
3
| 1 | 1 | 0.43 | 0.47 | 0.43 | 0.44 | 0.4489 | 0.4 | 0.434 | 0.51042 | 0.45 | 0.45 | 0.45 | 0.4 |  |  | 1
| 1 | 2 | 0.44 | 0.46 | 0.43 | 0.44 | 0.4472 | 0.42 | 0.43 | 0.46686 | 0.45 | 0.45 | 0.446 | 0.4 |  |  | #N/A
| 1 | 3 | 0.45 | 0.45 | 0.43 | 0.44 | 0.458 | 0.42 | 0.426 | 0.48076 | 0.44 | 0.45 | 0.444 | 0.41 |  |  | 16
| 1 | 4 | 0.43 | 0.44 | 0.43 | 0.45 | 0.4526 | 0.4 | 0.424 | 0.46476 | 0.45 | 0.47 | 0.447 | 0.4 |  |  | 0.43924
| 1 | 5 | 0.44 | 0.45 | 0.43 | 0.45 | 0.4496 | 0.41 | 0.427 | 0.46978 | 0.43 | 0.45 | 0.425 | 0.41 |  |  | 175
| 1 | 6 | 0.44 | 0.45 | 0.42 | 0.45 | 0.4658 | 0.4 | 0.424 | 0.46791 | 0.45 | 0.46 | 0.436 | 0.41
| Mean |  | 0.43833 | 0.45333 | 0.42833 | 0.445 | 0.45368 | 0.40833 | 0.4275 | 0.47675 | 0.445 | 0.455 | 0.44133 | 0.405
| Median |  | 0.44 | 0.45 | 0.43 | 0.445 | 0.4511 | 0.405 | 0.4265 | 0.46885 | 0.45 | 0.45 | 0.445 | 0.405
| Std Dev. |  | 0.0075277 | 0.010328 | 0.0040825 | 0.0054772 | 0.0070457 | 0.0098319 | 0.0038859 | 0.017423 | 0.0083666 | 0.0083666 | 0.009288 | 0.0054772
| Rel.Std.Dev. |  | 1.717% | 2.278% | 0.9531% | 1.231% | 1.553% | 2.408% | 0.909% | 3.655% | 1.88% | 1.839% | 2.105% | 1.352%
| PDM3 |  | -0.2061% | 3.209% | -2.483% | 1.312% | 3.289% | -7.036% | -2.673% | 8.54% | 1.312% | 3.588% | 0.4769% | -7.795%
| Z-Score (Absolute) |  | 0.29 | 0.61 | 0.9 | 0.11 | 0.63 | 2.1 | 0.95 | 1.62 | 0.11 | 0.71 | 0.11 | 2.3
| Analytical results for Tl in OREAS 922 (Certified Value 0.88 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.2 | 0.9 | 0.9 | < 2 | 0.79801 | 0.929 | 0.9 | 0.8207 | 0.9 | [5 empty] | 1
| 1 | 2 | 1.2 | 0.9 | 0.9 | < 2 | 0.86653 | 0.939 | 0.9 | 0.7879 | 0.9 | [5 empty] | #N/A
| 1 | 3 | 1.2 | 0.9 | 0.9 | < 2 | 0.73696 | 0.91 | 0.9 | 0.8188 | 0.9 | [5 empty] | 16
| 1 | 4 | 1.2 | 0.9 | 0.9 | < 2 | 0.80568 | 0.911 | 0.9 | 0.9059 | 0.9 | [5 empty] | 0.88234
| 1 | 5 | 1.2 | 1 | 0.9 | < 2 | 0.98774 | 0.916 | 0.9 | 0.7407 | 0.9 | [5 empty] | 176
| 1 | 6 | 1.2 | 0.9 | 0.9 | < 2 | 0.78863 | 0.921 | 0.9 | 0.8748 | 0.8
| Mean |  | 1.2 | 0.91667 | 0.9 |  | 0.83059 | 0.921 | 0.9 | 0.8248 | 0.88333
| Median |  | 1.2 | 0.9 | 0.9 |  | 0.80184 | 0.9185 | 0.9 | 0.81975 | 0.9
| Std Dev. |  | 0 | 0.040825 | 0 |  | 0.087403 | 0.011261 | 0 | 0.05927 | 0.040825
| Rel.Std.Dev. |  | 0% | 4.454% | 0% |  | 10.52% | 1.223% | 0% | 7.186% | 4.622%
| PDM3 |  | 36% | 3.89% | 2.001% |  | -5.865% | 4.381% | 2.001% | -6.521% | 0.1124%
| Z-Score (Absolute) |  | 10.74 | 0.6 | 0 | NA | 2.49 | 0.75 | 0 | 2.69 | 0.6
| Analytical results for Tm in OREAS 922 (Certified Value 0.51 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 0.54 | 0.5 | 0.5 | 0.6 | 0.449 | 0.6 | 0.4 | 0.48 | [6 empty] | 1
| 1 | 2 | 0.55 | 0.5 | 0.5 | 0.6 | 0.491 | 0.6 | 0.4 | 0.49 | [6 empty] | #N/A
| 1 | 3 | 0.54 | 0.5 | 0.5 | 0.6 | 0.46 | 0.6 | 0.5 | 0.49 | [6 empty] | 16
| 1 | 4 | 0.55 | 0.5 | 0.5 | 0.6 | 0.472 | 0.6 | 0.4 | 0.49 | [6 empty] | 0.51331
| 1 | 5 | 0.55 | 0.5 | 0.5 | 0.6 | 0.498 | 0.6 | 0.4 | 0.49 | [6 empty] | 177
| 1 | 6 | 0.55 | 0.5 | 0.5 | 0.6 | 0.459 | 0.6 | 0.4 | 0.49
| Mean |  | 0.54667 | 0.5 | 0.5 | 0.6 | 0.4715 | 0.6 | 0.41667 | 0.48833
| Median |  | 0.55 | 0.5 | 0.5 | 0.6 | 0.466 | 0.6 | 0.4 | 0.49
| Std Dev. |  | 0.005164 | 0 | 0 | 0 | 0.019378 | 0 | 0.040825 | 0.0040825
| Rel.Std.Dev. |  | 0.9446% | 0% | 0% | 0% | 4.11% | 0% | 9.798% | 0.836%
| PDM3 |  | 6.498% | -2.593% | -2.593% | 16.89% | -8.146% | 16.89% | -18.83% | -4.866%
| Z-Score (Absolute) |  | 0.84 | 0 | 0 | 1.79 | 0.51 | 1.79 | 1.79 | 0.21
| Analytical results for U in OREAS 922 (Certified Value 3.59 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 21 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 3.68 | 3.7 | 3.6 | 4 | 2.796 | 3.818 | 3.4 | 3.29 | 3.55 | [5 empty] | 1
| 1 | 2 | 3.67 | 3.7 | 3.6 | 4 | 2.6056 | 3.574 | 3.4 | 3.21 | 3.52 | [5 empty] | #N/A
| 1 | 3 | 3.7 | 3.7 | 3.7 | 4 | 2.6236 | 3.562 | 3.6 | 3.32 | 3.62 | [5 empty] | 16
| 1 | 4 | 3.71 | 3.7 | 3.6 | 4 | 2.6742 | 3.905 | 3.7 | 3.35 | 3.54 | [5 empty] | 3.5926
| 1 | 5 | 3.7 | 3.7 | 3.7 | 4 | 2.832 | 3.67 | 3.6 | 3.47 | 3.61 | [5 empty] | 178
| 1 | 6 | 3.69 | 3.6 | 3.5 | 4.5 | 2.8719 | 3.53 | 3.7 | 3.33 | 3.67
| Mean |  | 3.6917 | 3.6833 | 3.6167 | 4.0833 | 2.7339 | 3.6765 | 3.5667 | 3.3283 | 3.585
| Median |  | 3.695 | 3.7 | 3.6 | 4 | 2.7351 | 3.622 | 3.6 | 3.325 | 3.58
| Std Dev. |  | 0.01472 | 0.040825 | 0.075277 | 0.20412 | 0.11375 | 0.1532 | 0.13663 | 0.084951 | 0.057533
| Rel.Std.Dev. |  | 0.3987% | 1.108% | 2.081% | 4.999% | 4.161% | 4.167% | 3.831% | 2.552% | 1.605%
| PDM3 |  | 2.758% | 2.526% | 0.67% | 13.66% | -23.9% | 2.335% | -0.7217% | -7.356% | -0.2114%
| Z-Score (Absolute) |  | 0.76 | 0.67 | 0 | 3.88 | 8.93 | 0.61 | 0.51 | 2.92 | 0.32
| Analytical results for V in OREAS 922 (Certified Value 92 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 0
1
| 1 | 1 | 85 | 91 | 96 | 100 | 96 | 94.18 | 87 | 93 | 87.9 | 90 |  |  |  |  | 1
| 1 | 2 | 82 | 89 | 85 | 100 | 93 | 93.959 | 85 | 92 | 93.4 | 88 |  |  |  |  | #N/A
| 1 | 3 | 84 | 89 | 93 | 100 | 95 | 89.474 | 104 | 96 | 91.9 | 89 |  |  |  |  | 16
| 1 | 4 | 82 | 92 | 87 | 100 | 96 | 93.558 | 102 | 94 | 92 | 86 |  |  |  |  | 92.2057
| 1 | 5 | 82 | 90 | 80 | 100 | 99 | 93.613 | 102 | 93 | 93 | 91 |  |  |  |  | 179
| 1 | 6 | 83 | 87 | 86 | 100 | 98.7052 | 93.386 | 98 | 95 | 95 | 91
| Mean |  | 83 | 89.6667 | 87.8333 | 100 | 96.2842 | 93.0283 | 96.3333 | 93.8333 | 92.2 | 89.1667
| Median |  | 82.5 | 89.5 | 86.5 | 100 | 96 | 93.5855 | 100 | 93.5 | 92.5 | 89.5
| Std Dev. |  | 1.2649 | 1.7512 | 5.7764 | 0 | 2.273 | 1.765 | 8.2624 | 1.472 | 2.3891 | 1.9408
| Rel.Std.Dev. |  | 1.524% | 1.953% | 6.577% | 0% | 2.361% | 1.897% | 8.577% | 1.569% | 2.591% | 2.177%
| PDM3 |  | -9.984% | -2.754% | -4.742% | 8.453% | 4.423% | 0.8922% | 4.477% | 1.765% | -0.006154% | -3.296%
| Z-Score (Absolute) |  | 2.01 | 0.67 | 1.04 | 1.42 | 0.67 | 0.16 | 0.68 | 0.17 | 0.16 | 0.77
| Analytical results for W in OREAS 922 (Certified Value < 5 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 6 | < 0.7 | < 0.7 | 5 | 3.3714 | 4.165 | 3.8 | 6 | [6 empty] | 1
| 1 | 2 | 6 | < 0.7 | 5.2 | < 5 | 3.4422 | 4.306 | 5.7 | 5 | [6 empty] | #N/A
| 1 | 3 | 6 | < 0.7 | 5.1 | < 5 | 3.7219 | 4.308 | 3.8 | 3 | [6 empty] | 16
| 1 | 4 | 6 | < 0.7 | 3.3 | < 5 | 3.4023 | 6.193 | 3.8 | 2 | [6 empty] | < 5
| 1 | 5 | 6 | < 0.7 | 4.5 | < 5 | 3.8799 | 4.302 | 3.6 | 5 | [6 empty] | 180
| 1 | 6 | 6 | 25 | 3.6 | 5 | 4.4354 | 4.067 | 3.5 | 6
| Mean |  | 6 | 25 | 4.34 | 5 | 3.7088 | 4.5568 | 4.0333 | 4.5
| Median |  | 6 | 25 | 4.5 | 5 | 3.582 | 4.304 | 3.8 | 5
| Std Dev. |  | 0 |  | 0.86197 | 0 | 0.40899 | 0.8075 | 0.82624 | 1.6432
| Rel.Std.Dev. |  | 0% |  | 19.86% | 0% | 11.03% | 17.72% | 20.49% | 36.51%
PDM3
| Z-Score (Absolute) |  | 3.04 | 5.41 | 0.55 | 0.95 | 0.39 | 0.39 | 0.4 | 0.79
| Analytical results for Y in OREAS 922 (Certified Value 31.1 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 13 | 16 | 18 | 19 | 20 | 21 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 32.4 | 30.6 | 30.1 | 32 | 27.4543 | 31.326 | 32.3 | 30 | 25.6 | 30.2 |  |  |  |  | 1
| 1 | 2 | 32.1 | 30.2 | 29.1 | 33 | 26.3649 | 31.223 | 31.5 | 30 | 27.1 | 30.7 |  |  |  |  | #N/A
| 1 | 3 | 32.5 | 30.2 | 29.8 | 31 | 25.8115 | 31.58 | 32.8 | 31 | 28.4 | 31.2 |  |  |  |  | 16
| 1 | 4 | 32.6 | 30.7 | 29.6 | 32 | 26.1968 | 30.991 | 32.1 | 31 | 25.9 | 30.5 |  |  |  |  | 31.0675
| 1 | 5 | 32.6 | 30.3 | 29.9 | 32 | 26.9205 | 31.284 | 32 | 30 | 22 | 30.5 |  |  |  |  | 181
| 1 | 6 | 32.5 | 30.1 | 29.5 | 31 | 27.5847 | 31.136 | 31.5 | 30 | 25.6 | 30.6
| Mean |  | 32.45 | 30.35 | 29.6667 | 31.8333 | 26.7221 | 31.2567 | 32.0333 | 30.3333 | 25.7667 | 30.6167
| Median |  | 32.5 | 30.25 | 29.7 | 32 | 26.6427 | 31.2535 | 32.05 | 30 | 25.75 | 30.55
| Std Dev. |  | 0.18708 | 0.2429 | 0.35024 | 0.75277 | 0.71446 | 0.19818 | 0.49666 | 0.5164 | 2.1435 | 0.33116
| Rel.Std.Dev. |  | 0.5765% | 0.8003% | 1.181% | 2.365% | 2.674% | 0.634% | 1.55% | 1.702% | 8.319% | 1.082%
| PDM3 |  | 4.45% | -2.309% | -4.509% | 2.465% | -13.99% | 0.6089% | 3.109% | -2.363% | -17.06% | -1.451%
| Z-Score (Absolute) |  | 1.22 | 0.08 | 0.51 | 0.84 | 2.34 | 0.48 | 0.96 | 0.09 | 2.94 | 0.08
| Analytical results for Yb in OREAS 922 (Certified Value 3.17 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 09 | 16 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 3.2 | 3.5 | 3.3 | 3 | 2.97 | 3.6 | 2.9 | 3 | [6 empty] | 1
| 1 | 2 | 3.2 | 3.5 | 3.2 | 3 | 2.98 | 3.6 | 3 | 2.9 | [6 empty] | #N/A
| 1 | 3 | 3.2 | 3.5 | 3.3 | 3 | 2.84 | 3.7 | 2.6 | 2.9 | [6 empty] | 16
| 1 | 4 | 3.2 | 3.4 | 3.3 | 3 | 3.02 | 3.8 | 3.1 | 2.9 | [6 empty] | 3.1742
| 1 | 5 | 3.2 | 3.4 | 3.3 | 3 | 3.05 | 3.6 | 3 | 3 | [6 empty] | 182
| 1 | 6 | 3.2 | 3.5 | 3.5 | 3 | 2.92 | 3.5 | 2.9 | 2.9
| Mean |  | 3.2 | 3.4667 | 3.3167 | 3 | 2.9633 | 3.6333 | 2.9167 | 2.9333
| Median |  | 3.2 | 3.5 | 3.3 | 3 | 2.975 | 3.6 | 2.95 | 2.9
| Std Dev. |  | 4.8648e-16 | 0.05164 | 0.098319 | 0 | 0.075011 | 0.10328 | 0.17224 | 0.05164
| Rel.Std.Dev. |  | 1.52e-14% | 1.49% | 2.964% | 0% | 2.531% | 2.843% | 5.905% | 1.76%
| PDM3 |  | 0.8139% | 9.215% | 4.489% | -5.487% | -6.642% | 14.47% | -8.112% | -7.587%
| Z-Score (Absolute) |  | 0.39 | 1.43 | 0.7 | 0.39 | 0.53 | 2.07 | 0.71 | 0.65
| Analytical results for Zn in OREAS 922 (Certified Value 277 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 09 | 10 | 13 | 15 | 16 | 18 | 20 | 21 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*AAS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 260 | 290 | 280 | 300 | 285 | 261.2632 | 300 | 302.16 | 300 | 278 | 259 |  |  |  | 1
| 1 | 2 | 230 | 290 | 270 | 300 | 280 | 270.6703 | 300 | 264.607 | 290 | 266 | 266 |  |  |  | #N/A
| 1 | 3 | 280 | 280 | 270 | 300 | 266 | 251.4151 | 300 | 274.768 | 340 | 280 | 279 |  |  |  | 16
| 1 | 4 | 240 | 310 | 260 | 300 | 270 | 262.0627 | 300 | 278.132 | 330 | 263 | 269 |  |  |  | 276.6399
| 1 | 5 | 245 | 290 | 260 | 300 | 294 | 266.7267 | 300 | 270.805 | 320 | 271 | 284 |  |  |  | 183
| 1 | 6 | 245 | 330 | 270 | 300 | 280 | 260.8353 | 300 | 254.951 | 320 | 257 | 271
| Mean |  | 250 | 298.3333 | 268.3333 | 300 | 279.1667 | 262.1622 | 300 | 274.2372 | 316.6667 | 269.1667 | 271.3333
| Median |  | 245 | 290 | 270 | 300 | 280 | 261.663 | 300 | 272.7865 | 320 | 268.5 | 270
| Std Dev. |  | 17.6068 | 18.3485 | 7.5277 | 0 | 10.1275 | 6.5008 | 0 | 15.9475 | 18.619 | 8.8863 | 9.0037
| Rel.Std.Dev. |  | 7.043% | 6.15% | 2.805% | 0% | 3.628% | 2.48% | 0% | 5.815% | 5.88% | 3.301% | 3.318%
| PDM3 |  | -9.63% | 7.842% | -3.003% | 8.444% | 0.9134% | -5.233% | 8.444% | -0.8686% | 14.47% | -2.701% | -1.918%
| Z-Score (Absolute) |  | 1.35 | 0.99 | 0.33 | 1.44 | 0.28 | 0.67 | 1.44 | 0 | 2.37 | 0.28 | 0.16
| Analytical results for Zr in OREAS 922 (Indicative Value 206 ppm) | [15 empty] | Indicative
Zr | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 13: Lab |  |  |  | 1
| No | No | 02: 09; 04: 16; 05: 19; 13: 21 |  |  |  | ppm
|  |  | 02: PF*MS; 04: PF*OES/MS; 05: PF*MS; 13: PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 02: 220; 04: 206.17; 05: 211.7; 13: 199 |  |  |  | 1
| 1 | 2 | 02: 220; 04: 198.53; 05: 195.9; 13: 184 |  |  |  | 9
| 1 | 3 | 02: 210; 04: 209.29; 05: 212.9; 13: 216 |  |  |  | 16
| 1 | 4 | 02: 220; 04: 204.91; 05: 200.6; 13: 191 |  |  |  | 206.1071
| 1 | 5 | 02: 220; 04: 203.03; 05: 200.8; 13: 192 |  |  |  | 15
| 1 | 6 | 02: 210; 04: 203.94; 05: 205.8; 13: 211
| Mean |  | 02: 216.6667; 04: 204.3117; 05: 204.6167; 13: 198.8333
| Median |  | 02: 220; 04: 204.425; 05: 203.3; 13: 195.5
| Std Dev. |  | 02: 5.164; 04: 3.5722; 05: 6.736; 13: 12.4164
| Rel.Std.Dev. |  | 02: 2.383%; 04: 1.748%; 05: 3.292%; 13: 6.245%
| PDM3 |  | 02: 5.123%; 04: -0.8711%; 05: -0.7231%; 13: -3.529%
| Z-Score (Absolute) |  | 02: 2.85; 04: 0.04; 05: 0.04; 13: 1.31
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.